## Sheet1
Recall_Number | Date | Year | Trade_Name | Recalling_Firm | Recall_Class | Reason_Recall | Action
Z-0863-06 | May-16-2006 | 2006 | CONTAK RENEWAL 3 RF & 4 RF - CONTAK RENEWAL 3 RF (models H210; H215); CONTAK RENEWAL 3 RF HE (models H217; H219); CONTAK RENEWAL 4 RF (models H230; H235) CONTAK RENEWAL 4 RF HE (model H239) cardiac resynchronization therapy defibrillator (CRT-D); provide ventricular tachyarrhythmia and cardiac resynchronization therapies. Ventricular tachyarrhythmia therapy is for the treatment of ventricualar tachycardia (VT) and ventricular fibrillation (VF); rhythms that are associated with suden cardiac death (SCD). Cardiac resynchronization therapy is for the treatment of heart failure (HF) and uses biventricular electrical stimulation to synchronize ventricular contractions. The device also uses accelerometer-based adaptive-rate bradycardia therapy. These CONTAK RENEWAL RF devices feature enhanced telemetry communication with wandless ZIP telemetry. Cardioversion/defibrillation therapies include a range of low- and high-energy shocks using either a biphasic or monophasic waveform. Guidant Corporation; 4100 Hamline Avenue North; St. Paul; MN 55112-5798 USA. | Guidant Corporation | 3 | Guidant received 39 reports of devices that exhibited lower than expected battery voltage prior to implant; none of these devices were implanted. | Affected physicians were sent an Urgent Medical Device Safety Information & Corrective Action letter, dated 05/11/06. A press release was issued March 13, 2006 The letter describes the issue and emphasizes the importance of checking battery voltage prior to implant as stated in device labeling. For implanted devices, continue normal follow-up.
Z-0005-2007 | Oct-03-2006 | 2006 | Stryker Navigation System - Stryker Navigation System; eNlite System with Dell laptop computer model D800; Stryker # 7700-300-000. | Stryker Instruments Div. of Stryker Corporation | 2 | A component of the device is a Dell laptop computer; which contains batteries that are under recall because they can overheat and cause a fire. | A recall letter dated 9/8/06 was sent to each customer instructing them to remove the battery from their Dell laptop computer model D800, to check the batteries to determine whether they are affected by the recall, and to only use the computer on AC power until they have requested and received replacement batteries.
Z-0859-06 | May-10-2006 | 2006 | Electrosurgical; cutting & coagulation & accessories. - Cool-tip RF Generator; for use in percutaneous; laparoscopic; intraoperative coagulation and ablation of tissue. | Covidien LP | 2 | Electrosurgical ablation device may experience loss of power during use; when used with Cool-tip Switching controller in ablation mode. | Consignees were notified by telephone followed by a fax on 04/21/2006.
Z-0201-2007 | Nov-21-2006 | 2006 | Datex-Ohmeda Tec 6 Plus Desflurane Vaporizer - Datex-Ohmeda Tec 6 Plus Desflurane Vaporizer. It is an electronic vaporizer which delivers the anesthetic agent desflurane. | Datex - Ohmeda; Inc | 2 | There is a possibility of a vaporizer power failure; with audible and visual alarms; when using the Tec 6 Plus Desflurane Vaporizer. | An Urgent Medical Device Correction letter dated September 15, 2006, was sent to affected customers. The letter describes the issue, states that a GE representative will contact affected accounts to arrange for the vaporizer exchange at no charge.
Z-0220-2007 | Nov-28-2006 | 2006 | Konica Minolta PrintLink III - Konica Minolta PrintLink III Model - ID/IV; code no. 0770; 0771 Medical Imaging Communication DeviceModel IV - product code 5000230 Model ID; product code 5000330Konica Minolta Medical & Graphic; Inc.No. 26-2; Nishishinjuku 1-chrome; Shinjuku-ku; Tokoyo 163-0512; Japan | Konica Minolta Medical Imaging USA; Inc. | 2 | The power supply of a unit located in Japan overheated and caused a smoke condition and a small amount of flame coming off the unit. | Letters were sent to all customers by first class US mail, certified receipt requested on 10/10/2006. Customers were also notified by fax and e-mail and a phone call on the same day.
Z-0444-06 | Jan-31-2006 | 2006 | Colleague and Colleague CX Volumetric Infusion Pumps - Baxter Colleague Single Channel Volumetric Infusion Pumps; Baxter Healthcare Corporation; Medication Delivery Division; Deerfield; IL 60015 U.S.A.; Made in Singapore; product codes 2M8151; 2M8151R; BRM8151(Brazilian Portugese); DNM8151(French); HNM8151(German); PNM8151(Spanish); CNM8151(Swedish); GNM8151(Danish); TRM8151(Turkish); WNM8151(Dutch) ; 2M8161; 2M8161R | Baxter Healthcare Corp. | 1 | Interruption of therapy due to battery undercharging; the generation of air-in-line alarms due to IV administration set tugging; and gearbox wear; and underinfusion due to misalignment of the pump head components. | Baxter sent the 12/13/05 Urgent Device Correction letter to all Colleague Infusion Pump customers, to the attention of the Director of Nursing and the Biomedical Engineer, via first class mail to alert them to additional problems with the pump that may interrupt therapy. These issues include battery undercharging, generation of false air detected alarms due to IV administration set tugging, gearbox wear, underinfusion and non-detection of upstream occlusions. The letters included mitigating instructions to reduce the occurrence of these potential problems, and included a copy of the Battery Usage Guide, Infusion Management Guide, and the March 17, 2005 Buretrol Set Urgent Device Correction letter to be provided to all users of the Colleague pump. Baxter informed the accounts that they would be notified when the new release of the Colleague Volumetric Infusion Pump Operator''s Manual is available. Any questions were directed to Baxter''s Medication Delivery Services at 1-800-843-7867.
Z-1413-06 | Sep-12-2006 | 2006 | D-tron - Disetronic D-TRONplus Insulin Pump Battery Pack; Ref/Catalog no. 04697014001. | Disetronic Medical Systems; Inc. | 1 | The battery may turn the pump off without warning due to a design change in the battery. | Press release was issued on 7/13/06. Consignees, distributors and health care providers were issued a recall notice dated 7/13/06 instructing replacement of the power pack every two weeks until further notice.
Z-1522-06 | Sep-23-2006 | 2006 | Medline Strider Midi 3 Scooter - Medline Strider Midi 3 Scooter; a three wheeled battery operated scooter; manufactured by KYMCO; Kaohsiung; Taiwan for Medline Industries; Inc.; One Medline Place; Mudelein; IL 60060 USA; model MDS807500 - red and MDS807500B - blue | Medline Industries Inc | 3 | The P&G Solo 60A electronic control unit on the Midi Scooters may overheat. | Medline sent 'Voluntary Field Correction' letters dated 7/20/06 to the direct accounts on the same date, informing them that the Strider Midi 3 and Midi 4 scooters have an older version of the electronic control unit that may overheat, which Medline will replace with a newer version of the controller. Any questions were directed to 800-289-9793, press 1, then press 3 and ask to transfer to 4674 (8:00 AM to 5:00 PM CST).
Z-0718-06 | Apr-07-2006 | 2006 | Knee navigation trackers. - AxiEM EM Knee Kit; and associated workstation. | Medtronic Navigation; Inc | 3 | Knee Navigation Disposable Kits were programmed with an incorrect identification number on the tracker ROM chip which would not permit navigation during surgery. | Medtronic field personnel were notified by letter on 02/01/2006 and instructed to immediately contact each affected site and schedule an upgrade with the revised firmware.
Z-1453-06 | Aug-29-2006 | 2006 | TCm monitoring system - TCM4 Series Monitoring System (Base Unit) 391-876 (affected device) CompactFlash cards: 914-698 (Defective Device Component) | Radiometer America Inc | 3 | System shut down-When the TCM Monitor is turned on and the booting process begins; the device stops after the memory count and will not proceed further. | The firm began a recall/corrective action in response to this component defect on 4/7/2006, at which time, the firm issued a Field Action Notes memo (915-234). The firm is having their representatives: 1) assure that the use of the CompactFlashcards manufactured by Simple Tech has been discontinued at each customer location; 2) Replaace the CompactFlash card in all basic Units, 391-876, of R0036 through R0043; and 3) replace any spare part #914-698, which has been installed in Basic Units, 391--876, below R0036. Once completed, the firm has supplied their customers with a TELEFAX form to be completed, signed and dated by each customerand faxed back to Radiometer Medical ApS. The form documents that the instructions provided in theField Action Notes have been followed and that the necessary component replacements have been completed. The firm anticipates that the necesary corrective action will be completed by 7/1/2006.
Z-0861-06 | May-27-2006 | 2006 | Kinetra Implantable Neurostimulator - Medtronic Kinetra 7428. Dual Program Neurostimulator for Deep Brain Stimulation. Rx only. Medtronic; Inc.; 710 Medtronic Parkway; Minneapolis; MN 55432-5604 USA. Kinetra neurostimulator is a component of the Activa System. The Medtronic Activa System is an implantable; multiprogrammable system that delivers electrical stimulation to selected areas of the brain. The dual program Model 7428 Kinetra Neurostimulator generates electrical signals that are transmitted to the brain. The neurostimulator consists of electronic circuitry and a battery; which are hermetically sealed in a titanium case. | Medtronic Neurological | 3 | A specific subset of model 7428 Kinetra implantable neurostimulators may fail due to lifted wirebonds between the hybrid circuit and battery. This failure mechanism may present clinically as loss of functionality and a return of the patient's underlying medical conditions. | An Urgent Device Recall Important Patient Management Information letter was sent to implanting and following physicians of patients who have received the device. The letter describes the issue and provides recommendations.
Z-0547-06 | Feb-23-2006 | 2006 | LIFEPAK CR Plus - LIFEPAK CR Plus defibrillator | Medtronic Emergency Response Systems; Inc. | 2 | One of the printed circuit board assemblies may be contaminated with solder flux which can cause a short that may render the device inoperative. | On 1/31/06 the firm sent a letter flagged URGENT-MEDICAL DEVICE RECALL to their customers via certified return receipt. The letter advised customers of the problem and requested the customers inspect their units at least weekly. The customers are provided additional guidance if they see certain symbols as described in the letter. The firm''s service personnel will visit the customers, remove and replace units.
Z-1294-06 | Sep-28-2006 | 2006 | INSIGNIA Ultra and NEXUS Ultra - Guidant INSIGNIA I Ultra --family of pacemakers includes the following: SR (models1190); and DR (models1290; 1291). Intermedics NEXUS I Ultra family of pacemakers includes the following: SR (model 1390) and DR (model 1490; 1491). The Intermedics NEXUS I Ultra (models 1390; 1490; 1491) are not available in the US. The INSIGNIA I Ultra pacemakers are multiprogrammable pacemakers from Guidant.--The NEXUS I Ultra pacemakers are multiprogrammable pacemakers from Intermedics. Guidant Corporation; Cardiac Rhythm Management; 4100 Hamline Avenue North; St. Paul; MN 55112-5798 USA. or Intermedics; Inc.; 4100 Hamline Avenue North; St. Paul; MN 55112-5798 USA. Quantities Distributed by model are as follows: INSIGNIA ULTRA SR model 1190 = 4;590; DR model 1290 = 9;654; model 1291 = 4;038. NEXUS ULTRA SR model 1390 = 15; DR model 1490 = 52; model 1491 = 57. | Boston Scientific CRM Corp | 2 | Guidant has determined that low-voltage capacitors from a single component supplier may malfunction in a manner that can lead to premature battery depletion or loss of pacing output without warning in the affected devices. | A Press Release was issued on 06/26/06. An Urgent Medical Device Safety Information & Corrective Action letter, dated 06/23/06 was delivered to physicians describing preliminary information before Guidant has completed its own investigation and finalized patient care recommendations. Non-implanted inventory is being retrieved from hospital shelves and Guidant Represenatatives. Physicians are to perform an in-clinic follow-up exam as soon as possible with patients implanted with affected devices.
Z-0535-06 | Feb-18-2006 | 2006 | Sigma Pacemaker - Sigma implantable pulse generators (IPGs) Dual Chamber Rate Responsive Pacemaker; models SDR203; SDR303; SDR306 | Medtronic Inc. Cardiac Rhythm Managment | 2 | An issue exists with a specific subset of Sigma series pacemakers that may fail due to separation of interconnect wires from the hybrid circuit. This failure mechanism may present clinically as loss of rate response; premature battery depletion; intermittent or total loss of telemetry; or not output. There have been no reported patient injuries or deaths due to this issue. | Press Release was issued 11/29/05. Hospital/physician letters sent via registered mail on 11/29/05 for delivery on 12/01/05. Unimplanted devices are being retrieved. The Physician letter describes the issue, root cause, the probability of reoccurrence and provides recommendations. A list of affected serial numbers that the physician is following or has implanted is attached to the letter. The letter also provides medtronic''s website in which regular updates on the ongoing actual performance are listed in their Product Performance Report. Medtronic is providing a replacement device should the Physician and patient elect to replace the affected implanted device.
Z-1299-06 | Sep-28-2006 | 2006 | VITALITY and VITALITY 2 - VITALITY (models 1870; 1871; T125; T127; T135) and VITALITY 2 (models T165; T167; T175; T177). Implantable Cardioverter Defibrillators (ICD). Quantities Distributed by model are as follows: VITALITY model 1870 = 40; model 1871 = 1; model T125 = 613; model T127 = 49; model T135 = 536. VITALITY2 model T165 = 1;994; model T167 = 710; model T175 = 2;010; model T177 = 879. | Boston Scientific CRM Corp | 2 | Guidant has determined that low-voltage capacitors from a single component supplier may malfunction in a manner that can lead to premature battery depletion or incorrect sensing in the affected devices. | A Press Release was issued on 06/26/06. An Urgent Medical Device Safety Information & Corrective Action letter, dated 06/23/06 was delivered to physicians describing preliminary information before Guidant has completed its own investigation and finalized patient care recommendations. Non-implanted inventory is being retrieved from hospital shelves and Guidant Represenatatives. Physicians are to perform an in-clinic follow-up exam as soon as possible with patients implanted with affected devices.
Z-0290-2007 | Dec-19-2006 | 2006 | Medtronic Kinetra Implantable Neurostimulator - Medtronic Kinetra Dual Program Neurostimulator for Deep Brain Stimulation; model 7428. Rx only. Medtronic; Inc.; 710 Medtronic Parkway; Minneapolis; MN 55432-5604 USA. The dual program model 7428 Kinetra Neurostimulator generates electrical signals that are transmitted to the brain. These signals are delivered from the neurostimulator to the brain via DBS extension or DBS leads. The neurostimulator consists of electronic circuitry and a battery; which are hermetically sealed in a titanium case. The operation of the neurostimulator is supported by a clinician programmer; a therapy controller; and a control magnet. | Medtronic Neurological | 2 | A subset of Kinetra implantable neurostimulators may experience a failure of wire connections between the electronic hybrid circuit and battery which may lead to sudden cessation of therapy. Sudden cessation of therapy can result in the immediate return or worsening of underlying symptoms due to the progression of the disease state. | Letter dated July 2006. Recommendations were made to Health Care Providers of the following: 1.Consider the potential for this anomaly if your patient present with a device in a POR state. 2. Consider informing your patients of this risk and advise them to seek attention immediately if they experience a return of pre-implant or more advanced symptoms. 3. Determine whether device replacement is warranted based upon consultation with your patient, review of the patient''s medical history and consideration of the relative risks of abrupt cessation of stimulation versus the risk of a device replacement surgical procedure. 4. Report any malfunction, removal and replacement to Medtronic and FDA suing the medical device reporting system. 5. Acknowledge this communication using the reply form enclosed.
Z-1296-06 | Sep-28-2006 | 2006 | INSIGNIA AVT and NEXUS AVT - Guidant INSIGNIA I AVT --family of pacemakers includes the following: SSI (model 482); VDD (model 882); DDD (model 982); SR (model 1192 and DR (model 1292). Intermedics NEXUS I AVT family of pacemakers includes the following: VDD (model 1428); DDD (model 1432); SR (model 1392) and DR (model 1492). The INSIGNIA I AVT pacemakers are multiprogrammable pacemakers from Guidant.-- The NEXUS I AVT pacemakers are multiprogrammable pacemakers from Intermedics. Guidant Corporation; Cardiac Rhythm Management; 4100 Hamline Avenue North; St. Paul; MN 55112-5798 USA. or Intermedics; Inc.; 4100 Hamline Avenue North; St. Paul; MN 55112-5798 USA. Quantities Distributed by model are as follows: INSIGNIA AVT SSI model 482 = 303; VDD model 882 = 789; DDD model 982 = 656; SR model 1192 = 988; DR model 1292 = 1;408. NEXUS AVT VDD model 1428 = 67; SR model 1392 = 19; DR model 1492 = 39. ****The following devices are not available in the US. ***** | Boston Scientific CRM Corp | 2 | Guidant has determined that low-voltage capacitors from a single component supplier may malfunction in a manner that can lead to premature battery depletion or loss of pacing output without warning in the affected devices | A Press Release was issued on 06/26/06. An Urgent Medical Device Safety Information & Corrective Action letter, dated 06/23/06 was delivered to physicians describing preliminary information before Guidant has completed its own investigation and finalized patient care recommendations. Non-implanted inventory is being retrieved from hospital shelves and Guidant Represenatatives. Physicians are to perform an in-clinic follow-up exam as soon as possible with patients implanted with affected devices.
Z-0546-06 | Feb-22-2006 | 2006 | LIFEPAK 500 - LIFEPAK 500 automated external defibrillator (AED) | Medtronic Emergency Response Systems; Inc. | 2 | Component leads may have intermittent connection to the printed circuit board assembly pads - which could render the device inoperative. | On 2/1/06 the firm issued an UREGENT MEDICAL DEVICE CORRECTION letter to customers via certified return receipt. The letter advised customers of the issue and recommended customers inspect their AED at least weekly. The letter also advised customers what to do when readiness indicator displays certain symbols. The firm''s representatives will visit customers to replace the main circuit board assembly.
Z-0411-06 | Jan-28-2006 | 2006 | Baxter 6060 & Sabratek 6060 Homerun Multi-Therapy Infusion Pump - Baxter 6060 Multi-Therapy Infusion Pump; product codes 2M9832; 2M9832P and 2M9832R; and Sabratek 6060 Homerun Infusion Pumps; product codes 606000; 606000-40; 606000-40L; and 606000-40I; Manufactured by an affiliate of Baxter Healthcare Corp.; Deerfield; IL 60013 USA; made in Singapore | Baxter Healthcare Corp. | 2 | Possible shorting of some printed circuit boards could result in a possible interruption of therapy or cause the pump to fail the Power-On Self Test. | Baxter sent Urgent Device Correction letters dated 11/7/05 to the user accounts, advising them of the potential for the circuit boards in the 6060 pump to fail due to a problem that occurred in the supplier's manufacturing process. The letter included a list of 6060 pump serial numbers sent to the user that may contain an affected circuit board. To mitigate safety risks to the patient, Baxter advised customers to immediately discontinue use of the pump for therapies where interruption of the infusion could cause immediate patient harm. In addition, the pump is not to be used if it fails the Power-On Self Test. The accounts were instructed to contact Baxter Medication Delivery Services at 1-800-843-7867 for information regarding the return process for the affected 6060 pumps for correction. Baxter sent follow-up letters dated 2/16/06 to the accounts, to the attention of the Director of Biomedical Engineering, informing them that Baxter has sufficient quantities of replacement circuit boards to begin the repairs. The accounts were requested to review the enclosed list of serial numbers for accuracy and telephone Baxter at 1-800-843-7867 to schedule the return of any 6060 Infusion Pump that requires replacement of the circuit board.
Z-0308-2007 | Dec-27-2006 | 2006 | Philips/Laerdal HeartStart FR2+ Automated External Defibrillator - HeartStart FR2+ Defibrillator | Philips Medical Systems | 2 | Potential for shock treatment to be delayed or prevented due to board contamination that causes intermittent switch operation | On 12/19/06 the firm sent letters via return receipt to all domestic customers. The letter advised customers of the issue and if the user encounters the problem, to press the button again with additional force. The users are to contact Philips for replacement. The devices can be used until the user receives a replacement.
Z-1386-06 | Aug-10-2006 | 2006 | Blood Gas and Co-Ox Electrolyte & Metabolyte Analyzers - Blood Gas and Co-Ox2 Electrolyte & Metabolyte Analyzers; ABL-700 Series; manufactured from 8/1998 through 11/2004. | Radiometer America Inc | 2 | Under certain circumstances; the ABL 700 & 800 Series Blood Gas Analyzers may experience leakage current into the measuring system. Consequently; the analyzers intermittently provide incorrect (too low) result values for Calcium (Ca) and Sodium (Na). | The consignees were notified of this recall by a letter, dated 9/29/2005. The letter informs the customers of this potential problem and the availability of new updated software which will correct the problem. The letter goes on to inform the customer that a Radiometer Service Engineer wil be contacting them shortly in order to arrange for a convenient time to install the software update. Along with the letter, the customer is provided a copy of an additional page and a CD concerning the function and operation of the new software, which is to be added to the customer''s Operator''s Manual.
Z-1298-06 | Sep-28-2006 | 2006 | CONTAK RENEWAL TR and CONTAK RENEWAL TR2 - CONTAK RENEWAL TR (models H120; H125) and CONTAK RENEWAL TR2 (models H140; H145). CONTAK RENEWAL TR2 devices are not available in the US. Quantities Distributed by model are as follows: CONTAK RENEWAL TR model H120 = 359; model H125 = 235; TR2 model H140 = 613; model H145 = 390. | Boston Scientific CRM Corp | 2 | Guidant has determined that low-voltage capacitors from a single component supplier may malfunction in a manner that can lead to premature battery depletion without warning in the affected devices. | A Press Release was issued on 06/26/06. An Urgent Medical Device Safety Information & Corrective Action letter, dated 06/23/06 was delivered to physicians describing preliminary information before Guidant has completed its own investigation and finalized patient care recommendations. Non-implanted inventory is being retrieved from hospital shelves and Guidant Represenatatives. Physicians are to perform an in-clinic follow-up exam as soon as possible with patients implanted with affected devices.
Z-0514-06 | Feb-14-2006 | 2006 | Microlab AT; Microlab AT+; Microlab AT+2 - Microlab AT; part numbers 135001; 135001OOC; 135002; 135003; 135005; 135006; 135006R; 135007; 135008; 135016; 135028; 135029; 936400; 936400BLU. Microlab AT+ and Microlab AT+2; part numbers 142000; 142001; 142001R; 142002; 142003; 142005; 142006; 142006R; 142007; 142008; 142009; 142015; 142017; 142035; 142036 | Hamilton Co | 2 | Heat sinking on power boards may not operate as intended and the board may cease to operate. | On 12/15/05, the firm initiated the recall and its notification was via letters explaining the reason for the recall and requesting the product be returned.
Z-0047-2007 | Oct-19-2006 | 2006 | Esprit - Esprit Ventilator; Model Number V1000; Power Supply PN 1015852 and Power Supply Field Replacement Unit (FRU) PN; Continuous Ventilator/Continuous Ventilator with Respiratory Mechanics 1018246 | Respironics California Inc | 2 | This action is being initiated due to power supply snubber board failures on certain Esprit Ventilators in countries that utilize operating voltages of 200-240 Volts AC and will involve the replacement ofthe Esprit Ventilator power supply snubber board. Returned power supplies have shown evidence of overheating and in some cases in which a fan also failed; have shown signs of fire damage. | The field action will be conducted through consignees (distributors) in the affected countries. The distributors were notified of the field action by Federal Express au on August 8th, 2006. The distributors have been informed that Respironics will provide replacement circuit boards at no cost and will be instructed to upgrade the ventilator power supplies in their region. Instructions have been provided to the distributors and users on how to evaluate the ventilator to be able to continue to use it pending the completion of the upgrades. The distributors will document the replacement of the affected component by signing and dating the activity on the Ventilator Field Action Log (reference sample on Appendix IV). The serial number of the power supply destroyed will also be recorded on the log.
Z-1292-06 | Jul-27-2006 | 2006 | Guidant VITALITY - Guidant VITALITY 2 (models T165; T175); VITALITY DS (model T125); VITALITY EL (model T135); Implantable Cardioverter Defibrillators (ICD); are designed to detect and terminate ventricular tachycardia (VT) and ventricular fibrillation (VF) and provide bradycardia therapy (atrial and ventricular pacing). Therapies include both low- and high-energy shocks using either a biphasic or monophasic waveform. Vitality 2 devices also offer a wide variety of antitachycardia pacing schemes to terminate slower; more stable ventricular tachyarrhythmias. Bradycardia pacing; including adaptive-rate features; is available to detect and treat bradyarrhythmias and to support the cardiac rhythm after defibrillation therapy. Devices denoted with DR offer dual-chamber bradycardia features (atrial and/or ventricular pacing and sensing); and the devices denoted with VR offer single-chamber bradycardia features (ventricular pacing and sensing). | Boston Scientific CRM Corp | 2 | Guidant has identified a pattern of premature battery depletion in devices manufactured with a single lot of capacitors from a single supplier. If function of this low voltage capacitor is compromised; the battery may deplete prematurely. | An Urgent Medical Device Safety Information & Corrective Action letter, dated May 12, 2006, was sent to physicians. The letter gave a description of the issue, clinical implications, rate of occurrence, devices affected and recommendation. Physicians are asked to follow-up with affected patients as soon as possible because ERI to EOL time may be shortened. Physicians are asked to contact Guidant for instructions on performing a baseline 'save to disk' and provide this information to Guidant. Guidant will be able to analyze the device memory data and estimate remaining longevity and provide individualized follow-up and replacement guidelines. Subsequent follow-up visits should continue to monitor for evidence of premature battery depletion.
Z-0893-06 | May-23-2006 | 2006 | High Frequency Ventilator - Life Pulse High Frequency Ventilator (HFV) | Bunnell; Inc. | 2 | Potential loss of air flow or change in cycling of the ventilator. Valve Drive Regulator Board may not function properly at elevated temperatures. The increase in temperature could cause a reduction of power; causing the solenoid in the valve to function sporadically. | Notified consignees by letter on 04/04/2005.
Z-1103-06 | Jun-10-2006 | 2006 | Hitachi Altaire Magnetic Resonance Imaging Device - Hitachi Altaire Magnetic Resonance Imaging System | Hitachi Medical Systems America Inc | 2 | Circuit Failure-There is the potential risk of the Phase Loss Detector Safety Device (PHSEN PCB) malfunctioning in the event of a power loss. The circuit may fail to turn off the system main breaker when 2 of the 3 power phases are lost. | The recalling firm sent a Device Correction Letter dated 1/27/06 to all of their customers. The letter informs the customers of the problem in the (PHSEN PCB) and notifies the customers that an HMSA Field Service Technician will be performing a PCB modification to the unit, on-site, during their next scheduled preventative maintenance visit. Should the customer wish to have the modification performed prior to the representative''s scheduled visit, the letter says that they can telephone their local Field Service Rep. and make arrangements to have the modification performed at their convenience. .
Z-0300-2007 | Dec-19-2006 | 2006 | BCI Autocorr Plus Pulse Oximeter/ECG/Respiration Monitor - BCI 3404 Autocorr Plus Pulse Oximeter/ECG/Respiration Monitor. Smiths Medical PM; Inc. This device is designed to provide full featured monitoring capabilities in a tabletop design. The system features an ECG cable interface; an SpO2 probe interface; display of patient and waveform data via an EL panel; power status LED; and the function keypad area consisting of six keys (on/off; waveform/trend; alarm silence; menu/enter; up arrow & down arrow). The monitor has a serial port that is used for data communications to a printer or computer and for analog outputs. | Smiths Medical PM; Inc. | 2 | An error occurred at the board manufacturing site which may affect the following parameters: - An electrical noise observed on the SpO2 pleth waveform generating and displaying erratic heart rates. % SpO2 reading displays dashes. - The ECG waveform may appear noisy; ECG readings appear accurate. -Respiration waveform may appear noisy and the respiration readings will be displayed as dashes. | An Safety Action Bulletin, dated 11/03/2006 was sent to affected customers describing the issue, what parameters are affected on the monitor and provides informatin on how to have the monitor repaired. Smiths Medical MDPM requests that an acknowledgement form be returned to them.
Z-1387-06 | Aug-10-2006 | 2006 | Blood Gas and Co-Ox Electrolyte & Metabolyte Analyzers. - Blood Gas and Co-Ox Electrolyte & Metabolyte Analyzers; ABL-800 series; manufactured from 8/2004 through 9/2005. | Radiometer America Inc | 2 | Under certain circumstances; the ABL 700 &800 Series Blood Gas Analyzers may experience leakage current into the measuring system. Consequently; the analyzers intermittently provide incorrect (too low) result values for Calcium (Ca) and Sodium (Na). | The consignees were notified of this recall by a letter, dated 9/29/2005. The letter informs the customers of this potential problem and the availability of new updated software which will correct the problem. The letter goes on to inform the customer that a Radiometer Service Engineer wil be contacting them shortly in order to arrange for a convenient time to install the software update. Along with the letter, the customer is provided a copy of an additional page and a CD concerning the function and operation of the new software, which is to be added to the customer''s Operator''s Manual.
Z-1297-06 | Sep-28-2006 | 2006 | VENTAK PRIZM 2 - Guidant VENTAK PRIZM VR (model 1860) and VENTAK PRIZM DR (model 1861) Automatic Implantable Cardioverter Defibrillator (AICD); Dual Chamber; Sterilized with gaseous ethylene oxide. Guidant Corporation; Cardiac Rhythm Management; 4100 Hamline Avenue North; St. Paul; MN 55112-5798 USA. Quantities Distributed by model are as follows: VENTAK PRIZM VR model 1860 = 67; DR model 1861 = 108. | Boston Scientific CRM Corp | 2 | Guidant has determined that low-voltage capacitors from a single component supplier may malfunction in a manner that can lead to premature battery depletion in the affected devices. | A Press Release was issued on 06/26/06. An Urgent Medical Device Safety Information & Corrective Action letter, dated 06/23/06 was delivered to physicians describing preliminary information before Guidant has completed its own investigation and finalized patient care recommendations. Non-implanted inventory is being retrieved from hospital shelves and Guidant Represenatatives. Physicians are to perform an in-clinic follow-up exam as soon as possible with patients implanted with affected devices.
Z-0623-06 | Mar-02-2006 | 2006 | Powerheart AED G3 - Powerheart AED G3 Automated External Defibrillator Model 9300A (Automatic). Cardiac Science AEDs are manufactured by: Manufacturing: Cardiac Science; Inc.; 5474 Feltl Road; Minnetonka; MN 55343-7982. | Cardiac Science Corporation | 2 | Some Powerheart AED Model 9300A may contain a suspect electronic component and there is a remote likelihood that the function of this device may be interrupted during the charge. | Affected consignees were notified via telephone. If no phone contact could be made, a Urgent: Medical Device Recall letter dated 02/02/06 was sent with return receipt requested. Customers were asked to locate affected devices and remove them from service. A new AED would be sent to the customers and the customers are asked to return the affected device.
Z-1508-06 | Sep-19-2006 | 2006 | Electrosurgical Unit - Conmed System 5000 Electrosurgical unit; REF 60-8005-SYS; 60-8015-SYS; 60-8018-SYS; 60-8005-001; 60-8005-003; Conmed Corporation. | Conmed Electrosurgery | 2 | Wrong size capacitor installed in electrosurgical generator could malfunction; causing muscle stimulation or delay in therapy. | Domestic customers were notified by letter on 09/01/2006 and international accounts were notified by letter on 09/07/2006.
Z-0825-06 | May-02-2006 | 2006 | CONTAK RENEWAL 3 - Guidant CONTAK RENEWAL 3 (models H170;H175) CONTAK RENEWAL 3 HE (models H177; H179) Cardiac Resynchronization Therapy Defibrillator (CRT-D). Sterile EO. Guidant Corporation; 4100 Hamline Avenue North; St. Paul; MN 55112-5798 USA. Cardiac resynchronization therapy defibrillators (CRT-D) provide ventricular tachyarrhythmia and cardiac resynchronization therapies. Ventricular tachyarrhythmia therapy is for the treatment of ventricular tachycardia (VT) and ventricular fibrillation (VF); rhythms that are associated with sudden cardiac death (SCD). Cardiac resynchronization therapy is for the treatment of heart failure (HF) and uses biventricular electrical stimulation to synchronize ventricular contractions. Cardioversion/defibrillation therapies include a range of low-and high-energy shocks using either a biphasic or monophasic waveform. | Guidant Corporation | 2 | 77 distributed devices may have a cracked layer of insulation in a flexible hybrid circuit; leading to internal shorting during charging such as during routine capacitor formation or for defibrillation. This would likely render the device non-functional; resulting in no output and/or telemetry communications. | Guidant Sales Representatives are retrieving non-implanted devices April 10-12, 2006. If a device is implanted the physician is sent an Urgent Medical Device Safety Information & Corrective Action letter, April 11-14, 2006. This letter describes the issue, states that Guidant is retrieving all non-implanted devices and that Guidant is currently conducting tests on retrieved devices. Physicians will be advised once more information is available.
Z-0269-2007 | Dec-16-2006 | 2006 | Guidant CONTAK RENEWAL 3 RF & 4 RF - Guidant CONTAK RENEWAL 3 RF (models H210; H215); CONTAK RENEWAL 3 RF HE (models H217; H219) ;CONTAK RENEWAL 4 RF (model H230); and CONTAK RENEWAL 4 RF HE (model H239) cardiac resynchronization therapy defibrillator (CRT-D). Guidant Corporation; 4100 Hamline Avenue North; St. Paul; MN 55112-5798 USA. | Guidant Corporation | 2 | Shock Effectiveness: there is a potential for malfunction of a high voltage wire; which could compromise effectiveness of shock therapy. While no energy is lost; defibrillation thresholds may be higher (and safety margin may be reduced) because the shocking vector has changed. | An Important Medical Device Information letter, dated 10/09/2006 was sent to physicians following patients with a potentially affected device. The letter describes the issue, provides for a projected rate of occurrence, and provides recommendations. An October 9, 2006 Product Advisory Letter, beginning October 30, 2006, was also sent to provide an update that further laboratory testing and analysis indicates that the nominal probability of malfunction is higher than their earlier prediction. Non-implanted devices have been retrieved. .
Z-0035-2007 | Oct-12-2006 | 2006 | Haemonetics - SmartSuction Harmony Powered Suction DeviceModel Number: HAR-E-115-US | Haemonetics Corporation | 2 | Faulty circuit board may short circuit and cause electric shock to operator | Haemonetics issued recall leters on July 28, 2006 via Certified Mail, Return Receipt. Users are requested to return unit for repair or replacement.
Z-0517-06 | Feb-16-2006 | 2006 | Ultraview SL Capnography Module; Model 91517 - Ultraview SL Capnography Module; Model 91517A mainstream and sidestream gas analyzer designed to measure the concentration of carbon dioxide in a gas mixture; to aid in determining the patient''s ventilatory; circulatory; and metabolic status. | Spacelabs Medical Incorporated | 2 | The transducer used to compensate for pressure changes was mounted backwards on the printed circuit board assembly. The module will not compensate for pressure changes and CO2 readings will be inaccurate at higher altitudes. | On 7/20/05 Spacelabs called the consignees that received the units to notify them of the hazard and also the schedule a service call to repair the affected capnography modules. The firm will confirm that all consignees have been updated with the appropriate PCBA and will audit the device recall records for completion.
Z-0152-2007 | Nov-04-2006 | 2006 | Echelon Magnetic Resonance Imaging Device - Hitachi Echelon Magnetic Resonance Imaging System; Serial Number: V-001; equipped with CTL coil type: MR-CTL-150. | Hitachi Medical Systems America Inc | 2 | Component defect/overheating device; A malfunctioning decoupling circuit in the CTL coil caused high current flow and circuit overheating. | The recalling firm issued a Device Correction letter to the clinic customer on 9/15/2006 making them aware of this potential problem and of their completed corrective action on 8/14/2006, by a Service Rep visit.
Z-1195-06 | Jul-08-2006 | 2006 | Welch Allyn PIC 50 - Welch Allyn PIC 50 Portable Intensive Care System Multi-Parameter Monitor/Defibrillator; a portable; 12V internal battery powered defibrillator; Manufactured by MRL Inc.; A Welch Allyn Company; 1000 Asbury Drive; Buffalo Grove; Illinois 60089 | MRL; Inc.; A Welch Allyn Company | 1 | The Welch Allyn PIC 50 Defibrillators may display a 'Defib Comm' or 'Pace Comm' error message on the device display during use which may result in a terminal failure of the device to analyze the patient's ECG and deliver the appropriate therapy. | MRL sent Urgent Medical Device Recall letters dated 6/30/06 via certified mail to the direct accounts informing them of an intermittent connection between a pin and socket on a circuit board inside of the device that could result in a malfunction that prevents or unacceptably delays the delivery of a defibrillating shock, which possibly could result in failure to resuscitate the patient. The accounts were provided a list of the affected units' parts/serial numbers and requested to inspect their PIC 50 units for the affected serial numbers and return the identified units for servicing. The accounts were also instructed to respond to the notification within 5 working days if the unit has displayed the Defib Comm error, and within 30 days if it has not displayed the Defib Comm error by fax or mail. A loaner PIC 50 will then be sent to the consignee with instructions on how to return their unit for correction.
Z-1106-06 | Jun-22-2006 | 2006 | Welch Allyn AED 20 Defibrillator - Welch Allyn AED 20 Automated External Defibrillator; a portable; 12V internal battery powered defibrillator; Manufactured by MRL Inc.; A Welch Allyn Company; 1000 Asbury Drive; Buffalo Grove; Illinois 60089; Model AED20; part 972200 | MRL; Inc.; A Welch Allyn Company | 1 | The Welch Allyn AED 20 Defibrillators may display a 'Defib Comm' error message on the device display during use which may result in a terminal failure of the device to analyze the patient's ECG and deliver the appropriate therapy. | MRL issued a press release on 6/15/06, and began sending Urgent Medical Device Recall letters on the same date to the direct accounts informing them of the possibility of the defibrillator going into a terminal failure due to a U25 IC socket susceptible to defibrillation circuit board communication errors. The accounts were asked to respond to the notification witihin 5 working days by faxing or mailing the enclosed response form, or calling Welch Allyn-MRL Technical Support at 1-800-462-0777. A loaner AED 20 will then be sent to the consignee with instrucitons on how to return their unit for correction.
Z-0183-2007 | Nov-16-2006 | 2006 | Philips - M3001A Philips Multi Measurement Server (MMS) | Philips Medical Systems | 2 | Unexpected pulse oximetry (Sp02) readings (100%) over time when a sensor is not attached to the patient. In addition; with a sensor attached to a patient; an incorrect high Sp02 may be displayed when a pulse rate is in a range of 185BPM | Philips issued a Customer Letter and Urgent Device Correction letter to consignees on 09/21/06. Users were notified of the problem and provided with Instructions to Mitigate Risk. A software upgrade to: version D.00.61 and Sp02 FW A.01.46 is mandatory.
Z-0184-2007 | Nov-16-2006 | 2006 | Philips - M1020B PulseOximetry Module Philips FAST Sp02 and Pulse Oximetry Nellcor OxiMax Compatible used with the IntelliVue or M3/M4 patient monitor M1002 | Philips Medical Systems | 2 | Unexpected pulse oximetry (Sp02) readings (100%) over time when sensor is not attached to patient. In addition; with a sensor attached to a patient; an incorrect high Sp02 may be displayed when a pulse rate is in a range of 185BPM | Philips issued a Customer Letter and Urgent Device Correction letter to consignees on 09/21/06. Users were notified of the problem and provided with Instructions to Mitigate Risk. A software upgrade to: version D.00.61 and Sp02 FW A.01.46 is mandatory.
Z-0258-2007 | Dec-15-2006 | 2006 | Medrad 9500 Interface Cables - Invasive Blood Pressure Interface Cables: a)Medrad 9500 Multigas monitor for MRI; Invasive Blood Pressure Interface Cable (Part # 3006314) for use with Abbott/Hospira IBP Transducer b) Invasive Blood Pressure Interface Cable (Part # 3008772) for use with Braun IBP Transducers. The product is an accessory to patient physiological monitor. | Medrad Inc | 2 | Interference problems: certain invasive blood pressure transducers connected to these cables are susceptible to interference if exposed to static magnetic field strengths exceeding 300 gauss. In certain orientations; the transducer causes the invasive blood pressure readings on the physiological monitor to display inaccurate readings. | The recalling firm contacted all affected customers via telephone or letter dated November 20, 2006 to inform them of the problem and instruct them to return the cables or apply corrective labeling. .
Z-0860-06 | May-11-2006 | 2006 | ARCHITECT B12 Reagent - ARCHITECT B12 Reagent; list 06C09-22 (4 x 100 tests); 06C09-27 (1 x 100 tests); an in vitro diagnostic kit consisting of microparticles; conjugate; assay diluent; pre-treatment 1; pre-treatment 2; and pre-treatment 3; Abbott Laboratories; Abbott Park; IL 60064 | Abbott Laboratories MPG | 3 | The 30-day onboard storage information is not included in the barcode labels for these two lots. As a result; the ARCHITECT system software is unable to track how long a reagent kit has been stored onboard the ARCHITECT instrument. | Abbott affiliates in Canada, Germany and Australia were e-mailed copies of the recall letter on 3/8/06 for translation where needed and dissemination to their accounts receiving the suspect lots. The end users were advised that the 30-day onboard storage information is not included in the barcode for the two lots. As a result the ARCHITECT system software is unable to track how long the reagent kit has been stored onboard. The accounts were instructed to manually track the onboard storage time for the two lots, and discard the kits after 30 days of onboard storage. .
Z-0002-2007 | Oct-12-2006 | 2006 | Guidant Zoom Latitude Programming System - Guidant Zoom Latitude Programming System; Model 3120; Programmer/Recorder/Monitor (PRM). A portable cardiac rhythm management system designed to be used with Guidant implantable pulse generators. Guidant Corporation; 4100 Hamline Avenue North; St. Paul; MN 55112-5798 USA. | Guidant Corporation | 3 | Final software load did not occur prior to shipment of select programmers. | Sales Representatives were asked to retrieve/return identified monitors. .
Z-0389-06 | Jan-12-2006 | 2006 | Hospira - Hospira brand Plum A+I.V. Infusion Pump; List No. 11971-04;Hospira brand Plum A+I.V. Infusion Pump; List No. 12391-04;Hospira brand Plum A+3 I.V. Infusion Pump; List No. 12618-04; Hospira brand Plum A+3 I.V. Infusion Pump; List No. 12348-04;Infusion pump is a multi-function fluid delivery infusion system; consisting of pumping module and assortment of IV sets.Product is manufactured and distributed by Hospira Inc.; 755 Jarvis Dr.; Morgan Hill; CA | Hospira Inc | 2 | The pump may shut-off prematurely during battery operation without providing adequate response time when 'E320' service warning message appears. | On 5/18/05, the firm issued letters via Federal Express to all its direct consignees, informing them of the affected products and providing instructions on the recall.
Z-1306-06 | Aug-03-2006 | 2006 | FreeStyle - Abbott brand FreeStyle Blood Glucose Meter; distributed by Abbott Diabetes Care; 1420 Harbor Bay Pwky.Alameda; CA 94502.Manufactured by: Flextronics International; Long Jing Industrial Estate; Long Jing Road; Xi Li Town; Nantou District; Shenzhen; China | Abbott Diabetes Care; Inc. | 2 | The products may encounter display problem; 'Er 4' message; during prolonged use when the low battery symbol is displayed. The situation can render the meter either inoperable or operable with invalid user configuration data including selectable unit of measure; and strip calibration code. | On 5/22/06, the firm issued letters via Federal Express to all its direct consignees, informing them of the affected products and providing instructions for the corrective actions.
Z-0322-2007 | Dec-30-2006 | 2006 | Terumo APS 1 - Terumo Advanced Perfusion System 1; 100V -120V; 15 A (circuit breaker); 50/60 Hz (20A power source required) base; Model 801763. | Terumo Cardiovascular Systems Corp | 2 | Device module(s) may fail to properly initialize during start up; thus resulting in that single module being inoperable and a red 'X' being displayed on the monitor screen. | U.S. consignees were sent an addendum to the operators manual on October 12, 2004 instructing them to cycle power to the unit to clear the error. International distributors were requested to provide this information to their customers on 12/15/04, but Terumo has not yet received confirmation that this was completed.
Z-1250-06 | Jul-22-2006 | 2006 | Hamilton - Hamilton brand RAPHAEL Silver Ventilator (Software Version 2.2xS) Product is distributed by Hamilton Medical; Inc.; 4990 Energy Way;Reno; NV 89502 | Hamilton Medical Inc | 2 | ALARM FAILURE-Following an oxygen cell calibration; the user may inadvertently and unknowingly disable the alarm system. | On May 24, 2006 the manufacturer initiated this recall. The Recalling establishment/US Initial Importer/Distributor-received the manufacturer''s notice on May 31, 2006 and initiated their recall on June 1, 2006. The US distributor has issued a Medical Device Field Correction.
Z-0304-2007 | Dec-19-2006 | 2006 | Varian - 4D Integrated Treatment Console and Varis 1.4g Medical Charged Particle Radiation Therapy System;All 4D Integrated treatment console versions 8.0.15 and below (excluding v7.0.31) and all Varis Non-Varian Treatment versions 6.6.5027 and below; with Serial numbers H510946; H510651; H510190. | Varian Medical Systems Inc | 2 | Selecting an empty space next to the ''Override'' or ''Acquire Actual'' buttons may result in mistreatment to the patient. | On September 21, 2006, the firm issued letters via certified mailing with return receipt to all its direct consignees, informing them of the affected product and providing instructions on the recall.
Z-0352-06 | Jan-06-2006 | 2006 | - ADC Solo & Initia CR 1000; Path Speed CR SP1001; Centricity CR SP1001 and Centricity CR SP 100; image digitizer; components in Agfa''s computed radiography product line. | AGFA Corp. | 2 | Users of affected systems may experience a loss of images; the need to repeat the imaging procedure and possibly a short delay in diagnosis. | Consignees were notified by letter on/about December 1, 2005. .
Z-0042-2007 | Oct-14-2006 | 2006 | GE OEC 9900 Elite C-arm; Fluoroscopic system - OEC 9900 Elite Fluoroscopy System; Catalog No. 887208 and 887210; GE OEC Medical Systems; Inc. | GE OEC Medical Systems; Inc | 2 | X-ray system may lose or mix images and/or lose patient data. | Consignees will be notified by letter beginning 10/21/2006.
Z-1259-06 | Jul-26-2006 | 2006 | Beckman Coulter Cytomics FC 500 Flow Cytometry System - Cytomics FC 500 Flow Cytometry System Part Numbers 626553 (TN; FC500 with UPS); 6605627 (TN; 5 CLR; FC500 (100V)); 6605628 (TN; 5 CLR; FC500 (120V)) ; 6605629 (TN; 5 CLR; FC500 (220V)); 6605630 (TN; 5 CLR; FC500 (240V)) with CXP Software Versions 2.0 and 2.1 | Beckman Coulter Inc | 2 | Beckman Coulter has determined that under certain conditions a system error generated during a Pause and Rotate routine may result in a sample misidentification with a risk of reporting erroneous results. | A Product Corrective Action letter was sent via US mail the week of May 22, 2006 to all FC500 customers. They were informed that under certain conditions a system error generated during a Pause and Rotate routine may result in a sample misidentification with a risk of reporting erroneous results. The letter includes the sequence in which the issue occurrs: 1) During the acquistion of a tube, the Pause and Rotate Button is selected. 2) A system error is generated during the Pause and Rotate routine (The error message table may be viewed by going to the CXP Help Menu, select the CXP Help Menu Item, then to Section 13.5). 3) An error message appears in the Cytometer Status, an audible alarm sounds, the run aborts, the carousel ejects, and the system automatically transitions from Verification to Run Initialization and then to Awaiting Sample 4) The Worklist is continued or a new Worklist is loaded and started. 5) After the Worklist is continued, or loaded and started, during the acquistion of a tube, the Pause and Rotate Button is selected. 6) The carousel Pauses but does not Rotate. 7) When the Start/Continue or Restart button is selected, the carousel may move to the wrong tube and aspirate sample resulting in a sample misidentification. If no system errors are generated during a Pause and Rotate routine, the system will function correctly. The letter also provides steps for immediate user preventive action if a system error is generated during the Pause and Rotate Routine to remedy the issue: 1) Exit CXP Software. 2) Power off the Cytometer using the FC OFF utility 3) Turn the Cytometer back on. 4) Recovery should be complete and the laboratory may resume normal usage of the system. Customers are asked to share the information with their laboratory staff and to retain this notification as part of their laboratory Quality System documentation and to complete and return the enclosed Fax Response Form for the firm''s records.
Z-0438-06 | Jan-25-2006 | 2006 | Arcadis Orbic - Arcadis Orbic Mobile X-Ray System | Siemens Medical Solutions USA; Inc | 2 | patient information can be stored in another patient's file. | The recalling firm has issued a Customer Safety Advisory to the affected customers per Update Instructions SP030/05/S. The letter informs customers of the potential problem and provides a suggestion to avoid this issue.
Z-0857-06 | May-09-2006 | 2006 | S/5 Aneshesia Delivery Unit - Datex-Ohmeda S/5 Anesthesia Delivery Unit (ADU) | GE Healthcare | 2 | Alarm sound-It is possible that the audible alarms from the speaker from the Datex-Ohmeda S/5 Anesthesia Delivery Unit (ADU) machine may dampen or disappear as a result of normal handling of the anesthesia machine. Visual alarms on the ADU are unaffected by this condition. The dampening or absence of audible alarms on the ADU may result in a delay in treatment. | Customers were sent a letter dated November 16, 2005, informing them of the possible alarm failure and giving them short term and long term solutions. Short term, clinicians are to: 1) check the audio alarm prior to use; and 2) use the ADU in conjunction with functioning anesthetic agent monitor. In the long term a GE Healthcare representative will contact the facility and schedule a visit to implement the corrective action at no charge.
Z-0581-06 | Feb-28-2006 | 2006 | Allura Xper FD10/10 - Allura Xper FD10/10 x-ray system | Philips Medical Systems North America Co. Phillips | 2 | Potential lock-up of system requiring reboot of system. | On 2/6/06 the firm sent a Product Safety Notification letter to the customers that have devices installed. The letter advises of the problem, and to restart the device if it locks up. The letter also states that the firm will be performing upgrades of the devices.
Z-0041-2007 | Oct-25-2006 | 2006 | GE OEC 9900 Elite Fluoroscopic X-ray system. - GE OEC 9900 Elite Fluoroscopy System with Integrated Navigation;GE OEC Medical Systems Inc. | GE OEC Medical Systems; Inc | 2 | Missing; mixed or lost patient images may result after X-ray procedures. Error in navigation accuracy. | Consignees will be notified by letter beginning 10/16/2006.
Z-0565-06 | Mar-03-2006 | 2006 | AUDICOR 200 D - AUDICOR 200 D Inovise Part Number 30047 (console part number 20158). | Inovise Medical; Inc. | 2 | Use of the AUDICOR 200 D in-line with a defibrillator/monitor resulted in interruption of pacing therapy to a patient when the AUDICOR 200 D stopped working due to low battery power. | On January 12, 2006, 2 of 3 customers were notified by telephone, with follow-up email. All three customers were visited by the firm on January 16-17, 2006. Customers were instucted to not use the device in-line with a defibrillator/monitor.
Z-1445-06 | Aug-26-2006 | 2006 | Delphi IVantage - Delphi IVantage Volumetric Ambulatory Infusion Pump/System; Ref nos. 12864051FRE; Rev. 8 and 12864051DEL; Rev. 8. | Delphi Medical Systems | 2 | Potential for under-infusion without alarm; cassette rollers stop moving ; but the pump shaft continues rotating without alarm. | Distributors were notified via letter dated 3/29/05 instructing them to retrieve and return the product.
Z-0033-2007 | Oct-12-2006 | 2006 | IMPAX¿ - IMPAX¿ 5.2 Systems with CAD Capability (Computer Assisted Diagnosis) | AGFA Corp. | 2 | Failed CAD displayed as 'No Findings' | Consignees were notified by letter on/about 08/08/2006.
Z-0808-06 | May-02-2006 | 2006 | PLV Continuum Ventilator - Ventilator-PLV Continuum (PLVC I); Model Number: P1000. The PLVC I is a microprocessor controlled; compressor-based; mechanical ventilator. It is intended to control or assist breathing by delivering room air to the patient. It utilizes an internal compressor to generate compressed air for delivery to the patient. Breath delivery is controlled by software algorithms. The PL VC I user interface has a membrane keypad with indicator Light Emitting Diodes for the selection and acceptance of patient settings and for the display of alarm conditions. PL VC I provides the following types of ventilatory support:1) Positive Pressure Ventilation; delivered either invasively (via endotracheal or tracheostomy tube) or non-invasively (via mask or mouthpiece ).2) Assist/Control; Spontaneous Intermittent Mandatory Ventilation (SIMV)or Continuous Positive Airway Pressure (CPAP) modes of ventilation.3) Volume-Controlled (VC). Available in AlC and SIMV.4) Pressure-Controlled (PC). Available in AlC and SIMV.5) Pressure Support (PS). Available in SIMV and SPONT. | Respironics California Inc | 1 | Flow valve failures- When failure occurs; air flow from the ventilator ceases resulting in the PLVC I transitioning into a 'Vent Inop' mode whereby the safety valve is activated; opening the circuit to ambient air. Ventilator dependent patients may not receive adequate ventilatory support if this problem occurs. | Notification letters were mailed to customers on March 20, 2006 notifying them of this action. Customers were instructed to quarantine all PLV Continuum Ventilators in their possession and not to place patients on them. Respironics Customers Service will contact customers to arrange for the return of all PLV Continuum Ventilators. Customers were instructed to safely plan and transition patients in their care from the PLV Continum Ventilator onto to other comparable patient support devices. If customers did not have a suitable ventilator to use for their patients, Respironics will loan them a PLV 100, 102 or 102b Ventilator with no charge. If patient support goes beyond the capabilities of the PLV 100, 102, or 102b Ventilator is required, Respironics will make arrangements to loan the customer with a suitable substitute rental ventilator at no charge. The field action will involve the upgrade of all affected devices with a new flow valve assembly. All US and International units are being recalled to the factory in Carlsbad, California. The rework will be conducted on-site.
Z-1172-06 | Jun-27-2006 | 2006 | Voluson Pro/Expert - Voluson 730 Expert; Voluson 730 Pro; Voluson 730 ProV with software versions 4.0.0; 4.0.1; 4.0.2; 4.0.3; 4.0.4; 4.0.5; 5.00; 5.0.1. The device is a general-purpose diagnostic ultrasound system. | General Electric Med. Sys. Ultrasound | 3 | In the cardiac measurement section of the device the calculation of the 'PG Mean' (mean pressure gradient) leads to an erroneous result. It is recommended that customers with affected units do not use the measurement in question 'PG Mean' for any diagnostic or therapeutic decision; and 'PG Mean' values should not be included in any patient documentation. | An Urgent Medical Device Correction letter to customers 04/17/2006, states the product affected, provides a recommendation for affected units continued use until a software update is provided to the facility by a GE Healthcare Representative. Customers are recommended to not use the measurement in question 'PG Mean' for any diagnostic or therapeutic decision, and 'PG mean' values should not be included in any patient documentation. .
Z-0900-06 | May-25-2006 | 2006 | SYNCHRON LX 20 Clinical System - SYNCHRON LX 20; SYNCHRON LX 20 PROClinical Systems; Operating Software Version 4.5; Part Number: A23724 In vitro diagnostic. | Beckman Coulter Inc | 3 | It was confirmed that the ORDAC (Over Range Detection And Correction) feature (auto & manual) does not perform correctly for the glucose cup chemistry (GLUCm) when sample type of 'Other' is selected w/operating software version 4.5. Results for these samples could be elevated as much as twice the actual value. The ORDAC feature does function properly w/serum; plasma; urine & CSF. | User Notification: A Product Corrective Action (PCA) letter will be sent the week of MAR 15, 2006 to all Synchron LX20, LX20PRO, LXi 725, and UniCel DxC 600/800 customers to inform them of the limitation for the use of fluid type 'Other' when programmed in combination with automatic or manual ORDAC for a GLUCm test. Customers are requested to implement one of three workarounds described in the PCA. IN PROCESS Stock Control: 1) A STOP SHIP (BRE-06016-1) was placed on all stock in inventory. DONE 2) Current customers of LX and DxC will be informed of the limitation for the use of fluid type 'Other' when programmed in combination with automatic or manual ORDAC for a GLUCm. IN PROCESS 3) New customers will receive reworked new instrument shipments with documentation regarding the use of fluid type 'Other' when programmed in combination with automatic or manual ORDAC for GLUCm. IN PROCESS 4) A Technical Update is being created to notify worldwide services of the PCA. IN PROCESS *Submit initial corrections and removal report. DONE *Under Investigation. IN PROCESS *Failure analysis is ongoing and preventative actions will be determined based on the investigation results. IN PROCESS *In Process. A 75% effectivity is targeted based the class II level of action required per the Beckman Coulter Product Corrective Action Subcommittee review. Each user will be contacted initially and then once more (if needed) in order to determine the rate of response from the Response form provided. This check enables us to ensure that users have received the PCA letter. Records will be maintained of each official contact, and may be made available upon request.
Z-0398-06 | Jan-18-2006 | 2006 | ARCHITECT c8000 System - ARCHITECT c8000 System Software; Version 2.11; Product List/Model #5F48-12; distributed by Abbot Laboratories; Irving; TX. | Abbott Laboratories; Inc | 3 | System software assigns a calibrator default volume of 2.0uL when field is left empty by operator at time assay parameters set. Patient results could be affected if the volume required is not 2.0 uL. | Firm initiated recall to all consignees via letter dated 12/08/05. .
Z-0576-06 | Feb-28-2006 | 2006 | - Stat Profile Critical Care (CCX) Analyzer Catalog Number: 35942 | Nova Biomedical Corporation | 3 | Analyzers may fail to perform scheduled automatic two-point calibrations and Quality Control runs | Nova Biomedical Issued on 1/27/2006 an Advisory Notice to Stat Profile CCX Customers to report the malfunction and request the Good Laboratory Practices (GLP) be used and review be made to assure the calibration status and QC is being performed. A new software version is scheduled to be released to correct the error.
Z-0902-06 | May-25-2006 | 2006 | UniCel DxC 600/800 SYNCHRON Clinical Systems - Unicel DxC 600/800 SYNCHRON Clinical Systems; Operating Software Version 1.0 and 1.2; Part Number: A20463 and A27331 | Beckman Coulter Inc | 3 | It was confirmed that the ORDAC (Over Range Detection And Correction) feature (auto & manual) does not perform correctly for the glucose cup chemistry (GLUCm) when sample type of 'Other' is selected w/operating software versions 1.0 & 1.2. Results for these samples could be elevated as much as twice the actual value. The ORDAC feature does function properly w/serum; plasma; urine & CSF. | User Notification: A Product Corrective Action (PCA) letter will be sent the week of MAR 15, 2006 to all Synchron LX20, LX20PRO, LXi 725, and UniCel DxC 600/800 customers to inform them of the limitation for the use of fluid type 'Other' when programmed in combination with automatic or manual ORDAC for a GLUCm test. Customers are requested to implement one of three workarounds described in the PCA. IN PROCESS Stock Control: 1) A STOP SHIP (BRE-06016-1) was placed on all stock in inventory. DONE 2) Current customers of LX and DxC will be informed of the limitation for the use of fluid type 'Other' when programmed in combination with automatic or manual ORDAC for a GLUCm. IN PROCESS 3) New customers will receive reworked new instrument shipments with documentation regarding the use of fluid type 'Other' when programmed in combination with automatic or manual ORDAC for GLUCm. IN PROCESS 4) A Technical Update is being created to notify worldwide services of the PCA. IN PROCESS *Submit initial corrections and removal report. DONE *Under Investigation. IN PROCESS *Failure analysis is ongoing and preventative actions will be determined based on the investigation results. IN PROCESS *In Process. A 75% effectivity is targeted based the class II level of action required per the Beckman Coulter Product Corrective Action Subcommittee review. Each user will be contacted initially and then once more (if needed) in order to determine the rate of response from the Response form provided. This check enables us to ensure that users have received the PCA letter. Records will be maintained of each official contact, and may be made available upon request.
Z-1286-06 | Jul-27-2006 | 2006 | APS 1 - Terumo Advanced Perfusion System 1; 100V -120V; 15 A (circuit breaker); 50/60 Hz (20A power source required); Models 801763 and 801763 (Japan unit). | Terumo Cardiovascular Systems Corp | 3 | While on battery power; the system may flash a low battery warning even though the battery is adequately charged. | Customers were notified via letter dated October 19, 2005 that the Power Manager Software would be upgraded within the next six months. As of 7/7/06 all units within the U.S. have been completed and international units are still in the process of being upgraded.
Z-1235-06 | Jul-13-2006 | 2006 | KinetDx - KinetDx 4.0 Ultrasound Image Management System. | Siemens Medical Solutions; USA; Inc | 3 | The cardiologist's report comments may not be retained by the system due to a software bug. | In April 2006, Siemens'' field force was issued instructions to visit each customer location and to complete the software upgrades within six months.
Z-0911-06 | May-31-2006 | 2006 | Medtronic Paceart software - Medtronic Paceart 2004 Second Edition & Paceart 2005 First Edition versions of the Paceart Gateway Software. The Medtronic Paceart System is a personal computer based pacemaker-testing system; a 12-lead electrocardiograph; and a transtelephonic receiving station. It provides data that can be used to analyze implanted pacemaker performance based on electrocardiographic measurements; either taken directly from the patient or programmer. The system can measure; store; and display any of the 12 standard leads. Reports and charts are available by means of a laser printer. The system also includes a database that collects and stores patient data. Optional software allows the device to function as an unattended transtelephonic receiving station. | Medtronic Inc. Cardiac Rhythm Managment | 3 | The Paceart System could inappropriately insert data from a patient's Medtronic CareLink transmission into another patient's Paceart record. An issue associated with the automated batch transfer of implanted cardiac device data from the Medtronic CareLink network to the Paceart System exists. | Affected user and potential users of the Gateway product were notified via phone and/ or a customer notification. The notification offers interim options until a software update is available. Customers are asked to respond back to Metronic with an enclosed Acknowledgement form.
Z-0311-2007 | Dec-27-2006 | 2006 | Terumo APS 1 - Terumo Advanced Perfusion System 1; 100V -120V; 15 A (circuit breaker); 50/60 Hz (20A power source required) base; Model 801763. | Terumo Cardiovascular Systems Corp | 3 | The battery status light on the front panel may not accurately reflect the actual status of the battery. The battery may contain more charge than indicated. | U.S. consignees were visited and new software installed between May 2004 and November 2005. Currently, approximately 50% of exported units have been upgraded.
Z-0905-06 | May-25-2006 | 2006 | Medtronic Restore Neurostimulator - Medtronic Model 37711 Restore Neurostimulator. The Restore Neurostimulator is an implanted; rechargeable neurostimulator which provides pain management therapy. | Medtronic Inc; Neurological & Spinal Division | 3 | A limited number of Model 37711 Restore Neurostimulators have an incorrect internal memory parameter that causes the device to indicate it is fully recharged when it may only be 80% to 90% recharged. This anomaly does not affect patient safety. This anomoly can be corrected by reprogramming the Restore Neurostimulator using specific software. | A 'Customer Communication' letter was sent to Medtronic field personnel to hand-carry into the affected accounts; customer contact began on 11/22/2005. The letter explains the patients implanted with these devices likely will only be able to recharge their neurostimulator to about 80-90% of a full charge. This anomaly does not affect device safety or effectiveness but Medtronic will re-program them upon request for customer satisfaction. The letter will be followed by scheduled on-site visit to reset the internal memory parameter of the neurostimulator.
Z-1325-06 | Aug-03-2006 | 2006 | ABL800 blood gas analyzer - ABL800 Series Blood Gas Analyzer equipped with Software Version 5.21. | Radiometer America Inc | 3 | ABL800 Series Blood Gas Analyzer became inoperable when the 'restore default setup' button was touched. | The firm initiated this recall on 10/31/2005, at which time a Field Action Note was issued to their Service Reps. The notice instructed the service personnel to upgrade all ABL800 analyzer units presently running with software version V5.21 to the new version V5.23, in order to correct this problem. The Service Personnel began conducting visits at each customer site on 10/31/2005.
Z-0853-06 | May-06-2006 | 2006 | - Affymetrix GeneChip Microarray Instrumentation System; consisting of GeneChip 3000Dx scanner with autoloader; FS450Dx fluidics station and GCOSDX Software | Affymetrix; Inc. | 3 | Incorrect software version of instrument controller is not compatible with new configuration. This incorrect configuration may cause the instrument system to fail at start-up or during a run. | Afffymetrix notified the two hospitals by email dated 4/14/05 and by letter to the Distributor on 4/28/05. Affymetrix service personnel will be provided with discs to perform the software update. All field changes will be documented via Service Work Orders. .
Z-1256-06 | Jul-27-2006 | 2006 | IMPAX Diagnostic Workstation - IMPAX 4.5 Systems; used in the acceptance; transfer; display; storage and digital processing of medical images. | AGFA Corp. | 3 | Possible corrupted image appearing after System Start. | On 06/30/2006 Agfa upgraded the Software on the installed units using a Mandatory Service Bulletin DD+DIS225.06E.
Z-0364-06+A25 | Jan-07-2006 | 2006 | LANTIS software - LANTIS software with MedOnc Module;MedOnc module is an option of the LANTIS radiotherapy treatment planning software distributed as a component of the firm''s linear accelerator used to deliver radiation therapy to cancer patients.The MedOnc module is distributed by Siemens Medical Solutions USA ¿ Oncology Division; 4040 Nelson Avenue; Concord; CA 94520 with their linear accelerators software package called LANTIS. The MedOnc module option is distributed with all LANTIS software copies; but the module must be turned on by an electronic key that is purchased by individual hospitals.Siemens Medical Solutions USA purchases the LANTIS software (code) from IMPAC Medical Systems; Inc.; 100 West Evelyn Avenue; Mountain View; CA 94041 (REG#2950347; FEI: 1000123805). IMPAC distributes the same software under the name SEQUENCER. | Siemens Medical Solutions USA; Inc | 2 | When using the Creatine Clearance dose calculations a software bug will allow the algorithm to assume a patients age to be in years when it is entered in months. This miscalculation may result in patients less than two years old to be seriously overdosed with radiation. | The firm has sent out notification letters to the 18 sites with the activated MedOnc option. All 18 site had the letters delivered by Siemens Medical Solutions field service personnel by 1/6/2005. .
Z-0417-06 | Jan-24-2006 | 2006 | Accu-Chek Aviva - Roche brand Accu-Chek Aviva Care Kit; Catalog number 04528280001. (Distributed within the U.S.) | Roche Diagnostics Corp. | 2 | Incorrect patient blood glucose test results may be reported due to a software problem. | Notification letters dated 11/9/05 were issued to Users, Mail Order Distributors, Clinical Trial Sponsors, Pharmacists and Health Care Professionals. Press release issued 1/17/06. .
Z-0682-06 | Mar-29-2006 | 2006 | RapidLab1200 System - Brand Name: RapidLab 1200 System; Model 1240; Part Number 05060298; All RapidLab 1200 Systems with Software Versions 1.0; 1.01; or 1.02. --- Common Name: Blood gases; electrolyte and blood pH test system. | Bayer Healthcare; LLC (Diagnostics Division) | 2 | When stored data is recalled from memory; the results do not agree with the original results. Software anomaly: a calculation error occurs when using the correlation adjustment feature on the RapidLab 1200 Systems with software versions 1.0; 1.01; or 1.02. The correlation feature applies user-defined slope and offset parameters each time stored test results are recalled to the display. | Support Bulletins and Customer Bulletins were e-mailed to all Bayer HealthCare LLC Branches on 1/20/06 for distribution to the end user facilities (medical facilities).
Z-0298-2007 | Dec-19-2006 | 2006 | Identity - St. Jude Medical Identity SR Model 5172; Identity DR Model 5370; & Identity XL DR Model 5376; Pulse Generators (pacemakers) | St Jude Medical CRMD | 2 | St. Jude Medical has identified a low-frequency anomaly in the software used in the APS III Model 3500/3510 and Merlin PCS Model 3650 programmers that can lead to incorrect reporting of battery voltage; expected battery longevity and Elective Replacement Indicator (ERI) status. | The notification material was sent to physicians by FedEx. The advisory notification was sent by FedEx on Oct. 12, 2006, to arrive at the physicians' offices on Oct. 13th. In the US, acknowledgement of receipt of the notification material will be monitored by St. Jude Medical CRMD in Sylmar, California. For the rest of the world it will be by various St. Jude offices and local distributors with this information being fed back to St. Jude Medical CRMD in Sylmar. In addition, St. Jude Medical issued a press release on October 13, 2006. Plus information concerning this advisory can be found on St. Jude's website, www.sim.com Completion of this advisory is estimated to be three months after its initiation.
Z-1355-06 | Aug-05-2006 | 2006 | ACCU-CHEK Advantage - ACCU-CHEK Advantage Complete Diabetes Monitoring Kit; Cat. No. 860. Ref. no. 080086000 with meter model number 870; Ref. no. 03002531001 with meter model number 33304510; Ref. no. 03002531001 with meter model number 2138930; Ref. no. 03002531001 with meter model number 2138930. | Roche Diagnostics Corp. | 2 | The meter gives a strip error message that can "actually" mean either a problem with the strip or a blood glucose too low to measure; but the meter error message only reports that there is a strip error and the manual only instructs the user to retest with a new strip using a larger drop of blood. | A recall notification letter dated April 24, 2006 was sent to each customer alerting them to the fact that this error code has a dual meaning. New software will be installed in all meters when it becomes available.
Z-1356-06 | Aug-05-2006 | 2006 | ACCU-CHEK Advantage - ACCU-CHEK Advantage Blood Glucose Meter; Cat. No. 2137500; Ref. no. 0385633001 with meter model no. 768. | Roche Diagnostics Corp. | 2 | The meter gives an error message that can "actually" mean either a problem with the strip or a blood glucose too low to measure; but the meter error message only reports that there is a strip error. | A recall notification letter dated April 24, 2006 was sent to each customer alerting them to the fact that this error code has a dual meaning. New software will be installed in all meters when it becomes available.
Z-1357-06 | Aug-05-2006 | 2006 | ACCU-CHEK Advantage - ACCU-CHEK Advantage Blood Glucose Meter; Cat. No. 3332586; Ref. no. 03332586001 with meter model no. 2138930. | Roche Diagnostics Corp. | 2 | The meter gives an error message that can "actually" mean either a problem with the strip or a blood glucose too low to measure; but the meter error message only reports that the test strip may be damaged or test was not performed correctly. | A recall notification letter dated April 24, 2006 was sent to each customer alerting them to the fact that this error code has a dual meaning. New software will be installed in all meters when it becomes available.
Z-1360-06 | Aug-05-2006 | 2006 | ACCU-CHEK Advantage - ACCU-CHEK Advantage (Managed Care Kit) 100- value memory with time and date; Not for Retail Sale; Cat. No. 3000389; Ref no. 03000389001 with meter model no. 2138930. | Roche Diagnostics Corp. | 2 | The meter gives an error message that can "actually" mean either a problem with the strip or a blood glucose too low to measure; but the meter error message only reports that there is a bad strip. | A recall notification letter dated April 24, 2006 was sent to each customer alerting them to the fact that this error code has a dual meaning. New software will be installed in all meters when it becomes available.
Z-1367-06 | Aug-05-2006 | 2006 | AccuData GTS Plus - AccuData GTS Plus AccuData GTS For Blood Glucose Monitoring and Automated Data Management; contents Accu-Chek Advantage repalcement Meter and module; cat. no. 454; Ref. no. 03144798001 contains meter model no. 777. | Roche Diagnostics Corp. | 2 | The meter gives an error message that can "actually" mean either a problem with the strip or a blood glucose too low to measure; but the meter error message only reports that there is a bad strip. | A recall notification letter dated April 24, 2006 was sent to each customer alerting them to the fact that this error code has a dual meaning. New software will be installed in all meters when it becomes available.
Z-1487-06 | Sep-12-2006 | 2006 | Philips Medical - IntelliVue MultiMeasurement Server (MMS)/Multimeasurement Server; physiological patient monitoring system.Model: M3001A | Philips Medical Systems | 2 | Patient monitor may display inaccurate reading when the Disposable Sp02 Sensor is not attached | Philips is using the contract services of Sterling, Indianapolis, IN to contact all consignees regarding the correction/upgrade to the device. The letter notification dated 7/21/06 was mailed UPS Delivery Tracking the week of 07/24/06. Additionally, Philips included with the Urgent Device Correction Notice a procedure to mitigate risk until the upgrade is installed.
Z-1518-06 | Sep-23-2006 | 2006 | UniCel Dxl Access Immunoassay System - UniCel Dxl 800 Access Immunoassay Systems Wash Buffer; Part Number 8547197 Wash Buffer | Beckman Coulter Inc | 2 | The neck of the wash buffer cube can be extended higher than intended possibly causing the system to aspirate air instead of buffer before the buffer container is empty without the system immediately indicating that the supply is empty. This may lead to incorrect assay results. | A Product Corrective Action (PCA) letter was sent on July 28, 2006 to all Dxl customers to inform them that due to a recent change in packaging it is possible that the neck of the wash buffer cube can be extended higher than intended. This condition can cause the system to aspirate air instead of buffer before the buffer container is empty without the system immediately indicating that the supply is empty. If they are running System software version 2.2.2 (distributed beginning March 2006) and greater it will detect the premature out-of-buffer state, thus preventing the potential for erroneous results. The letter also informed the customers that if they run system software version 2.2.1 or earlier with these wash buffer lot numbers, the premature out-of-buffer state will not be detected immediately, leading to the potential for erroneous results. They were informed that Installation of software version 2.2.2 or greater is mandatory. To ensure immediate upgrade to version 2.2.2, customers are to contact the Clinical Support Center or their local Beckman Coulter representative. After system software version 2.2.2 or greater is installed, the system will detect the out-of-buffer state, stop sample processing, and suppress results. One of the following Event Log messages will be posted: ''Air or excessive restriction has compromised wash buffer delivery'' or ''Air detected in the wash buffer supply line'' To prevent out-of-buffer events users are to complete the enclosed procedure: Calibrating the Wash Buffet Level Sensor. Completing the enclosed procedure before using any of the affected lots of Part Number 8547197 will eliminate unnecessary workflow interruptions. Customers are to share the information with laboratory staff and retain this notification as part of their laboratory Quality System documentation. They are also instructed to complete and return the enclosed response form within 10 days so that the firm can be assured they were notified.
Z-0601-06 | Mar-09-2006 | 2006 | LifeCare PCA3 - Hospira LifeCare PCA3 Infusion System Software; versions 4.08 and 4.09 | Hospira Inc | 2 | Incorrect values for dosage; dose limit; concentration or rate of delivery may display on screen. | Field corrections initiated November 2, 2005. The firm has sent notification letters to consignees via Federal Express. .
Z-0295-2007 | Dec-19-2006 | 2006 | LIFEPAK 20 - LIFEPAK 20 automatic external defibrillator | Medtronic Emergency Response Systems; Inc. | 2 | LIFEPAK 20 may lock-up when attempting to power-up on DC within 2 seconds after removing AC power. | Letter dated December 2006 alerts users to the lock-up problem, and suggesting a procedure to prevent lock-up. Firm promised another letter in January 2007 detailing the plan to provide software upgrade.
Z-1310-06 | Aug-02-2006 | 2006 | VIGILANCE MONITOR - EDWARDS VIGILANCE MONITORS(FORMERLY BAXTER VIGILANCE MONITORS)with software release 5.3 or earlier only. Patient monitor which measures cardiac output. Models: VG; VGS; VGS1; VGS2; VG2; & IVM. | Edwards Lifesciences Llc | 2 | Edwards Lifesciences Vigilance monitors with software release 5.3 or earlier may improperly cause the monitor to deliver power to the Continuous Cardiac Output (CCO) catheter without alerting the user to this situation. This can result in overheating and thermal damage to the CCO catheter and serious patient injury. | Strategy to contact affected physicians accounts: Edwards will communicate to all accounts with affected monitors using distribution and sales records to identify those accounts. Letter dated 07/07/2006. In addition Edwards will analyze all CCO catheter sales and identify all accounts not on the monitor distribution list and include those accounts on the communication listing. Communication will be send via FedEx to the Risk Manager at each account with sufficient copies for the Heads of Departments of Surgery, Anesthesia, ICU, Nursing, Risk Management and Biomedical Engineering. Edwards will follow up by phone with each account that does not initiate contact per the Recall Letter. Edwards will make three attempts to contact the account by phone and complete the recall process.
Z-0055-2007 | Oct-20-2006 | 2006 | ApexPro Telemetry System - GE ApexPro CH Telemetry System. The ApexPro Telemetry System is composed of 6 major components. -The patient worn data acquisition transmitters; -The receiver antenna system infrastructure; -The receivers; -The receiver subsystem; -A computer platform running the Apex Pro Telemetry Application; -A computer platform running a central station application (which may be the same computer platform running the ApexPro Telemetry Application) | Ge Healthcare | 2 | When a patient being monitored with ApexPro/Apex Pro CH Telemetry with ST monitoring disabled is in a pre-existing condition of continuous MESSAGE or ADVISORY level alarm preceding a SYSTEM WARNING level alarm; the SYSTEM WARNING audible alarm and flashing yellow border around the patient panel at the CIC (Clinical Information Center) does not occur. | An Urgent Medical Device Correction letter dated August 04, 2006 was sent to customers with affected devices. The letter described the issue affecting devices containing software versions v3.8 and earlier, provides short term recommendations and a long term solution of updating the software.
Z-1502-06 | Sep-13-2006 | 2006 | LIFEPAK - LIFEPAK 20 external defibrillator/monitor | Medtronic Emergency Response Systems; Inc. | 2 | Devices with v38 system software do not display a ''LOW BATTERY: CONNECT TO AC POWER'' message when the monitor is on backup (DC) battery power and may shut down without warning. | On August 30, 2006 the firm issued a letter dated August 2006 and titled MEDICAL DEVICE CORRECTION. The letter advises of the issue, explains how to resume power if power is shut down without warning, and advsies that a service representative will contact the consignee to make arrangements to install new software. On 10/19/06 the firm mailed an identical letter dated October 2006 to customers for units that were later identified as being subject to the recall.
Z-0387-06 | Jan-11-2006 | 2006 | GE LOGIQ Book - LOCIQ Book System models: 2399921; 2399923; 2349933; 2349934 and 2349935 | General Electric Med Systems LLC | 2 | The Estimated Fetal Weight (EFW) result will be incorrect if the operator fails to perform the 'end exam' or new patient registration step before starting an EFW measurement procedure on the next patient. | Devices will have software upgrades done by GE Field Service Technicians. The upgrade will include a dialog box which will display on the screen when prior data exists and a new patient is being registered.
Z-0353-06 | Jan-06-2006 | 2006 | - CR QS (versions 3.0.207; 2.0.210 SU1; 3.0.229) ADC QS and Centricity CR Online Processing Software (version 2.x) | AGFA Corp. | 2 | Users of affected systems may experience one or more of four (4) problems the clinical impact could be a loss of image(s); the need to repeat the imaging procedure and; possibly; a short delay in diagnosis. | Consignees will be notified by letter on/about December 1, 2005 and offered upgraded software.
Z-0515-06 | Feb-14-2006 | 2006 | Babylog 8000 plus with software version 5.n - Babylog 8000 plus with software version 5.n. Neonatal Respirator-Continuous Ventilator | Draeger Medical; Inc. | 2 | Failure to generate an alarm when an unintended disconnection occurrs | Recalling firm has issued a medical device recall letter to consignees to the attention of the Director, Neonatal Intensive Care Unit via Certified Mail to alert them of additional monitoring requirements if a PEEP setting of less than 2.5 mbar is used .
Z-0378-06 | Jan-11-2006 | 2006 | Instrumentation Laboratory - GEM PREMIER 3000 | Instrumentation Laboratory Co. | 2 | Software error may report higher results for glucose and lactate | Instrumentation Laboratory notified customers by letter dated December 2005 on 12/16/2005. Customers were notifed if the error message occurs the lab should diable the glucose and lactate channels or remove cartridge. A software upgrade will be released in early 2006. .
Z-1266-06 | Jul-27-2006 | 2006 | APS 1 - Terumo Advanced Perfusion System 4 inch diameter Roller Pump; Model 801040. | Terumo Cardiovascular Systems Corp | 2 | The pump may stop during the autodose delivery without completion of the dose delivery to the patient. | Service technicians were notified on 9/30/05 to upgrade the software on these units by March 31, 2006 as follows: pump software to version 1.20, power manager software to version 1.20 and monitor software to version 2.2.1. U.S. software has all been upgraded and international upgrades are ongoing as of 7/7/06.
Z-0070-2007 | Oct-25-2006 | 2006 | syngo Imaging - syngo Imaging; picture archiving and communications system. Model number 7502029 | Siemens Medical Solutions USA; Inc | 2 | Potential issue with annotation function. Annotations; shown as overlays; may change font size when the image is saved. Text may appear shifted; or lines appear thicker; and their position may not be as intended. | The recalling firm has issued a Customer Safety Advisory via certified mail to affected customers per Update Instructions HS096/06/S and HS097/06/S. The letter, dated September 5, 2006, informs customers of this issue and provides recommendations to avoid its occurrence.
Z-1317-06 | Aug-02-2006 | 2006 | Blood Gas and Co-Ox Electrolyte & Metabolyte Analyzers - Blood Gas and Co-Ox Electrolyte & Metabolyte Analyzer; ABL-700 series. | Radiometer America Inc | 2 | Software defect. pO2 and pCO2 sample test results run on the firm's ABL700/800 Series Blood Gas Analyzers are not being properly flagged during the sample calibration phase. | The customers were apprised of this situation with a recall letter issued by the firm on 6/17/2005. In addition, new Operators Manuals were sent to the customers. In the letter, the customers were advised that a Service Engineer would be contacting them shortly in order to arrange for the installation of the software upgrade. The firm''s Service Engineers visited each customer location and installed software upgrades. Version 3.831 software was installed on the 700 Series Analyzers and Version 4.14 was installed on the 800 Series Analyzers. According to the firm, these software corrections were verified as having been completed by 3/10/2006.
Z-0910-06 | May-31-2006 | 2006 | Altair; Airis Elite; AIRIS II MRI System - Magnetic Resonance Imaging Device | Hitachi Medical Systems America Inc | 2 | The device has a software anomaly which causes the operating system to 'lock up' and data to become corrupted during service procedures. | The recalling firm plans to notify their customers of this recall/soft ware correction during a scheduled visit to be conducted at each customer site by a Hitachi Field Service Technician. During the visit, the necessary corrective software will be installed. The visits/corrections began on: 2/21/06.
Z-1245-06 | Jul-18-2006 | 2006 | Coulter LH500 Hematology Analyzer - Coulter LH500 Hematology Analyzer: Part Numbers 178832; 178833; 178834 | Beckman Coulter Inc | 2 | Beckman has confirmed that erroneous results could be reported when the workstation database crashes. | A Product Corrective Action (PCA) letter was sent week of April 17, 2006 via US mail. The letter outlines action to be taken to avoid potential erroneous results. Firm targeting 75% effectivenss checks.
Z-1440-06 | Aug-26-2006 | 2006 | UniCel DxC Synchron Clinical Systems - UniCel DxC 600/600PRO/600i/800/800PRO Synchron Clinical Systems ; Part Numbers: A20463-Software version 1.0; A27331-Software version 1.2; A29764-Software version 1.4 | Beckman Coulter Inc | 2 | Incorrect Reagent Status-When a new cartridge of Infinity Lithium Reagent is loaded on to the UniCel DxC System; the Reagent Status screen will incorrectly show 'Days Left' as 21. The correct Reagent Status for a new Lithium reagent cartridge is 14 days. Lithium reagent used past 14 days may produce low quality control and/or patient results. | A Product Correction Action (PCA) letter was sent the week of May 15, 2006 to all Unicel DxC customers informing them not to leave any Lithium reagent cartridge onboard a UniCel DxC System past 14 days. They also were instructed to complete and return the enclosed response form. The Letter was sent by US Mail.
Z-1464-06 | Aug-29-2006 | 2006 | RapidLab¿ 1245 System - RapidLab¿ 1200 Systems; Model 1245 - Blood gases; electrolyte and blood pH test system; Bayer HealthCare LLC; Diagnostics Division | Bayer Healthcare; LLC (Diagnostics Division) | 2 | Invalid CO-oximeter values; Bayer HealthCare determined that under very specific conditions; the RapidLab? 1245 or RapidLab? 1265 systems may report invalid CO-oximeter values to an LIS; Rapidlink?; or Rapidcomm&#153; data management system. | On 4/04/06, Bayer HealthCare LLC (Diagnostics Division) sent Support Bulletins to all affected Bayer HealthCare LLC Branches via e-mail in order to inform them of the issue, and to provide them with instructions for the interim work-around until the new software version is installed. Customer Bulletins were prepared and sent to the Branches both in the United States and globally for communication with the affected customers.
Z-0197-2007 | Nov-21-2006 | 2006 | Cytomics Flow Cytometry System - Cytomics FC 500 Flow Cytometry System with CXP Software Versions 2.0 & 2.1 | Beckman Coulter Inc | 2 | Mis-identification-If a panel or protocol is added to an existing worklist but the tube location is not specified; the CXP Acquisition software will run the last specified tube through the remaining protocols and will generate results and printouts of these runs leading to a Mis-Identification condition. | Customer letter -- A Product Corrective Action (PCA) letter was sent the week of July 24, 2006 to all accounts that have FC500 with CXP Software 2.0 and 2.1 to inform them that a potential sample misidentification may occur when running a Work list with missing tube Location(s) in Automatic MCL mode. The letter outlines actions to be taken to avoid mis-identification of patient results.
Z-1147-06 | Jun-21-2006 | 2006 | - LADARVision Excimer Laser System. | Alcon Laboratories; Inc. | 2 | Alcon has identified a software error associated with the use of the Measuring Mode (ruler tool) on the LADARVision4000 Excimer Laser Systems. Use of this feature after alignment of the horizontal reference line will negate compensation for cyclotorsion on CustomCornea procedures; possibly adversely affecting clinical outcomes. | A Medical Device Safety Alert dated August 1, 2005, has been issued to all consignees. Each consignee was issued a Response To Safety Alert form to acknowledge receipt of the Safety Alert. A Field Service Engineer will visit to install the revised software.
Z-0889-06 | May-19-2006 | 2006 | Extended Brilliance Workspace Imaging Workstation - Extended Brilliance Workspace Imaging Workstation. An independent diagnostic viewing and processing workstation; for CT images. | Philips Medical Systems (Cleveland) Inc | 2 | Due to a software deficiency that occurs in the combine mode; a patients image can be filmed over a different patients image. | Revised, corrected software has been shipped to consignees beginning June, 2004.
Z-0054-2007 | Oct-21-2006 | 2006 | ApexPro FH Telemetry System - ApexPro FH Telemetry System: composed of six major components (as follows); Accessories to the patient worn acquisition transceivers; the patient worn data acquisition transceivers; the transceiver access points with antenna; the network infrastructure; A computer platform running the ApexPro Telemetry Application and a computer platform running a central station application (which may be the same computer platform running the ApexPro Telemetry Application). GE Medical Systems information Technologies; 8200 West Tower Avenue; Milwaukee; WI 53223. | Ge Healthcare | 2 | System Warning Alarm failure: When a patient being monitored is in a pre-existing condition of continuous MESSAGE or ADVISORY level alarm preceding a SYSTEM WARNING level alarm; the SYSTEM WARNING audible alarm and flashing yellow border around the patient panel at the CIC does not occur. | An Urgent Medical Device Correction letter, dated June 23, 2006, was sent to customers notifying them of the issue, provided short term recommendations and provided the long term solution as a software upgrade that will correct this situation. A reply card is asked to be returned to initiate the upgrade process.
Z-0029-2007 | Oct-12-2006 | 2006 | MRI systems - MRP-7000 and AIRIS Magnetic Resonance Imaging Systems; Software Versions: V7.0A to V7.0J. | Hitachi Medical Systems America Inc | 2 | Image orientation error. When a 3D Maximum Intensity Projection (MIP) image data set is transferred from the MRI system to a computer workstation via the DICOM protocol; the anatomical markers do not change between the images as they rotate. | On 8/16/2006, Hitachi issued two (2) Device Correction letters. The first letter was addressed to the five (5) customers whose system units are currently ''de-installed''. This letter informs the customers of the problem and asks that the customer(s) please contact the firm''s Customer Service Center on their toll-free tel. # 1-800-800-4925, in order to schedule a device correction should the customer ever re-activate/install their system(s). The second letter, issued to the customers with activated MRP-7000 and/or AIRIS MRI systems, informs the customers of the problem, and states that a HMSA Service (person) will be contacting the customer in order to schedule a device correction which should address the problem.
Z-1262-06 | Jul-27-2006 | 2006 | Radiance STAT Analyzer Management System - Radiance Data Management System Software Version 2.42 (Stand-alone software package) for the Radiance STAT Analyzer Management System; in vitro diagnostic. | Radiometer America Inc | 2 | Software anomaly. Data generated by blood gas equipment did not correlate with data shown in the Radiance Data Management System. | The firm initiated this recall with letters to their customers on 3/2/2005. After this notification, the firm''s Field Service Reps visited each affected customer and updated their software to version 2.42. According to the firm, these actions took place from 3/12/2005 thru 4/15/2005. The Field Reps. of Radiometer America reported to Radiometer Medical ApS in Denmark that the necessary corrections had been competed as of: 4/15/2005.
Z-0809-06 | Apr-26-2006 | 2006 | AXIOM Artis Angiographic X-Ray System - AXIOM Artis X-ray systems running software versions VB11K and VB22G | Siemens Medical Solutions USA; Inc | 2 | The recalling firm became aware of a potential problem when imaging patients on their stomach on the AXIOM Artis X-ray Systems running software versions VB11K and VB22G. In some instances; the orientation labels could be incorrect. | The recalling firm has issued a Customer Safety Advisory to the affected customers per Update Instructions AX011/06/S and AX013/06/S, on 3/10/2006. This informs customers of the potential problem and provides a recommendation to avoid this issue.
Z-1105-06 | Jun-10-2006 | 2006 | MRP-7000 AIRIS magnetic resonance imaging devices - MRP-7000; AIRIS Magnetic Resonance Imaging Systems | Hitachi Medical Systems America Inc | 2 | Software anomaly. A software defect in the referenced systems can cause a slice indicator (reference line) to be mis-positioned on the cross-reference image when it is magnified. This image error could result in a possible patient mis-diagnosis. | On 3/17/2006, the firm began conducting a survey of all afftected customer sites in order to determine if corrected software has been installed on the unit(s) at each customer location. If the corrected software has not been installed, the firm''s Field Service Engineers will install software upgrades as necessary. .
Z-1268-06 | Jul-27-2006 | 2006 | GammaWin - GammaWin software; part number GM11019110; versions up to and including 1.62; for the GammaMedplus radionuclide applicator system; for radiation therapy. | Varian Medical Systems | 2 | Medical device software for brachytherapy may cause erroneous data to be recorded and affect patient radiation treatments. Erroneous data may be listed on the treatment history report; and the default step size may lead to a misadministration if treatment data is entered manually. | The recalling firm notified end users by letter on 05/01/06 and advised users of possible communication error between afterloader devices and software. The notification also advised that the control program for the devices has the potential to lead to a misadministration of treatment in cases where data is manually entered by the user. The firm additionally advised that user may continue to use the afterloader devices if following the listed cautionary statements. A software upgrade is planned to be implemented at the user site''s next regularly scheduled radiation source exchange. A response form is enclosed with the notification and users are requested to return the customer receipt verification via fax. .
Z-1516-06 | Sep-23-2006 | 2006 | ADVIA Centaur CP System - ADVIA Centaur CP System; Automated Immunoassay Analyzer; Part /Catalogue Number 086-A001. | Bayer Healthcare; LLC (Diagnostics Division) | 2 | Sample/Patient mis-identification (software defect)-the system can associate a test result and sample identification (SID) with an incorrect patient name when the Patient Demographics feature is used and the Patient Identification (PID) field is left blank | On 8/04/06, a Support Bulletin was e-mailed to all affected Bayer HealthCare LLC Branches worldwide in order to inform them of the issue and to provide them with instructions for this 'field correction' until the new software version is issued. A Customer Bulletin and Confirmation Form (FaxBack) were prepared and sent to the affected Bayer Healthcare LLC Branches on 8/04/06 for communication with the affected customers worldwide. .
Z-0854-06 | May-09-2006 | 2006 | Image Processing System - Advantage CT/PET Fusion for Advantage Windows Workstation with Volume View CT Colonography/Navigator 2 software. (Version 4.0; 4.1;or 4.2) | General Electric Med Systems LLC | 2 | Display in fused mode- viewer allows 2 studies ( PET/CT) registered in the same spatial domain to be viewed simultaneously. These 2 images should be visually aligned to a precision of less than 1/2 PET Voxel. A problem associated with this version of software may introduce a shift of up to 1.5 PET Voxels; which can be as much as 8 mm with certain zoom factors. | An Urgent Safety Notice letter, dated October 8, 2004,(5121309-147, rev.E) was distributed to affected customers and a validated software patch that corrected the problem was installed on affected systems via Field Modification Instruction (FMI) 80114.
Z-0380-06 | Jan-11-2006 | 2006 | Cytomics FC500 with CXP Software - Cytomics FC500 withCXP Software | Beckman Coulter Inc | 2 | Under certain conditions an incorrect Tube ID (Barcode) can be displayed and printed on the Runtime Panel Report in CXP Cytometer Software Version 2.0 | A Product Corrective Action letter will be mailed the week of Nov 28, 2005 to all FC500 wi CXP Software customers. Each user will be contacted initially and then once more (if needed) in order to determine the rate of response from the Response form provided.
Z-0622-06 | Mar-15-2006 | 2006 | Fujifilm - Synapse Software Versions 3.1.0 and 3.1.1 used witth the Fuji Synapse PACS | Fujifilm Medical System USA; Inc. | 2 | Image orientation and orientation markers may not match when an image is rotated or flipped when the CT Image is preset. | FujiFilm Medical Systems USA notified users by letter dated 2/3/06. Users are requested to deactivate the CT Image Processing functionality SWAT (Synapse Web Administration Tool).
Z-0219-2007 | Nov-28-2006 | 2006 | CR NX Modality Workstation - CR NX Modality Workstation; Computed Radiography System (Software versions NX 1.0.2402 and NX 1.0.2405). | AGFA Corp. | 2 | Users may experience one or more of four (4) problems; 1. Image quality problem with full leg full spine exams. 2. Unable to print 2 or more images on 1 film. 3. Possible image loss when printing. 4. Study date and time not showing up on Imexius Web. | A Service Bulletin to address the issues was created on 09/01/2006. Customers were notified of the plans to upgrade via telephone and email on 10/13/2006.
Z-1252-06 | Jul-26-2006 | 2006 | SoftPath - SoftPath ASCII Software; used in pathology labs to track and report specimens and diagnostic results. Releases 1.2; 2.1; and 2.3 | SCC Soft Computer | 2 | The client had made a decision to use the SoftPath Module in a nonstandard manner. Under specific circumstances; incorrect text may appear on a patient report. | The firm issued a software correction letter to its consignees 1/9/2001, advising them of the problem and supplied them with a software patch ICC number 3618 also referred to as release 3.0. The firm advised its clients that the installation of this software to their system was mandatory. .
Z-1353-06 | Aug-03-2006 | 2006 | EasyVision - Easy Vision Workstation Family (for the manipulation and displaying of x-ray images) with Multi Planar Reformat (MPR) Option | Philips Medical Systems North America Co. Phillips | 2 | Potential for measurements; lines; texts; etc to be stored incorrectly. | On 7/1/06 the firm issued a letter to consignees stating the problem and how to prevent the problem. The firm reports a software update will be issued no later than September 2006 to correct the problem.
Z-0713-06 | Apr-06-2006 | 2006 | SonoCalc IMT 3.0 - SonoCalc IMT 3.0; a windows based software used in conjunction with images from high-resolution ultrasound systems (MicroMaxx; TITAN; and SonoSite180PLUS). This software generates a report with the patients intima media thickness (IMT) based on average thickness of the carotid artery. | Sonosite; Inc. | 2 | When SonoCalc IMT 3.0 software is used with the SonoSite 180Plus ultrasound system; the software miscalculates the intima-media-thickness; a measurement of risk for cardiovascular or cerebrovascular events. Underestimates percentage of stenosis; false negative result. | On October 26, 2005, the company sent their customers a User Guide Addendum that told customers how to verify that the image calibration data was correct.
Z-0529-06 | Feb-16-2006 | 2006 | Lumenis - Lumenis brand Lumenis One System with Multi-Spot Nd:YAG; LightSheer and Universal IPL; Manufactured by Lumenis; Ltd; 13 Hayetzira Street; Yokneam Industrial Park; Yokneam; Israel 20692 | Lumenis Inc. | 2 | Software anomaly recorded incorrect filter in patient database and this may lead to patient skin burns. | On 4/8/05, the firm initiated the recall and its notification was via letters informing its customers of the corrective action. .
Z-1303-06 | Sep-30-2006 | 2006 | RUS Software - RUS Tool Version Software | GE OEC Medical Systems; Inc | 2 | X-ray systems could provide output which exceeds the 20 R/minute limit. | Consignees were notified by certified letter on 08/14-15/2006.
Z-0831-06 | May-03-2006 | 2006 | RenalSoft Observational Study v.2.0 - Baxter Healthcare Corporation; RenalSoft Observational Study v.2.0 - HD Module; PD Module and Medical Record Module; clinical data management software; Baxter Healthcare Corporation; McGaw Park; Illinois 60085 | Baxter Healthcare Renal Div | 2 | Software Error: There are software discrepancies that may increase the possibility of a medication error if the dialysis clinic software is used to track; record; or coordinate the administration of medications. | Baxter sent Urgent Device Correction letters dated 3/29/06 to those customers who had their RenalSoft and Renal Software Suite software configured with the affected modules. The letters indicated in a table which software anomalies were associated with the affected software modules in each software version. A description of each anomaly was provided, along with interim actions to take to mitigate the anomaly until the new software version RenalSoft V.2.0.x can be issued. Individual customer training will be performed at each facility for those issues identified in the letter specific to their software version. At that time the mitigation provided in the letter will be reviewed with the customer.
Z-0437-06 | Jan-24-2006 | 2006 | - Panorama Patient Monitoring Network; Cardiac Arrhythmia Monitor. | Datascope Corp | 2 | Software anomaly that affects the Panorama Telepack where if communications are lost during standby; certain alarms which appear set could be actually off. | Datascope sent letters to the distributors and end users on 12/7/2005. This communication states that a Datascope Service Representative will conduct and document the corrective action during a visit to the customer.
Z-0399-06 | Jan-18-2006 | 2006 | ISite PACS - ''iSite'' PowerScribe/iSite Integration; Picture Archiving and Communication System; software version 3.3 | Stentor Inc | 2 | When using a third party dictation accessory with the iSite PACS; a malfunction may occur that would result in the petient/exam selected and shown may be different on the radiologist console and diagnostic display screens | On 12/20/05, the firm initiated the recall and its notification was via letters informing its customers of the corrective action. .
Z-0358-06 | Jan-06-2006 | 2006 | System; Nuclear Magnetic Resonance Imaging - Gyroscan NT MRI System rev 4.x; 5.x; and 6.x software. System; Nuclear Magnetic Resonance imaging. | Philips Medical Systems North America Co. Phillips | 2 | The Gyroscan NT MRI system has a computer software problem which results in images from one patient being placed into another patient's image record. | Product safety notification sent certified mail to each customer explaining the problem. A mandatory software upgrade to correct the problem will be installed starting the first quarter of 2006 and be completed by the third quarter of 2006. .
Z-0520-06 | Feb-14-2006 | 2006 | RotoProne Therapy System - RotoProne powered patient rotation bed; part #209500; distributed by Kinetic Concepts; Inc.; San Antonio; TX. | Kinetic Concepts; Inc | 2 | To correct software issues in software versions 1.3 and 1.5 and to upgrade hatch latch handles to a more ergonomic design. | All KCI Service Centers notified via Technical Bulletin of upgrades. 1/24/2005.
Z-0161-2007 | Nov-08-2006 | 2006 | Allura Xper FD 10/10 - Allura Xper FD 10/10 x-ray; angiographic system | Philips Medical Systems North America Co. Phillips | 2 | Potential for x-ray system to freeze during an examination; and inability of table to be re-positioned. | A letter dated September 1, 2006 was sent to customers on September 5, 2006. The letter advised customers of the problem and advised a service representative will visit the consignee to install a new software upgrade.
Z-0797-06 | Apr-26-2006 | 2006 | SOMATOM Sensation CT Systems - SOMATOM Sensation CT Systems | Siemens Medical Solutions USA; Inc | 2 | Firm became aware of a potential problem that could affect scan results and require patient exam to be repeated when using the customized ECG-gated scan protocols in the CT2006A and CT2006G software of their SOMATOM Sensation CT Systems. | Firm has issued an advisory letter (3/28/2006) to affected customers per Update Instructions CT008/06/S. This letter informs customers about the potential problem and provides instructions to avoid this issue. A Siemens Service Engineer will visit affected sites to hand-deliver this letter and to delete any affected scan protocols, which can then be replaced with default protocols or with new customized protocols.
Z-0757-06 | Apr-11-2006 | 2006 | Digital mammography system. - SenoScan Full Field Digital Mammography System--Product Number 94001G-3; 105200G | Fischer Imaging Corporation | 2 | Software upgrade to preclude lost images during mammography procedure. | Consignees were notified by letter on 03/27/2006.
Z-1305-06 | Jul-28-2006 | 2006 | TriActiv ProGuard Embolic Protection System - TriActiv ProGuard Embolic Protection System; Part number 61000-01. The product is shipped as a procedure kit. It is used with a standard 6F guide catheter using the standard femoral approach. | Kensey Nash Corp | 2 | Alarm activiation-A priming issue involving the flow control unit due to a software problem has caused false postive Extraction Line Block (ELB) alarming. Use of this product may result in a health hazard resulting from early termination of the flush and extraction procedure after stenting | The recalling firm issued a draft recall letter to their European distributor who translated the letter for their customers.
Z-0554-06 | Mar-31-2006 | 2006 | Epic - Stryker Epic Critical Care Bed; Model 2030. | Stryker Medical | 2 | The bed exit may not alarm as appropriate. | The firm began visiting accounts to install new software on 1/17/06 and expects completion of installation at all accounts by 3/31/06.
Z-0862-06 | May-16-2006 | 2006 | REGIUS IM; RIM - REGIUS Image Manager (AKA REGIUS IM; RIM). Catalogue Number 5900554. Centralized CR Image Quality Control. Patient throughput optimization via offloaded processing from various CR readers. Technologist review and audit. Key processing for specialized exams; totally automated processing for hands off operation. Centralized workflow device for dispersed readers environments. HIS/RIS patient status updates. Demographic Correction. | Konica Minolta Medical Imaging USA; Inc. | 2 | If the patient demographic data is entered into the CS-1 manually; and the user fails to enter any data in one of the demographic data fields labeled as Birthday; Sex; Name or Patient Comments; etc; the RIM will populate the fields where no data has been entered with data from the previous patient. | The firm phoned the customers on 2/6/2006 to inform them of the recall and the corrective action associated to the recall. This will be followed up with a fax, e-mail and letter. The final step will be visiting the location to install the correct software.
Z-1375-06 | Aug-05-2006 | 2006 | VARiS Version 6.6 - Varis Vision Treatment 6.6/RTP Exchange/Siemens Accelerators | Varian Medical Systems Oncology Systems | 2 | A software anomaly may occur which can lead to patient treatment with the wrong field. The anomaly is reported only to occur when this software version (6.6.5022) is used with Elekta or Siemens linear accelerators . | On 09/20/2005, Varian Medical Systems issued a notification letter via certified mail and FedEx. These letters were issued with instructions on what to do regarding the corrections. Varian Medical Systems intends to release new, updated software.
Z-0223-2007 | Nov-29-2006 | 2006 | Roche Modular Drive - Roche Modular Drive (rchmdlri); analytical laboratory data interface. | Data Innovations; Inc. | 2 | Software of modular driver may incorrectly report patient results as Quality Control Results | Data Innovations notified users by email listserv and by letter dated 10/25/06.
Z-1161-06 | Jun-22-2006 | 2006 | OBSERVA - OBSERVA R02 (version R02.00.17) computers using PSC 6000 or HHP Barcode Scanners which are connected to BacT/ALERT 3D | bioMerieux Inc | 2 | Software problem with scanner inter-character delay which may result in false positive or false negative results after incorrect scanning of bottle ID's | The recalling firm telephoned customers who ordered the MB bottles and use the OBSERVA software ver. 2.0 to inform them of the issues. Recall letters were issued to all customers via overnight mail on 5/22/06 informing them a field representative would be visiting to make a short-term correction to the system. The letter also informed the customer a permanent correction to the software was scheduled for release later in the year.
Z-0165-2007 | Nov-14-2006 | 2006 | Allura - Allura X-per FD 10/10 x-ray; and fluoroscopy | Philips Medical Systems North America Co. Phillips | 2 | Potential for generator to lock-up. Fluoroscopy and x-ray is no longer possible; and system must be restarted to be able to continue. | On 10/12/06 the firm issued a letter to their consignees advising of the problem and how to avoid the problem by using specific image formats. If the system should lock up, consignees are to switch the system off and then restart. The firm will be conducting software upgrades to the devices once software is ready (estimated time for FD 10/10 is November 2006 and estimated time for FD 20/10 is January 2007). Later, the firm identified 8 additional units and issued a letter on 12/11/06 to affected customers.
Z-0795-06 | Apr-26-2006 | 2006 | SOMATOM CT System - SOMATOM CT System Sensation 40 and Sensation 64. Computed tomography x-ray system. | Siemens Medical Solutions USA; Inc | 2 | When using the LungCare feature in the SOMATOM CT Systems some marks made on CT images may not be saved. | The recalling firm has issued a software patch to the affected customers per Update Instructions CT086/05/S and CT087/05/S. By letter 3/9/2006. In addition, a Siemens Service Engineer will visit affected sites to install this software patch which will resolve this issue.
Z-0856-06 | May-09-2006 | 2006 | Diagnostic Ultrasound Transducer - Convex Array Transducer used with B-K Ultrasound ScannersModel number: 8667 | B & K Medical Systems; Inc. | 2 | Due to a problem with calibration; the device can produce up to 50% elevated acoustic output and up to 2 degrees higher surface temperature | BK Medical notified customers by letter on 4/10/06 via FedEx advising user of the problem and scheduling the software correction by a Service representative visit .
Z-1401-06 | Aug-11-2006 | 2006 | LANTIS Verify and Record software with XLINK option for Varian MLC - LANTIS XLINK (Impac MultiAccess Sequencer) on Varian accelerators with the MLC3PI interface | Siemens Medical Solutions USA; Inc | 2 | Improper treatment may be delivered if the machine configuration file is edited/set for the Varian MLC revision H file format. | The firm has issued a safety advisory letter to all affected consignees, and will distribute an updated software version as soon as it is available.
Z-0557-06 | Feb-23-2006 | 2006 | - The ETI-Mumps IgG Enzyme Immunoassay Test Kit is sold with assay software version 1.0. that is designed to work exclusively with the ETI-Max 3000 automated platform (version 1.51.1). | Diasorin Inc. | 2 | Potential for false negative results.A software anomaly may occur when performing the ETI-Mumps IgG assay on the ETI-Max 3000 automated platform. The error occurs when running a combined plate on the ETI-Max 3000; with the ETI-Mumps IgG assay as the second assay on a combined plate with ETI-Measles IgG and/or ETI-VZV IgG. | Consignees were contacted by telephone and then the firm sent a fax titled 'Diasorin Customer Notification' dated 8/8/05 that documented the information as provided by telephone. .
Z-1398-06 | Aug-10-2006 | 2006 | GammaMed software - GammaMed software program; version 5.07; vers. 5.08; vers.-usa; vers.-tst and vers.-row ; part number GM11005400; for the GammaMed model 12i radionuclide applicator system. | Varian Medical Systems | 2 | Software control program for a medical device used in radiation treatment may cause practitioners to incorrectly administer the treatment plan to cancer patients. There is a hazard when entering the treatment plans manually that the user neglects to change the default step size and/or the origin or accidentally enters incorrect parameters. | The recalling firm notified healthcare practitioners using the medical device by letter on 4/17/06 and continuing through 5/1/06. The notification was flagged as a 'Medical Device Recall' and instructed users to adjust system parameters in the autoexec.bat file. The notification further recommended using the identified treatment plans and that manual input should be used for QA purposes only.
Z-1255-06 | Jul-27-2006 | 2006 | ExacTrac Robotics - ExacTrac Robotics; sold alone as Robotic Tilt Module catalog 49720 and in a bundle as catalog 49700; for installation on the Varian Exact Couch; a component of the Novalis Shaped Beam Surgery System; a table tilt device - powered radiation therapy patient support assembly; BrainLab AG; Ammerthalstrasse 8; 85551 Heimstetten; Germany | Brainlab AG | 2 | During ExacTrac Robotics installation on the Varian Exact Couch; the couch height position indication is re-calibrated to read out the correct positions. This re-calibration could cause the eventual failure of the vertical lift mechanism of the couch; which might result in patient injury or death. | BrainLab issued a Product Notification Letter to the hospitals on 6/2/06 to advise the users not to lower the couch below 50 cm, and sent a second letter to the users on 6/7/06, informing them that BrainLab would remove the ExacTrac Robotics from the Exact Couch and re-calibrate the couch to restore the software limit switch. The Robotics will be re-installed once a permanent solution is available.
Z-1079-06 | Jun-03-2006 | 2006 | CXR-4CT (including Sceptre P3 PET/CT) - Hitachi Emission Computed Tomography SystemHitachi CX-4 CT (including SceptreP3 PET-CT) | Hitachi Medical Systems America Inc | 2 | The device has a software anomaly which causes blank images to be created during Multiplanar Reconstruction processing or results in an inherent filming error in which a reference scale is improperly sized. | Hitachi Medical Systems America, Inc. was informed of a software-related problem with the 50mm scale displayed on filmed images. This problem was reported to the firm by their Japanese manufacturer. In cases where the image has been magnified before filming, the 50mm scale is not correctly sized to the magnified image. The firm issued a Device Correctioin Notice to their customers on 2/21/2006. The notice concerns the referenced software problem which results in the resizing of the images, as well as a second previously discovered problem which results in blank multi-planer reconstructions (MPR''s). According to the Device Correction Notice, a software upgrade has already been installed on all of the units in order to correct the MPRs problem [Ref: RES #35081, Correction Removal Report #1528028-01/31/06-003-C] In response to the resizing issue, a second software upgrade [which the firm states address both issues] will be released on or before 2/27/2006, and this software upgrade will be performed by HMSA''s service personnel during each customer''s next scheduled PM. The Device Notification Letter also instructs the customers to call their Field Service Representatives or the HMSA Applications Helpline if they have any questions about the notice, or if any additional problems are encountered. .
Z-1078-06 | Jun-03-2006 | 2006 | Hitachi CXR-4 CT (including SceptreP3 PET/CT) - Hitachi CXR-4 CT Computed Tomography X-ray System (including SceptreP3 PET/CT) | Hitachi Medical Systems America Inc | 2 | The device has a software anomaly which affects the system's ability to create Multiplanar Reconstructions (MPR). Some images from the resulting series do not display correctly; or display at all; when MPR was run on them. | Hitachi Medical Systems America, Inc. became aware of an intermittent problem in the Multi-planar Reconstruction process when covering long ranges that could result in blank frames in the displayed reconstructions. HMSA also discovered a second software-related problem with the 50mm scale displayed on filmed images. In cases where the image has been magnified before filming, the 50mm scale is not correctly sized to the magnified image. Software corrections have already been installed on all units in order to correct the first of these problems (MPRs). According to the firm''s Device Notification Letter issued on 2/21/2006, a second software upgrade [which the firm states address both issues] will be released on or before 2/27/2006 and this correction will be performed by HMSA''s service personnel during each customer''s next scheduled PM. The Device Notification Letter also instructs the customers to call their Field Service Representatives or the HMSA Applications Helpline if they have any questions about the notice, or if any additional problems are encountered.
Z-0314-2007 | Dec-27-2006 | 2006 | Terumo APS 1 - Terumo Advanced Perfusion System 1; 100V -120V; 15 A (circuit breaker); 50/60 Hz (20A power source required); Models 801763. | Terumo Cardiovascular Systems Corp | 2 | The gas system may fail calibration prior to use due to incorrect gas system flowmeter software. | U.S. consignees were visited and corrections made between May 2004 and November 2005. Foreign affiliates were instructed to make corrections on international units, but the firm has not received confirmation yet that these units have been corrected.
Z-1466-06 | Aug-31-2006 | 2006 | Beckman Coulter FP1000 Cell Preparation System - FP1000 Cell Preparation System Part Number 624922 | Beckman Coulter Inc | 2 | During the cleaning cycle performed during the shutdown procedure of the Beckman Coulter FP1000 Cell Preparation System; fluid (diluted bleach) may drip from the probe in the location of the back reagent rack potentially resulting in bleach and/or water dripping into the reagents contaminating them and resulting in possible incorrect results. | A Product Corrective Action letter was sent on June 26, 2006 informing users that during the cleaning cycle performed during the shutdown procedure of the Beckman Coulter FP1000 Preparation System, fluid (diluted bleach) may drip from the probe in the location of the back reagent rack potentially resulting in bleach and/or water dripping into the reagents contaminating them and resulting in possible incorrect results. To reduce the risk of incorrect results from dripping fluid, users were advised to remove all patient samples, reagents, controls, and calibrators before initiating the shutdown cycle of the Beckman Coulter FP1000. Software will be modified in the next version to prevent the possibility of dripping bleach during the shutdown cycle. The Product Corrective Action letter is to be maintained as part of the customer''s laboratory Quality System documentation. The enclosed Response Form is to be completed and returned within 10 days to the recalling firm. Customers are also advised to contact the firm or their local Beckman Coulter Representative if there are questions.
Z-1437-06 | Aug-30-2006 | 2006 | TCM4 - TC Monitor - Transcutaneous Monitor for p02 and pC02; Model #TCM4; (Cutaneous Oxygen Monitor) | Radiometer America Inc | 2 | Alarm may be silenced- Rapid switching; on and off; of the TCM4 monitor may cause the built-in acoustical alarm to remain silent when tripped. | The firm initiated a recall action for the suspect device with a letter dated 2/2/2005. During a site visit to the U.S. customers conducted by the firm''s Field Service reps, the reps informed the user/customers of this problem and requested that all use of the device be discontinued, pending the formulation of a more permanent solution. Beginning on 2/8/2005, a second site visit was made to the same user/customers, at which time, the affected units were upgraded with operating software version 2.12. The firm''s recall/corrective action was scheduled for completion by 3/1/2005.
Z-0821-06 | Apr-27-2006 | 2006 | CALM perinatal monitoring system - CALM software; version 02.09.01 This software is intended for use as a central monitoring system with signal analysis and remote repeaters in the perinatal clinical environment. It interfaces with standard fetal and maternal monitors and provides fetal surveillance; monitoring of labor progress; charting; reporting and data analysis; display and archiving functions. | LMS Medical Systems Ltd | 2 | A deficiency in the software code can; in certain circumstances; cause data for the wrong patient to be displayed in the Tracing Review screen when a user is acknowledging an Alert | Consignees were first notified by telephone on 01/11/2005. A brief document was sent via email or fax on/about 01/13/2005.
Z-0069-2007 | Oct-25-2006 | 2006 | Sienet Cosmos - Sienet Cosmos V15C; picture archiving and communications system. Model number 10018969 | Siemens Medical Solutions USA; Inc | 2 | Potential loss of images. Images may not successfully be sent to the RAID system. | The recalling firm has issued a software patch to resolve the issue via Update Instructions HS110/06/S. The Update Instructions direct a Siemens Service Representative to visit the affected sites, install the software patch to resolve the issue, and hand deliver a Customer Advisory Notice, dated September 20th, 2006, to the customer to notify them of the issue.
Z-0533-06 | Mar-29-2006 | 2006 | - Toshiba America Medical Systems; Ultimax System MDX-8000A equipped with the EPS-Plus/ADR-1000A/E2D Digital Spot Film Device | Toshiba American Med Sys Inc | 2 | When the error condition occurs the x-ray exposures will terminate immediately. The digital system (EPS-Plus/ADR-1000A) will immediately stop acquiring images or will continue to acquire blank images. The remote console will lock-up upon release of the hand switch. The operator must reboot the system. This will result in the loss of the images acquired during the contrast portion of the study. | Recall letter to be sent to to customers of affected devices beginning February 23, 2006. Corrected software to be issued by Toshiba America Medical Systems Inc, for all affected devices. .
Z-0115-2007 | Oct-28-2006 | 2006 | Radiance Flex Blood Gas Anayzer - Radiance Data Management System; Software Version 2.60; modification for the ABL800 Flex Blood Gas Analyzer | Radiometer America Inc | 2 | Incorrect FI02 values: programming issue can cause incorrect values to be transmitted to the LIS when : 1)- The FIO2 result is edited in the manual sample processing mode--2)- An existing result is opened and FI02 is then edited and sent.....In both cases RADIANCE will transmit the original FI02 value; not the value that was edited. | The recalling firm initially contacted their customers and informed them of the software problem and the upcoming field correction for the software, via telephone on 7/6/2006. A follow-up letter dated 7/21/2006, was also sent to the customers.
Z-1104-06 | Jun-10-2006 | 2006 | AIRIS II; Altaire Magnetic Resonance Imaging Devices - AIRIS II; Altaire Nuclear Magnetic Resonance Imaging Device Systems | Hitachi Medical Systems America Inc | 2 | Software anomaly-A software defect can cause the slice line indicator to become mis-positioned during Multi-Planar Reconstructions (MPR) with the potential for patient mis-diagnosis. | The firm''s Field Service personnel will apprise the customers of this situation during scheduled visits at each customer location. During this visit, the corrected software will be installed by the Hitachi Field Service Technician. The installation of the software upgrades by the service technicians began on 4/28/2006. The firm anticipates that the correction of all units in their inventory as well as those units installed out-in-the-field will becompleted by 6/15/2006. .
Z-1243-06 | Jul-27-2006 | 2006 | SoftPath - Product is SoftPath ASCII Software Releases 2.3.0 and 4.3.7. The software is used in pathology labs for administrative and clinicial documentation and data processing. | SCC Soft Computer | 2 | Text from one case is overwriting another. This could cause an incorrect diagnostic report to be sent to a clinician/surgeon. | Consignees were notified and the products were corrected by the firm.
Z-0532-06 | Feb-23-2006 | 2006 | Toshiba America Medical Systems; Inc.; Digital Radiography System Model DFP8000D - Toshiba Digital Radiography System Model DFP8000D | Toshiba American Med Sys Inc | 2 | In certain conditions the FOV of the lateral plane does not coincide with the X-ray exposed area; and consequently a part of the X-ray field may not be displayed in the magnification mode. | An Advisory Letter will be sent to customers via certified mail. Corrected software will be provided at no charge. Customers can contact their local Toshiba Service person to schedule the software update.
Z-0855-06 | May-09-2006 | 2006 | - CR DX-S; DX-S; Computed radiography system (Digitizer) | AGFA Corp. | 2 | Corrupted image after system start; Corrupted image after workflow interruption; An erasure unit problem can cause system to stop. | Consignees were notified at the time of upgrading. The software was upgraded on 11/3/2005, 12/19/2005 and 12/28/2005. A Mandatory Service Bulletin, No. 11, explaining the issues was shared with the accounts. .
Z-0061-2008 | Oct-18-2007 | 2007 | Battery Pack portion of the LifeVest device - Battery Pack portion of the LifeVest device. The LifeVest is a wearable cardioverter defibrillator worn by a patient at risk for sudden cardiac arrest (SCA). | Zoll Lifecor Corporation | 3 | Possible malfunction of the battery pack. A significant electrostatic discharge (ESD) can cause the battery pack to malfunction. The malfunction causes the battery pack to discharge more rapidly and prevents the pack from recharging. | Patients currently using the LifeVest should continue to do so while the product removal is underway. Advisory notices will be sent via first class mail 9/14/2007 to all active patients using the LifeVest device. Product replacements began on 8/31/07 by shipping out revised battery packs for new patient use. Shipments of new battery packs will continue based on production capacity and material availability. Any patients expected to be using the device longer than 5 months will be provided with 2 replacement Battery Packs. They will be instructed to return their existing packs in the same shipping box.
Z-0988-2007 | Jul-03-2007 | 2007 | MediSpectra - MediSpectra Luma Cervical Imaging SystemModel Number: 3-02600 | Medispectra Inc. | 3 | Power supply may fail and render the unit inoperable. | MediSpectra notified users by telephone on 5/15/07 and follow-up letter. Users are requested to select a date for field servce to replace the power supply.
Z-0886-2007 | Jul-04-2007 | 2007 | LIFEPAK 20 defibrillator/monitor - LIFEPAK 20 defibrillator/monitor. The device is an AC powered defibrillator device with battery backup. The defibrillators are primarily used by healthcare providers in device user facilities such as hospitals and clinics. | Medtronic Emergency Response Systems; Inc. | 2 | Defibrillator internal battery cell may contain a short which could reduce battery capacity and cause the device to shut down without warning while operating on backup (DC) battery power. | Firm sent letters on 03/30/2007 dated 03/2007 via certified return receipt to consignees with affected LIFEPAK 20 devices. The firm advised consignees to use the Operator's Checklist in the Operating Instructions to perform daily user tests and function checks while operating the defibrillator on battery power. In addition consignees are advised to replace their battery every 1 to 2 years.
Z-0048-2008 | Nov-24-2007 | 2007 | Guidant VITALITY AVT ICD - Guidant VITALITY AVT; Models A135 and A155. Implantable Cardioverter Defibrillator; ICD. Atrial and Ventricular Therapies. Guidant Corporation Cardiac Rhythm Management; 4100 Hamline Avenue North; St. Paul; MN 55112-5798; USA. | Boston Scientific CRM Corp | 2 | Erroneous Display/Atypical Charge Time Behavior: 1)-End of Life (EOL) or ERI is displayed during mid-life (typically 24-48 months); even though battery capacity remains available. 2)- Extended ERI charge time limits: Charge times during mid-life may remain below a normal extension of the ERI charge time . ****** Prophylactic replacement of these devices prior to ERI is not recommended. As stated | A Product Update titled "ERI Charge Time Limit Extended During Mid-Life and Mid-Life Display of Replacement Indicators", dated 03/10/2007 was posted to Company's website on 03/13/2007. A CRM Product News letter was sent to physicians beginning 03/16/2007. The CRM Product News Letter gave an overview of the content of the Product Update and stated the web location for finding the complete Product Update. The Product Update identified affected devices, gave an overview of normal charge time behavior at mid-life for specific ICD and CRT-D devices. The letter also described the atypical charge time behavior, projected rate of mid-life display of replacement indicators and Patient Management Considerations. ****** Prophylactic replacement of these devices prior to ERI is not recommended. As stated in the Product Update, "Devices that have triggered charge time-based ERI or EOL during mid-life have several months, and in most cases more than one year of remaining battery voltage and capacity, which allows the devices in this pattern to continue to provide brady and LV pacing and maximum energy shocks. However, if ERI or EOL is triggered, device replacement should be scheduled."*****Recalled units only cover devices implanted prior to July 2005****.
Z-0047-2008 | Nov-24-2007 | 2007 | Guidant VITALITY ICD - Guidant VITALITY VR/DR; Implantable Cardioverter Defibrillator; ICD; Models 1870 and 1871. Guidant VITALITY DR+; Implantable Cardioverter Defibrillator; ICD; Model 1872. | Boston Scientific CRM Corp | 2 | Atypical Charge Time Behavior: 1)-End of Life (EOL) or ERI is displayed during mid-life (typically 24-48 months); even though battery capacity remains available. 2)- Extended ERI charge time limits: Charge times during mid-life may remain below a normal extension of the ERI charge time . ****** Prophylactic replacement of these devices prior to ERI is not recommended. As stated in the Product Upd | A Product Update titled "ERI Charge Time Limit Extended During Mid-Life and Mid-Life Display of Replacement Indicators", dated 03/10/2007 was posted to Company's website on 03/13/2007. A CRM Product News letter was sent to physicians beginning 03/16/2007. The CRM Product News Letter gave an overview of the content of the Product Update and stated the web location for finding the complete Product Update. The Product Update identified affected devices, gave an overview of normal charge time behavior at mid-life for specific ICD and CRT-D devices. The letter also described the atypical charge time behavior, projected rate of mid-life display of replacement indicators and Patient Management Considerations. ****** Prophylactic replacement of these devices prior to ERI is not recommended. As stated in the Product Update, "Devices that have triggered charge time-based ERI or EOL during mid-life have several months, and in most cases more than one year of remaining battery voltage and capacity, which allows the devices in this pattern to continue to provide brady and LV pacing and maximum energy shocks. However, if ERI or EOL is triggered, device replacement should be scheduled."*****Recalled units only cover devices implanted prior to July 2005****.
Z-1192-2007 | Sep-19-2007 | 2007 | Newport HT50 - Newport HT50 Ventilator (all Models) | Newport Medical Instruments Inc | 2 | Battery Life: With use over time; the internal battery begins losing its ability to hold a charge resulting in shortened ventilator operation time. Specifically; Newport has received a number of reports concerning the internal battery not lasting as long as expected followed by ventilator shut down with less advance warning than when the battery was new. | An Important Medical Device Correction Letter was issued on September 14, 2007 as a short term fix. The firm has submitted a long term correction as a supplement to its 510 K which is pending review by FDA.
Z-0598-2007 | Jun-30-2007 | 2007 | Terumo APS 1 - Terumo Advanced Perfusion System 1; 100V -120V; 15 A (circuit breaker); 50/60 Hz (20A power source required) base; Model 801763. | Terumo Cardiovascular Systems Corp | 2 | The power manager subsystem may not initialize properly during start-up; which will cause the battery status monitoring and battery switching operations to not be fully functional. | Firm repaired the affected units beginning in November 2005, upon receipt of complaints of failure. The firm issued a Device Correction Letter dated 6/6/07 notifying all users of the problem and advising them that they will be visited and their units upgraded as replacement components become available.
Z-0761-2007 | Apr-20-2007 | 2007 | Ceegraph/Sleepscan Netlink Traveler - Bio-logic-Ceegraph/Sleepscan Netlink Traveler; Bio-logic Systems Corp.; a nebus company; Mundelein; IL 60060; Model 580-T2ASM2. (a Digital EEG/Sleep Recorder Electroencephalograph that includes a built-in pulse oximeter; body position sensor; snore monitor; chest; abdominal and air flow transducers and electrode array for EEG; EMG and EOG) | Bio-Logic Systems Corp | 2 | Battery and Battery Compartment; incorrect placement in the Netlink Traveler unit which caused a short in the wiring of the battery pack and overheating. | On 2/26/07, Bio-logic Systems shipped replacement batteries, marked with part number 520-LIBAT4-3 and a label reading 'THIS SIDE OUT', and a new battery compartment cover with a label inside the cover to clarify the correct installation of the battery to each of the affected customers. The accounts were requested to remove and replace the old batteries, part 520-LIBTA4-2, and compartment covers, and return the old batteries and covers to Bio-logic Systems in the enclosed pre-paid shipping container, along with the completed accounting form.
Z-0840-2007 | Jun-07-2007 | 2007 | Endoscope light source. - Smith & Nephew; 420XL Xenon Light Source; 180 Watt; for endoscopic use. Catalog/Part #72200510; manufactured for Smith & Nephew Endoscopy; Andover; MA 01810. | Smith & Nephew; Inc Endoscopy Division | 2 | Several units were manufactured using a power supply that contains a 1 amp fuse where a 5 amp fuse is required. When the unit's power supply draws more that 1A; the fuse will fail causing the light to cease operating. The problem can occur anytime: from the initial turning on (powering up); to anytime during use. | Notification sent 04/09/07 via FedEx to distribution centers, sales reps and user facilities. Users asked to send light source to Smith & Nephew GMBH Service Center in Germany for repair (contact Ute Dieling at +49 (0) 7462-208-313).
Z-0286-2008 | Dec-20-2007 | 2007 | Ellex Solitarie Opthalmic Laser System - Ellex Solitarie Opthalmic Laser System; Model No: LP 4532 - Ellex Medical PTY; LTD 82 Gilbert Street; Adelaide; S.A. 5000; AUSTRALIA. | Ellex USA | 2 | Electric Shock from Power Supply : The power supply has a defect which can cause an electrical shock to user. | Ellex, Inc. called each customer on 6/13/07, about the problem (defect which can cause electric shock) and advised that they would be performing modification of the power supply in the laser system and scheduling a field service visit to complete the work. Upon completion of the modification, the service technician completed the service record and left a copy with each client and informed them of the completed work. A copy of each completed service record was sent to Ellex Medical Pty. Ltd. On 9/27/07, Ellex sent a letter to all consignees, informing them that all modifications by service technicians' have been completed, and that their Solitaire LP4532 laser and/or Intregre Duo LP1RG is up to the manufacturers specifications and is in good operating order.
Z-1203-2007 | Sep-07-2007 | 2007 | Kion Anesthesia Workstation - Kion Anesthesia WorkstationMaquet Critical Care ABSE-171-95 Solna; Sweden | MAQUET Inc. | 2 | Overcharging of batteries and the release of acid fumes | Safety Alert letters were sent out by FedEx tracking. A Service Express letter will be sent out emphasizing that only qualified batteries must be used in KION equipment and will list all qualified battery types. A Field Action - update of existing operating manuals will include warning on battery charging on KION. This action is currently scheduled to start 9/1/07 and be completed in US by 10/1/07. After FDA consultation the recalling firm issued a 9/5/2007 URGENT DEVICE CORRECTION Letter stating that KION will visit consignees and replace batteries older than three years, verify that the batteries are ones of the approved type, and verify battery condition.
Z-0287-2008 | Dec-20-2007 | 2007 | Ellex Integre Duo - Ellex Integre Duo; Model No: LP1RG - Ellex Medical PTY; LTD 82 Gilbert Street; Adelaide; S.A. 5000; AUSTRALIA. (The Integre Duo is a dual-wavelength surgical photocagulator laser intended for use in a clinic; outpatient clinic; or a retinal specialist's office. ) | Ellex USA | 2 | Electric Shock from Power Supply: The power supply has a defect which can cause an electrical shock to user. | Ellex, Inc. called each customer on 6/13/07, about the problem (defect which can cause electric shock) and advised that they would be performing modification of the power supply in the laser system and scheduling a field service visit to complete the work. Upon completion of the modification, the service technician completed the service record and left a copy with each client and informed them of the completed work. A copy of each completed service record was sent to Ellex Medical Pty. Ltd. On 9/27/07, Ellex sent a letter to all consignees, informing them that all modifications by service technicians' have been completed, and that their Solitaire LP4532 laser and/or Intregre Duo LP1RG is up to the manufacturers specifications and is in good operating order.
Z-0602-2007 | Jun-30-2007 | 2007 | Terumo APS 1 - Terumo Advanced Perfusion System 1; 100V -120V; 15 A (circuit breaker); 50/60 Hz (20A power source required) base; Model 801763. | Terumo Cardiovascular Systems Corp | 3 | The power supply may fail to function due to a defective capacitor. (Note; this device incorporates a redundant separate power supply not affected by this capacitor) | The power supplies were replaced as complaints were received between August 2005 and February 2006. The firm issued a Device Correction Letter dated 6/6/07 notifying all users of the problem and advising them that their units will be corrected when replacement components become available.
Z-1242-2007 | Sep-26-2007 | 2007 | HealthPax Cranial Electrotherapy Stimulator - HealthPax Cranial Electrotherapy Stimulator | Health Directions; Inc | 3 | Peak voltage test result not in acceptable limits. | The recalling firm sent a recall letter 07/30/2007 to all their customers. The letter indicated that it is necessary to return the product for testing and the device will either be returned or if necessary replaced. Each unit will be tested with an oscilloscope device to determine whether it meets or fails design specifications. The recalling firm will reimburse shipping costs.
Z-0872-2007 | Jun-06-2007 | 2007 | Smith & Nephew; Well Leg Holder - Well Leg Holder With Boot; part of the Supine Universal Hip Distractor System; Ref (Part) #72200632; Manufactured for Smith & Nephew; Inc.; Endoscopy Divsion; Andover; MA 01810. | Smith & Nephew; Inc Endoscopy Division | 3 | The device was not manufactured to correct specification causing interference when used in conjunction with Patient Transfer Board. | Firm notified consignees by letter and phone call on 11/1/2006. Consignees asked to return out-of-spec Well Leg Holder for a replacement.
Z-1007-2007 | Jul-07-2007 | 2007 | Medtronic MiniMed Paradign Insulin Infusion Pumps - Medtronic MiniMed Paradigm Insulin Infusion Pumps for the following Paradigm models: MMT-511; MMT-512; MMT-712; MMT-515; MMT-715; MMT-522 and MMT-722 | Medtronic MiniMed | 2 | Exposure to Magnetic Resonance Imaging (MRI) has resulted in damage to the component that monitors and controls movement of the motor in the MiniMed Paradigm insulin infusion pump. Although there were alarms as a result of the damage; some users cleared these alarms and continued using the pump. Under such conditions; the pump will significantly over-deliver; potentially causing severe hypoglycemi | Firm mailed Medical Device Safety letters on April 24, 2007, to healthcare professionals, existing pump users and MRI facilities reiterating the existing warning in the pump User Guide to avoid exposing the pump to MRI (or similar high strength electromagnetic fields) and strengthening previous warning by specifically mentioning the potential for over-delivery and severe hypoglycemia. Firm is also including an insert with this information with any Paradigm infusion pumps shipped to new customers.
Z-0483-2007 | Jan-11-2007 | 2007 | Colleague and Colleague CX Volumetric Infusion Pumps - Baxter Colleague Single Channel Volumetric Infusion Pumps; Baxter Healthcare Corporation; Medication Delivery Division; Deerfield; IL 60015 U.S.A.; Made in Singapore; product codes 2M8151; 2M8151R; 2M8161; 2M8161R | Baxter Healthcare Corp. | 2 | Baxter has identified the potential for a corruption of the memory chip in the pump to occur due to linear accelerator radiation exposure. If corruption occurs; it could result in an interruption of therapy with an audible and visual alarm notification to the user. | Baxter sent the 12/6/06 Safety Alert letter to all Colleague Infusion Pump customers, to the attention of the Director of Radiation Oncology with copies to the Medical Physicist and Vice President of Nursing, via first class mail to advise them that the Colleague pump is not recommended for use in Linear Accelerator radiation therapy suites. The letters informed the users that there is a potential for a corruption of the memory chip to occur due to linear accelerator radiation exposure. If corruption occurs, it could result in an interruption of therapy with an audible and visual alarm notification to the user. The letter listed the error codes that may occur if the pump has been exposed to this type of radiation: 18:200, 18:416, 36:416, 19:113, 41:113, 314:432 and 302:435:XXXX:0006. If the pump displays any of the codes, the user was requested to submit the pump to their service facility. Failure code 199 may also occur and requires reconfiguration of the pump settings by the user's biomedical services organization. Any questions were directed to Baxter Global Technical Services at 1-800-843-7867, prompt 2.
Z-1053-2007 | Jul-11-2007 | 2007 | VITALITY 2 - VITALITY 2 (model T165; T175); VITALITY DS (model T125; T135); VITALITY EL (model T127); Implantable Cardioverter Defibrillators (ICD); are designed to detect and terminate ventricular tachycardia (VT) and ventricular fibrillation (VF) and provide bradycardia therapy (atrial and ventricular pacing). Therapies include both low- and high-energy shocks using either a biphasic or monophasic waveform. Vitality 2 devices also offer a wide variety of antitachycardia pacing schemes to terminate slower; more stable ventricular tachyarrhythmias. Bradycardia pacing; including adaptive-rate features; is available to detect and treat bradyarrhythmias and to support the cardiac rhythm after defibrillation therapy. Devices denoted with DR offer dual-chamber bradycardia features (atrial and/or ventricular pacing and sensing); and the devices denoted with VR offer single-chamber bradycardia features (ventricular pacing and sensing). | Boston Scientific CRM Corp | 2 | Boston Scientific has recently identified low-voltage capacitors from a former supplier that may be subject to degradation at a low rate of occurrence. These capacitors may cause accelerated battery depletion and may reduce the time between elective replacement indicator (ERI) and end of life (EOL) to less than 3 months. | An Urgent Medical Devivce Information letter dated April 05, 2007 was sent to implanting and following physicians. The letter describes the issue, identifies patients with affected devices and provides recommendations.
Z-0641-2007 | Mar-22-2007 | 2007 | HeartStart FR2+ - Philips HeartStart FR2+ Defibrillator Automated External Defibrillator (AED)Catalog Numbers: M3860A/M3861A/M3840A/M3841A | Philips Medical Systems | 2 | Failure of tactile switches on one of two printed circuit board assemblies (PCBA) used in the FR2+ device. Potential delay or prevention of shock treatment due to PCBA board contamination; causing intermittent switch operation. | On 02/28/2007 the firm sent letters via return receipt to all domestic and foreign customers. The letter advises the customer of the issue and if the user encounters the problem, to press the button again with additional force. In addition, the firm sent with the letter a new Quick Reference Card. Customers are instructed to replace the old Quick Reference Card with the new one which instructs them to press 'On' and 'Shock' buttons firmly.
Z-0760-2007 | Apr-24-2007 | 2007 | Puritan Bennett 840 Ventilator System - Puritan Bennett 840 Ventilator System; Catalogue Numbers:4-84O12OXXXX-XX; 4-840220XXXX-XX; 4-NPB84O-XX; DL4-NPB84O-XX; DS4-NPB840-XX & 4-070550-SP(Note: the Xs indicate any combination of alpha characters which specify option and language packs) | Nellcor Puritan Bennett | 2 | In rare instances; the safety valve may remain open after an occlusion is resolved. If the safety valve does not reset; the ventilator will detect a circuit disconnect condition. When this happens; the ventilator will annunciate a high-urgency audible alarm and the ventilator will not resume normal operation. | The correction will extend to the end user. All users will be contacted by mail or through personal contact. In the U.S., all customers were mailed the 'Urgent Device Correction' letter along with a serial number listing of those units which records show were consigned to them beginning on March 12, 2007. The customer communication informs customers of the issue and provides instruction for continued safe use of the device pending the correction.
Z-0376-2007 | Jan-25-2007 | 2007 | D-10 Dual Program - Bio-Rad D-10 Dual Program Calibrator Diluent; a component of the Bio-Rad D10 Dual Program diagnostic test. This is a specific part of the kit; and only the calibrator diluent is defective. It comes in two kits. | Bio-Rad Laboratories Inc | 2 | The Diluent in the D-10 Dual Program Calibrator reorder packs is defective; and can cause early calibration failure or elevated control and sample recovery results. | The firm has issued phone calls and faxes to its US consignees starting 11/17/2006 and International notices began on 11/20/2006 explaining the reason for the recall and requesting the product be disposed.
Z-1002-2007 | Jul-07-2007 | 2007 | Terumo perfusion - Terumo Perfusion System 8000 Roller Pump; 115V; Model 8000; Catalog No. 16402. | Terumo Cardiovascular Systems Corp | 2 | Roller Pump: Due to a circuit board malfunction the roller pumps may stop during use without displaying an error message; may stop during use with an overspeed error posted on the local display; or may fail to start/restart resulting in the pump(s) being non-functional. | As the complaints were received (between 07/04 and 01/06), the firm visited the customer filing the complaint and repaired the pumps demonstrating the problem. The firm will correct other units as complaints are received. The firm modified its recall strategy and issued a Device Correction Letter dated 6/22/07 notifying all users of the problem.
Z-1005-2007 | Jul-07-2007 | 2007 | Terumo perfusion - Terumo Perfusion System 8000 Roller Pump; 1040V; Model 8000; Catalog No. 16411. (Not distributed within the United States). | Terumo Cardiovascular Systems Corp | 2 | Due to a circuit board malfunction the roller pumps may stop during use without displaying an error message; may stop during use with an overspeed error posted on the local display; or may fail to start/restart resulting in the pump(s) being non-functional. | As the complaints were received (between 07/04 and 01/06), the firm visited the customer filing the complaint and repaired the pumps demonstrating the problem. The firm will correct other units as complaints are received. The firm modified its recall strategy and issued a Device Correction Letter dated 6/22/07 notifying all users of the problem.
Z-0416-2007 | Feb-07-2007 | 2007 | ARCHITECT i2000SR Processing Module - ARCHITECT i2000SR Processing Module; for in vitro diagnostics; List Number 3M74-01; Temperature Controller Board Part #78560-109; distributed by Abbott Laboratories; Irving; TX. | Abbott Laboratories; Inc | 2 | The re-designed ARCHITECT Temperature Controller Board has been identified as being susceptible to electromagnetic interference (EMI) in the laboratory. This can cause a board reset condition; stopping temperature control function of the board. Situation does not stop analyzer operation and does not generate an error condition alerting operator when condition occurs. | Firm contacted consignees by telephone and fax with instructions to verify proper functioning of the Temperature Controller Board starting December 13, 2006.
Z-0343-2007 | Jan-19-2007 | 2007 | Welch Allyn PIC 50 Portable Monitor/Defibrillator - Welch Allyn PIC 50 Portable Intensive Care System; Multi-Parameter Monitor/Defibrillator; a portable 12V internal battery powered defibrillator; Manufactured by MRL Inc.; a WelchAllyn Company; 1000 Asbury Drive; Buffalo Grove; Illinois 60089-Part Numbers-971081;971082;971083;971084. | MRL; Inc.; A Welch Allyn Company | 2 | Shock Delay:The monitor-defibrillator may display an 'ECG Comm' error message on the display during use; which may prevent or cause an unacceptable delay the delivery of a defibrillation shock. This delay could result in failure to resuscitate the patient. | MRL sent Urgent Medical Device Recall letters dated 11/8/06 via certified mail on 11/15/06 to the direct accounts informing them of an intermittent connection between the J5 cable and its mating connector on a circuit board inside of the device that could result in a malfunction that could prevent or unacceptably delay the delivery of a defibrillating shock, which possibly could result in failure to resuscitate the patient. The accounts were provided a list of the affected units'' parts/serial numbers and were requested to inspect their PIC 50 units for the affected serial numbers and return the identified units for servicing. Until their PIC 50 is serviced, the accounts were instructed to check the unit before each shift by using the self-test, which is automatically initiated by powering up the device. If the 'ECG Comm' error occurs during use, cycling the power off-on may clear the error. Even if power cycling clears the device, the accounts were instructed to contact Welch Allyn - MRL to obtain priority servicing of the affected unit. The accounts were also instructed to respond to the notification within 30 days by fax or mail, or by calling Welch Allyn - MRL Technical Support at 1-800-462-0777, to schedule their PIC 50 for preventative service. Welch Allyn will provide the accounts with a loaner PIC 50 after receiving the form with instructions to send back the PIC 50 in the same shipping box accompanying the loaner unit. Any questions were directed to Welch Allyn-MRL Technical Support at 1-800-462-0777.
Z-0304-2008 | Dec-19-2007 | 2007 | c.cam Gamma Camera - c.cam Gamma Camera; Model No: 9VIR1200; Tomography System (Emission Computed); Siemens Medical Solutions USA; Inc.; Molecular Imaging; 2501 N. Barrington Road; Hoffman Estates; IL 60195-5203; Made in Denmark | Siemens Medical Solutions USA; Inc. | 2 | Unintended Motion: A hardware fault in the c.cam system can cause the system to experience unintended motion; thereby causing a collision hazard. | Urgent Field Correction Recall letters dated 9/11/07, on Siemens letter head were sent to the end users via U.S. certified mail, informing them of the potential risk of unintended motion of the detector head resulting in a collision hazard. The end users were warned not to press any of the detection motion buttons while simultaneously holding the collision over ride button on the c.cam hand controller, and were provided steps to take if the system stops during an acquisition and displays the Collision Error message. They were told that a Siemens service representative will contact them within 90 days to schedule the repair of the c.cam at no cost to the end user. Any questions regarding the Important Safety Notice were directed to their service representative or Siemens office at: 1-800-888-7436 (USA).
Z-0929-2007 | Jun-30-2007 | 2007 | Terumo APS 1 - Terumo Advanced Perfusion System 1; 100V -120V; 15 A (circuit breaker); 50/60 Hz (20A power source required) base; Model 801763. | Terumo Cardiovascular Systems Corp | 2 | System 1 power supply may electrically short to the main console; disrupting normal system operation by causing unexpected re-setting or shutdown of pumps and modules; false alarms and/or overheating and failure of the central control monitor. | The firm corrected the complaint units at the time of the complaints in 2005. The firm issued a Device Correction Letter dated 6/6/07 notifying all users of the problem and advising them that their system will be corrected when replacement components become available.
Z-0323-2008 | Dec-01-2007 | 2007 | Terumo APS 1 - Terumo Advanced Perfusion System 1 Integrated Centrifugal System Control Unit; Catalog Number 801046; Terumo Cardiovascular Systems; Ann Arbor; MI 48103 | Terumo Cardiovascular Systems Corp | 2 | Failure to operate: The unit display may go blank and reset when the start button is pressed; due to an internal electrical short; resulting in a centrifugal pump that fails to start as expected. | Consignees were notified via an Urgent Medical Device Correction letter dated 9/10/07, given instructions to follow if the problem were to manifest itself, and informed the firm would contact them to schedule an upgrade.
Z-0075-2008 | Nov-08-2007 | 2007 | Suinsa Flexi-DT Table - Sunisa Flexi-DT Table; Model Numbers: 51335678; 2269647; 2242088-2; 5141609-2; 5115763; and 5128145;Manufactured by Suinsa C Primavera; 39 Pol. Ind. Las Monjas; Tarrejon de Ardaz Madrid 28850 Spain | GE Healthcare | 2 | Design Problems Including: 1) Hot Brake (Longitudinal Overheating) potential risk of burns 2) Brake Locks and Sudden Releases 3) Emergency Stop Circuit Button; open instead of closed 4) Longitudinal Brakes Release does not engage properly 5) Tilting Movement causing tabletop to slide; may cause risk of pinched fingers to patient and/or operator 6) Lodging Pins on floor Uncovered; may cause potenti | Recalling firm initiated correction by visit on August 2007. Corrections include replacement of one or more of the following: charger board, pin lodging, rear bumper. Also, replace the interface board and remove resistor mounted on chassis, remove brake pedal and assembly of cosmetic brake or apply caution labels.
Z-0322-2008 | Dec-01-2007 | 2007 | Terumo Advanced Perfusion System 1 - Terumo Advanced Perfusion System; 1 Roller Pump; 6 inch diameter; Catalog No. 801041; Terumo Cardiovascular Systems; Ann Arbor; MI 48103 | Terumo Cardiovascular Systems Corp | 2 | Failure to Operate: The pump may fail to operate; and be difficult to hand crank; due to an internal electrical short condition. | Consignees were notified via an Urgent Medical Device Correction letter dated 9/10/07, which provided them guidance in the event the problem occurs and advised them they would be contacted to schedule replacement of the circuit boards.
Z-0687-2007 | Mar-31-2007 | 2007 | Boston Scientific Thermal endometrial ablation system - HTA Console System; 110 Volt; Product Number: 56000 (HTA system is a software-controlled hysteroscopic thermal endometrial ablation system) | Boston Scientific Corporation | 2 | Defective integrated circuit board could result in loss of the system pump and patient injury (hot fluid 90 degree C into uterus) if in use during the recirculation/heating phase of treatment | Boston Scientific notiifed accounts by letter dated January 24, 2007 via federal Express Mail. Accounts were requested to examine inventory, remove and return product.
Z-1090-2007 | Jul-23-2007 | 2007 | Medtronic Performer CPB - Medtronic Performer CPB-Universal R2100375. Catologue Number R5100052. The Performer CPB System is an electromechanical extracorporeal support device used during cardiopulmonary bypass procedures. Manufactured by: RanD S.r.l 41036 Mendolla (MO) Italy. Distributed exclusively by Medtronic; Inc. | Medtronic Perfusion Systems | 2 | Device failure requiring use of back up systems (malfunctioning printed circuit board (PCB)) | There are two separate letters dated June 4, 2007 for the customers, based on the configuration of the unit being inspected/updated in the field. These letters explains the situaiton and informs the customers that Medtronic will be visiting each customer site to inspect/update the units. Medtronic Service Technicians will complete a Field Action Certicate acknowledging the successful completion of the update.
Z-0395-2007 | Feb-01-2007 | 2007 | Boston Scientific iLab¿ - Boston Scientific brand iLab¿ Ultrasound Imaging System; for intravascular use;Model Number(s): iLab120lNS; iLab2401NS; Distributed by and Manufactured by:Boston Scientific Corporation IVUS; Technology Center;47201 lakeview Parkway; Fremont; CA 94538 | Boston Scientific Target | 2 | The product fails to comply with applicable electronic product performance standard; in that the electromagnetic energy is in excess of the labeled conformance standards for radiated emissions. | A customer notification was sent December 1, 2006 indicating the need for the correction.
Z-0321-2008 | Dec-01-2007 | 2007 | Terumo Advanced Perfusion System 1 - Terumo Advanced Perfusion System; 1 Roller Pump; 4 inch diameter; Catalog No. 801040; Terumo Cardiovascular Systems; Ann Arbor; MI 48103 | Terumo Cardiovascular Systems Corp | 2 | Failure to operate; The pump may fail to operate; and be difficult to hand crank; due to an internal electrical short condition. | Consignees were notified via an Urgent Medical Device Correction letter dated 9/10/07, which provided them guidance in the event the problem occurs and advised them they would be contacted to schedule replacement of the circuit boards.
Z-0129-2008 | Nov-06-2007 | 2007 | Welch Allyn AED10 - Welch Allyn AED 10; Automated External Defibrillator; a portable 12V internal battery powered defibrillator; Model AED10; Part Numbers: 970302E;970308E;970309E; 970310E and 970311E; | Welch Allyn Protocol; Inc | 1 | Failure to Deliver Shock; a defective capacitor may cause the delay or non-delivery of the defibrillating shock which may result in failure to resuscitate the patient | The firm issued a press release on 10/30/07, and has set up a call center for customers, and plans to send three (3) certified mailings to notify affected customers, alert them of the issue, and provide instructions on how to return their device. The first letter dated 10/26/07 was sent via certified mail to customers on 10/26/07. The letter informs customers of the firm recalling the product to correct faulty capacitors on the circuit board, and which could result in a malfunction that prevents or delays delivery of a defibrillating shock. The letter indicates the faulty AED10 devices were manufactured between 3/29/07 and 8/9/07. The letter also indicates the serial numbers affected, and where the serial number is located - upper left hand corner of the back panel of the device. In addition the letter includes a response form, the firm's contact information, and indicates that the firm will exchange the recalled defibrillator with a replacement and new five (5) year warranty.
Z-1006-2007 | Jul-03-2007 | 2007 | ResMed S8 Continuous Positive Airway Pressure (CPAP) Flow generator - ResMed S8 Continuous Positive Airway Pressure (CPAP) Flow generator with the following model names and product codes: S8 AUTOSET VANTAGE; #33112; S8 COMPACT; #33030; S8 ELITE; #33021; S8 ESCAPE; #33007. | Resmed Corporation | 2 | S8 devices manufactured between July 2004 and May 15; 2006 are susceptible to fatigue in the connection of the AC appliance inlet connector and the power supply (power cord) leading to possible failure (short circuit; electrical failure; sparks; smoke). | ResMed began to broadly distribute information regarding the recall on 4/23/2007. A press release was issued on 4/23/2007 with information regarding the recall and information was posted at ww.resmed.com/s8program. The firm is providing a Consignee notification package and press release to include(1) specific device names subject to the recall and affected serial numbers; specific information to be provided in each package to the end-customer. The specific information to the end user customers will include: (1) specific device names subject to the recall; (2) affected serial numbers; (3)continuing use and safety information; and (4) contact information for requesting additional details. Patients are advised that they may continue to use their S8 flow generators until a replacement unit is provided. ResMed is instructing patients, sleep labs, and hospitals to continue using the device, placing it on a hard clean surface, with a clear area around the device. ResMed intends to replace the affected products with either a new device or a device in which the affected components have been replaced.
Z-0491-2007 | Feb-14-2007 | 2007 | Terumo APS 1 - Terumo Advanced Perfusion System 1; 6 inch diameter Roller Pump; Model 801041. | Terumo Cardiovascular Systems Corp | 2 | The pump may lose communication with the central control monitor; causing the monitor to display a service message "System Computer Needs Service" and to become non-operational. | Consignees were notified of the problem via letter dated 11/10/06 and informed that they would be contacted to schedule a service call to install software upgrades. Installation of software upgrades began at customer locations on 10/16/06.
Z-0600-2007 | Jun-30-2007 | 2007 | Terumo APS 1 - Terumo Advanced Perfusion System 1; 4 inch diameter (small) Roller Pump; Model 801040. | Terumo Cardiovascular Systems Corp | 2 | The internal ribbon cable may loosen from the roller pump display board; resulting in loss of local control and/or local display. | The firm corrected units in the field or brought them back for evaluation and repair as complaints were received between 9/04 and 4/06. The firm issued a Device Correction Letter dated 6/6/07 notifying all users of the problem and advising them that they will be visited and their units upgraded as replacement components become available.
Z-0492-2007 | Feb-14-2007 | 2007 | Terumo APS 1 - Terumo Advanced Perfusion System 1; Integrated (Sarns) Centrifugal System Control Unit; Catalog number 801046. | Terumo Cardiovascular Systems Corp | 2 | The pump may lose communication with the central control monitor; causing the monitor to display a service message "System Computer Needs Service" and to become non-operational. | Consignees were notified of the problem via letter dated 11/10/06 and informed that they would be contacted to schedule a service call to install software upgrades. Installation of software upgrades began at customer locations on 10/16/06.
Z-0039-2008 | Oct-11-2007 | 2007 | STA Compact and STA Compact CT - STA Compact and STA Compact CT; Haemostasis Analyzer; in vitro diagnostic. | Diagnostica Stago; Inc. | 2 | Incorrect results. Internal investigation determined the presence of a defective measurement head which can impact the results of clotting tests by either yielding no result or rendering an error message. | Recall notification letters were mailed to consignees on 7/30/07 by first class mail.
Z-0709-2007 | May-23-2007 | 2007 | Precision Link Software - Abbott Precision Link Software (Version 2.1); (Version 2.2); (Version 2.3); (Version 2.4); (Version 2.5); and (Version 2.6); Serial Number(s): ALL; | Abbott Diabetes Care; Inc. | 2 | If the time and date are not set correctly on both the user's meter and computer; there is the potential for incorrect times and dates to be assigned to the results in the Precision Link Software. | Firm initiated its recall on 12/21/2006. The firm provided customer letters with additional information (eg. flyer or insert) to customers that they had addresses for via fedex. In addition, the product notification was posted on the firm's internet website. Further, a flyer is to be included as a supplement to the user manual.
Z-0342-2007 | Jan-25-2007 | 2007 | Cyberonics VNS Therapy System - Handhelds-Cyberonics VNS Therapy System; Model 250-''Handhelds'' preprogrammed with software v6.0; v6.1; v6.1.7; v7.0 and v7.1.3; manufactured by Cyberonics; Inc.; Houston; TX. | Cyberonics; Inc | 2 | During programming; pulse generator may be inadvertently set to 8.0 mA output; regardless of the mA range selected by the clinician | Firm sent safety alert to physicians on 11/14/06. Safety Alert informs physicians of the problem and provides instructions on how to avoid an event and correct an event should it occur.
Z-0064-2008 | Nov-03-2007 | 2007 | AxSYM Digoxin III Reagent Pack - Abbott Diagnostic International; Ltd.-- AxSYM Digoxin III Reagent Pack; List (6L07-20); For In Vitro Diagnostic use. This reagent system is composed of one bottle of Digoxin-Alkaline Phosphatase Conjugate; one bottle of Anti-Digoxin (Rabbit) Coated microparticles; one bottle of MEIA buffer; and one bottle of Digoxin Probe wash solution. | Abbott Diagnostic International; Ltd. | 2 | Incorrect Results -- Error codes 1062; 1063; 1113; and 1118 observed when running patient samples on the AxSYM Digoxin III assay; rather than a concentration; is reported for the patient sample. | Abbott Diagnostics International Limited, Ltd., mailed a Product Correction Letter to customers on 7/20/02 advising an additional option(s) has been identified to resolve Error Codes: - Perform one of the two options to resolve these error codes 1062, 1063, 1113, and 1118 observed when running patient samples on the AxSYM Digoxin III assay. - Dilute the sample one to one (1 part sample plus 1 part Calibrator A) and repeat analysis. - The final concentration is equal to the reported concentration of the sample multiplied by the dilution factor - The minimum final concentration that may be reported out is 0.6 ng/mL, which corresponds to the sensitivity of the AxSYM Digoxin III assay (0.3 ng/mL) multiplied by the dilution factor (2). OR - Centrifuge the sample at 8,000 to 10,000 RCF for 10 minutes and repeat analysis.
Z-0372-2007 | Jan-25-2007 | 2007 | SERVO-i - Ventilator System SERVO-i V3.1 PC1777Maquet Critical Care | MAQUET Inc. | 2 | Reported failures of the SERVO-i ventilators user interface system causing technical errors and increased fault frequencies. Switches may be activated unintentionally; or a function not activate when the switch is selected. Direct access knobs may not function or change the parameter in the wrong direction compared to how they are turned. | Maquet sent an ''Extended Warranty'' letter on May 23, 2006. On March 13, 2007, Maquet issued a Recall Notification letter to all customers stating that they will replace all PC1777 boards in the serial number range 2400-3300, which have not previously been replaced. A safety alert letter was included outlining the failure indication and effect for the user of the failures, as well as, how to identify whether the Servo-i unit has shown any failure indications.
Z-0059-2008 | Nov-10-2007 | 2007 | BV Pulsera - Philips Medical System - BV Pulsera Mobile C-arms with Remote Control Option; (Release 2) | Philips Medical Systems North America Co. Phillips | 2 | System Shut Down -- Unexpected Shut Down during use of the Optional Remote Control. | On 8/23/07, the firm sent out a Safety Notification letter, via certified mail. On 8/29/07, the firm sent out an URGENT - DEVICE CORRECTION letter via regular mail, dated 8/20/07. Both letters advised customers that if the device shuts down during use, it can be restarted within about 1 minute by switching the system off and then on again, using the system on/off buttons. The letter(s) states that software is being developed to correct the problem.
Z-0356-2008 | Dec-08-2007 | 2007 | Stryker inTouch Critical Care - Stryker inTouch Critical Care bed; Model 2130; Stryker Medical; Portage; MI 49002 | Stryker Medical Div. of Stryker Corporation | 2 | Multiple defects: 1) beds exhibit reverse motions when the Fowler; gatch or foot section is activated 2) a battery charging defect which may result in a depletion of battery power 3) the touch screen on the footboard freezing; resulting in controls to powered; pressure-relief mattresses being difficult to access 4) beds equipped with zoom drive systems may have an overdrive problem which could re | Consignees were notified via an Urgent Medical Device Correction letter dated 11/2/07, and instructed to remove the beds from service if they display any unintended motions and informed that one of the firm's field service representatives will visit the facility to make corrections to all beds under recall.
Z-0848-2007 | Jun-30-2007 | 2007 | GE OEC 9900 Elite Fluoroscopic x-ray system - GE OEC 9900 Elite Fluoroscopic X-ray System; GE Healthcare; Salt Lake City; Utah. | GE OEC Medical Systems; Inc | 2 | X-ray systems were not configured to display Air Kerma Rate and Cumulative Air Kerma. | Consignees were notified by letter on 06/12/2007 and told that a service representative would arrange for an inspection and would modify their systems.
Z-0056-2008 | Oct-13-2007 | 2007 | Angiomat Illumena - Angiomat Illumena Injector with Auxiliary Console; Mallinckrodt; Inc.; Hazelwood; MO. | Mallinckrodt Inc | 2 | Synchronization problems: units that have both a main programmable console and an auxiliary programmable console; when simultaneously; injection parameters are programmed into the unit's main console and the unit's auxiliary consoles; mis-synchronization could occur resulting in a variance between the intended injection parameters (those programmed) and the actual parameters of the injection deliv | On 9/20/2007, a notification was sent from Mallinckrodt to all direct accounts via certified mail. The notification details the potential for mis-synchronization of injection parameters if the main and auxiliary consoles are simultaneously programmed. The notification reminds end-users that simultaneous programming should not be performed and to ensure it is not performed, a notification label (provided in the notification) is to be affixed on both consoles of the injector system. Once the label is affixed to the system, end-users are instructed to complete the notification survey and return the survey to the Mallinckrodt. In addition, the notification includes an updated chapter insert (Chapter 11.1) for the injector system's product manual that emphasize the separation between programming of the main and auxiliary consoles. Questions regarding this matter may be directed to Mallinckrodt's Product Monitoring Department at 1-888-744-1414, menu option 1 and then menu option 2.
Z-1089-2007 | Jul-23-2007 | 2007 | BCI Advisor - BCI Advisor Vital Signs Monitor (Model/Catalog Number: 9200) | Smiths Medical PM; Inc. | 2 | Turn off unexpectedly/ Not turn on : monitors turned off unexpectedly sounding the critical failure alarm. Advisor Vital Signs monitors that were off; would not turn back on. | Consignees will be contacted via a Safety Action Bulletin 07-SAB02 dated June 5, 2007. Technical Worksheet, 07-TW01 including a kit will be sent to the International consignees whom Technical Service has deemed qualified to perform the repair. The repair will be performed at the consignee site. International consignees who require or request the repair be done at the Smiths Medical PM, Inc. Technical Service department will return their monitors for repair.
Z-1245-2007 | Sep-26-2007 | 2007 | daVinci S Surgical System IS2000 - daVinci S Surgical system; patient side cart Auxiliary power board (APB) (printed circuit board assembly); endoscopic instrument control system. | Intuitive Surgical; Inc. | 2 | Product may malfunction and fail to start up on AC power. | As of 9/12/2007, the firm has conducted a field inspection of any involved units via service visit. No other notice is planned or indicated by the firm's submission. The firm feels that the risk associated with this issue is minimal, and therefore, any advanced notification is unneccessary. If any system failed one or more of the inspection points, its APB (3510404-04) was replaced with a verified APB (351040-04 or 351040-05).
Z-0890-2007 | Jun-05-2007 | 2007 | Sienet Magicstore - Sienet MagicStore. Radiological Digital Image Communications System | Siemens Medical Solutions USA; Inc | 2 | System could crash and result in data loss. | The recalling firm issued a Safety Advisory letter to their customers (May 2, 2007) to inform them of the problem. A company representative will be contacting the customers to arrange for replacement of the RAID component.
Z-0660-2007 | Mar-28-2007 | 2007 | EndoWrist - 8mm EndoWrist Bipolar Maryland Instrument; Part Number: 400172-04; Distributed and Manufactured by Intuitive Surgical; Inc. 950 Kifer Road Sunnyvale; CA | Intuitive Surgical; Inc. | 2 | The products were incorrectly programmed as training instruments; which allows for 30 uses (surgical procedures) instead of 10 uses. | The firm initiated its recall in March 24, 2005, when the firm sent a letter via Federal Express mail to all 18 consignees, informing them of the issue, how to identify affected product, and instructions to return the affected instruments for a full credit. The form included a reply form that consignees were requested to fill out and return.
Z-0296-2008 | Nov-27-2007 | 2007 | Xoran MiniCAT for ENT. - Xoran MiniCAT for ENT; Xoran Technologies; Incorporated; Ann Arbor; MI 48103. This device is a cone beam CT scan for sinue; temporal bones and maxillofacial imaging. | Xoran Technologies; Inc. | 2 | The hard drive may crash; resulting in loss of data from patient CT scans. | Consignee accounts were visited between 7/25/06 and 8/10/07 and the hard drives on their systems were replaced.
Z-1223-2007 | Sep-19-2007 | 2007 | Welch Allyn AED 20 Defibrillator - Welch Allyn AED 20 Automated External Defibrillator; a portable; 12V internal battery powered defibrillator; Manufactured by MRL Inc.; A Welch Allyn Company; Buffalo Grove; Illinois; Model AED 20; Catalog Numbers 972213; 972214; 972215; 972216 (Part Numbers 972200E; 972211E; 972212E; 972213E; 972214E; 972215E; 972216E) | Welch Allyn Protocol; Inc | 1 | Failure to deliver therapy: The Welch Allyn AED20 Defibrillators may display a "Defib Comm" error message on the device display during use which may result in failure of the device to analyze the patient's ECG and deliver the appropriate therapy. | On 8/23/07 the firm sent a press release to the Associated Press for release. On 8/29/07 the firm began sending URGENT MEDICAL DEVICE RECALL letters to their consignees informing them of the possibility of the defibrillator to fail to resuscitate the patient. The accounts were asked to respond to the notification within 5 working days by faxing or mailing an enclosed response. End users will be provided with a loaner device while their unit is being serviced. Distributors were sent a separate letter request they provide their customer names to the firm so that the firm can send letters to those customers.
Z-0854-2007 | Jun-05-2007 | 2007 | LADAR6000 - LADAR6000 Excimer Laser System | Alcon Refractive Horizons; Inc. | 1 | Observed 'central islands' in patients following custom myopia laser procedures using the LADAR6000 Excimer Laser. (On May 11; 2007; Alcon notified consignees that the device software programs for CustomCornea Myopia and CustomCornea Myopia with Astigmatism of the device would be deactivated) | Consignees were notified of the recall via two letters. On Feb. 21, 2007, consignees were requested to immediately stop performing all custom cornea myopia procedures and to return an acknowledgement card indicating they have received the alert and understand the required actions. On May 11, 2007, Alcon notified consignees that the device software programs for CustomCornea Myopia and CustomCornea Myopia with Astigmatism of the device would be deactivated. After meeting with CDRH's Office of Compliance on August 27, 2007, the recalling firm committed to total product removal of the device. Notification will begin in Sept 2007 and will be made by personal visits to each consignee.
Z-1001-2007 | Jul-03-2007 | 2007 | Ascensia Contour Diabetes Care System - Ascensia Contour Blood Glucose Monitoring System ( Bayer )Diabetes Care System; Product 7152A.---NO USA Units are affected---. | Bayer Healthcare LLC | 1 | Incorrect Unit of Measure displayed : The meters report results in the wrong unit of measurement for Canadian users. | The firm began calling Canadian customers on or about 6/06/07 to inform them of the problem, to tell them that a meter with the correct measurement units will be sent to them, and to request that they return the meter currently in their possession.
Z-0470-2007 | Jan-13-2007 | 2007 | ProTimePro Microcoagulation System - International Technidyne Corporation (ITC) ProTime Microcoagulation System Instrument-Catalogue Number-Protimeint & Protimepro; Protimeintrf & Protimeprorf; Protimeintl & Protimeprol. | International Technidyne Corp. | 3 | Incorrect Test Results-ProTime ITC''s Microcoagulation System: Professional Instrument may report incorrect test results when more than 50 tests are in the memory and the technician uses the AUTO Send or PRINT RESULTS feature. | Medical Device Correction letters were sent on 12/6/2006 by first class mail along with a replacement Operator''s manual.
Z-0253-2008 | Nov-17-2007 | 2007 | NEMIO - NEMIO Ultrasound System; Model SSA-550A ; Software versions: V3.1.000B; V3.5.0; V3.5.000A; V3.5.000B; V3.6.0; V3.6.000A; V4.0.0; V4.0.000A; V4.l.0; V4.l.000A; V4.2.0; V4.3.0; V4.5.0; V4.5.000A; Toshiba America Medical Systems; Inc.; Tustin; CA. | Toshiba American Med Sys Inc | 3 | Multiple problems including : (1) The surface temperature of the transducer tip can rise and exceed the limit set by the international standard. (2) Incorrect display of the TI indices on the display monitor. The displayed value is greater or smaller than the true TI value; and (3) Manuals contain incorrect values for MI/TI values. | Urgent Medical Device Notification letters were sent to customers on August 29, 2007. Customers were notified of the several issues related to the Nemio SSA-550A and Nemio SSA-580A recall and asked to share this information with all users of the device. Customers were informed that Toshiba is preparing software and corrected manuals for distribution to address these issues. If customers have any questions they were instructed to contact their Toshiba Customer Service Engineer or call (800) 421-1968 and ask for the Regulatory Affairs Department.
Z-0081-2008 | Nov-06-2007 | 2007 | Medtronic Paceart - Medtronic Paceart System; 2006 Edition; Medtronic; Inc. 710 Medtronic Parkway; Minneapolis; MN 55432-5604 USA. The Paceart System can act as a Practice Management software application and is intended for use by facilities that provide services such as event monitoring; trans-telephonic pacemaker evaluation; and in-office evaluation of implanted cardiac devices. | Medtronic Inc. Cardiac Rhythm Managment | 3 | An issue with the Paceart System - 2006 First Edition (Get Connected Edition) exists. Under certain circumstances; Paceart System Generic ICD reports may not accurately reflect patient VT/VF; SVT/NST and Mode Switch/AT/AF episode detection data. When episode detection data is imported into the Paceart System from certain sources; the Paceart System Generic ICD reports display a zero; suggesting | A Medtronic letter was sent to consignees on 8/27/07. The letter described the issue and the product involved. It also notified customers that Medtronic is developing a software update that will correct the issue. An Acknowledgement form was attached to be filled by consignees as confirmation and understanding of the information received.
Z-1095-2007 | Jul-24-2007 | 2007 | ABL 800 FLEX analyzers - ABL 800 FLEX series blood gas analyzers containing software version 5.24. | Radiometer America Inc | 3 | Software anomaly. Protein removing program incorporated in the operating software for Versions: V3.833 and V5.24; of the ABL700 and ABL800 Blood Gas Analyzers; may allow for protein build up on the electrode membranes of blood gas analyzers. This condition can influence analyzer results. When the protein removal program is run too frequently; more than once a week; or when running the protein rem | The firm issued a 'Safety Alert Notice' [FAN 915-248] to all customers which received the recalled ABL700 and ABL800 Analyzers equipped with the referenced software. The 11/13/2006 notice describes the problem concerning the Protein Removal program to the user/customers and emphasizes that the frequency of running the program in these software versions must be no more than once a week in order to prevent the protein build-up problem. It also contains additional recommendations for the user/customers for operation of these analyzer and informs the customers that a Radiometer Service Rep. will be coming out and making a software upgrade to these systems within the next few weeks. The alert notification is accompanied by a 'FAX FORM' to be signed and completed by each customer in order to verify that the alert notification has been received and that the staff members responsible for operating and maintaining the analyzer(s) on-site have been informed of this situation and that the necessary changes in the Operator's Manual concerning the Cleaning, Protein Removal, and Decontamination programs have been made.
Z-0380-2007 | Feb-15-2007 | 2007 | NicoletOne monitoring system version 5.20 software - NicoletOne version 5.20 software (also known as NicoletOne Version 5.2 Software) used with the NicoletOne LTM; Sleep; nEEG; vEEG and ICU monitoring Systems. | Nicolet Biomedical Div of Viasys Healthcare | 3 | A software anomaly exists in the 5.20.1038 NicoletOne LTM; Sleep; nEEG; vEEG and ICU Monitor systems using the M; and/or C series amplifiers. If this anomaly occurs; the system will display the data for channel one in all channels on NicoletOne LTM; Sleep; nEEG; vEEG or ICU Monitor systems. | An Advisory Notice, dated August 25, 2006, was sent to all affected customers. The notice described the issues, states to discontinue use of 5.20.1038 software until updated 5.21.1039c software is installed.
Z-0700-2007 | Mar-31-2007 | 2007 | Precision Spinal Cord Stimulation System - The Precision Implantable Pulse Generator; Model Number SC-1110 | Advanced Bionics Corp | 3 | The Precision Implantable Pulse Generator with firmware Revision 3.01 incorrectly reports an error condition for a valid output from its self internal check. | March 2007 by letter to physicians who had implanted the units notifying them of the recall and future upgrade by wireless RF telemetry. .
Z-0718-2007 | Apr-05-2007 | 2007 | CADD - Prizm - Deltec branded CADD-Prizm PCS II ambulatory infusion pump and the CADD-Prizm VIP system ambulatory infusion pumps; (Both) Model 6101; Order #21-8861-01 (PCS II) and 21-8821-01 (VIP). Smiths Medical MD; Inc.; St. Paul; MN 55112 USA; RX ONLY. | Smiths Medical MD; Inc. | 3 | The year of manufacture had been incorrectly entered as 1980 into the pumps software during the manufacture process. As a result the pump may go into an immediate 'Clock battery is Low/Service Immediately' or 'Clock Battery needs service soon' alarm message indicating that the pump requires servicing. | All consignees who have received the affected CADD-Prizm PCS II ambulatory infusion pumps and the CADD-Prizm VIP System ambulatory infusion pumps were contacted via telephone by February 1, 2007. The problem will be explained to the consignee and they will be asked to return the pumps to Smiths Medical MD in order for the problem to be corrected.
Z-0041-2008 | Oct-11-2007 | 2007 | PEMView - PEMView Software used with Naviscan PEM Flex Solo I and PEM Flex Solo II PET Scanner; for nuclear medicine imaging. | Naviscan PET Systems | 3 | Naviscan has identified a potential safety issue with all versions of PEM View Software commonly used for analyzing patient images acquired on the PEMFIex product line.1. The Positron Emission Mammography Standardized Uptake Value (PEMSUV) and distance measurements for the product mentioned above may be inaccurate if a zoomed image has been opened; measurement tools are used; and the applet rema | Telephone notification regarding issues #1 and #2 were carried out on August 1 and 2, 2007 and included instructions for a validated work-around to avoid the software issues. Written notifications with these same validated work-around instructions were sent on August 2, 2007 via certified mail. Customers were notified that Naviscan has identified a potential safety issue with all versions of PEM View Software. Issue #3 was identified after the notification to customers for issues #1 and #2. Telephone notification regarding issues #3 was carried out via telephone and certified mailing beginning August 16, 2007. Customers were informed that Naviscan is preparing software to resolve these issues and that they would be contacted by Naviscan Customer Support to ensure they thoroughly understand the issues and the recommended procedure to prevent these issues from occurring. If customers have any questions they were instructed to contact Naviscan Customer Support at 1-888-628-4159 or service@navsicanpet.com.
Z-0768-2007 | Apr-26-2007 | 2007 | Medfusion 3500 - Medfusion 3500 Syringe Infusion Pump with PharmGuard Medication safety Software versions 4.1.3 and 4.1.4. Medex; 4350 Green Parkway; Suite 200. Duluth; GA 30096 | Smiths Medical MD; Inc. | 2 | Software anomaly that causes the pump to experience a sudden and unexpected failure. This anomaly occurs during the pump's startup sequence when the pump's Preventative Maintenance Alarm goes off; preventing the pump from completing its power up routine. Lock-up occurs and the pump will not function. | Consignees were initially notified of the recall via a Customer Information Bulletin which was sent via electronic mail 2/3/07 and confimed with a telephone conversation. A Product Recall Notification and a Product Recall Bulletin will be sent to all consignees via the mail on 2/16/07.
Z-0671-2007 | Mar-29-2007 | 2007 | LIFEPAK - LIFEPAK CR Plus defibrillator. The device is a battery operated; portable automated external defibrillator (AED). The devices are designed for non-professional users and are primarily used in public access defibrillator environments such as airports; schools; or public libraries. | Medtronic Emergency Response Systems; Inc. | 2 | LIFEPAK CR Plus defibrillators with software version 1.16 or lower may miscalculate the HLC battery capacity and result in the CHARGE-PAK battery charger icon/indicator to display and falsely indicate battery needs replacement. | Firm sent letters on 03/02/2007 dated 03/2007 via certified return receipt to consignees with affected LIFEPAK CR Plus devices. The firm advised consignees to upgrade their devices' software, and indicated that each consignee will be provided a user-installable software upgrade for their device(s) within 45 days.
Z-0773-2007 | May-03-2007 | 2007 | BacT/ALERT - BacT/ALERT Classic System Software; Blood Culturing System | bioMerieux; Inc. | 2 | Plastic bottle samples assigned less sensitive default glass bottle algorithm instead of the plastic bottle algorithm on BacT/ALERT Microbial Detection Systems. | Consignees were notified via A Field Corrective Action Notification on/about 10/18/2006. A second notification was sent on October 31, 2006.
Z-1000-2007 | Jul-07-2007 | 2007 | Viatron T-Series Dual Chamber pacemakers - Vitatron T-Series Dual Chamber pacemaker T60 DR; model T60A1. The Vitatron C60 DR is a dual chamber rate responsive pacemaker (activity sensing using an accelerometer) for permanent atrial and ventricular pacing. T-Series software VSF07/VSF08 Version 1.2. Vitatron is a division of Medtronic; Inc.; Sales Office is Vitatron USA; 7000 Central Ave.; NE; Minneapolis; MN 55423-3576. The head office is Vitatron B.V.; P.O. Box 5227; 6802 EE Arnhem; The Netherlands | Medtronic Inc. Cardiac Rhythm Managment | 2 | Pacing rate: Software anomaly affecting Vitatron dual chamber C-series and T-series pacemakers; related to automatic retrograde conduction testing; can inhibit the device from pacing if the patient's intrinsic rate falls below the programmed lower rate. | Advisory letters to physicians were mailed in February 2007. Sales respresentatives will follow-up with physicians by visit. A programmer software update will be released upon regulatory approval. The Advisory Letter provided short-term direction to reduce problems pending software upgrades.
Z-0891-2007 | Jun-06-2007 | 2007 | CIC Pro System - GE Healthcare CIC Pro Clinical Information Center; Central Station; Software Version 4.1. ; GE Medical Systems Information Technologies; Inc.; 8200 W. Tower Ave. Milwaukee; Wisconsin USA. (System hardware consists of a processing unit and a monitor) | General Electric Medical Systems Information Technology | 2 | Alarm and Reboot problems: Two separate issues could occur with the CIC Pro system when used with Telemetry. Each issue occurs in software version v4.1 only; A) the system may experience a loss of audible alarms when operating system memory resources reach 170MB and/or B) the system may intentionally or unexpectedly reboot . | GE Healthcare sent a certified letter titled "Urgent Medical Device Correction" to customers on 3/13/07 that explained the short and long term solutions and that the company would be providing a software update that will rectify the issue specific to CIC Pro v4.1 software.
Z-0999-2007 | Jul-07-2007 | 2007 | Viatron C-Series dual chamber Rate Responsive Pacemakers - Vitatron C-Series C60DR DDDR Dual Chamber Rate Responsive Pacemaker includes the following models: C60A1 and C60A3. The Vitatron C60 DR is a dual chamber rate responsive pacemaker (activity sensing using an accelerometer) for permanent atrial and ventricular pacing. C-series software is VSF11/VSF12 Version 1.0. Vitatron is a division of Medtronic; Inc.; Sales Office is Vitatron USA; 7000 Central Ave.; NE; Minneapolis; MN 55423-3576. The head office is Vitatron B.V.; P.O. Box 5227; 6802 EE Arnhem; The Netherlands. | Medtronic Inc. Cardiac Rhythm Managment | 2 | Pacing rate; Software anomaly affecting Vitatron dual chamber C-series and T-series pacemakers; related to automatic retrograde conduction testing; can inhibit the device from pacing if the patient's intrinsic rate falls below the programmed lower rate. | Advisory letters to physicians were mailed in February 2007. Sales respresentatives will follow-up with physicians by visit. A programmer software update will be released upon regulatory approval. The Advisory Letter provided short-term direction to reduce problems pending software upgrades.
Z-1170-2007 | Aug-18-2007 | 2007 | CELL-DYN 1700 Hematology Analyzer - CELL-DYN 1700 Hematology Analyzer; List Numbers 03H53-01 and 03H53-03; software controlled medical device for diagnostic use.; and CELL DYN Control Assay Disk; LIst Numbers 01H91-01 and 01H92-01 when used with above CELL DYN 1700; Abbott Laboratories Diagnostics Division | Abbott Laboratories | 2 | Abnormal results may not be "flagged" appropriately because of a software anomaly. (Flags for the Hemoglobin (HGB); MCH and MCHC parameters when the upper assay paramter limit has been exceeded) | The firm initially sent reports (Product Information Letters) of this issue to the field on May 7, 2007. The firm sent a notice with customer notification response form on May 16, 2007 (Product Correction Letter). In addition, product information letters are being packed with the instruments accessory kit and quality control products used with the Cell-Dyn 1700 system.
Z-0262-2008 | Nov-24-2007 | 2007 | Edwards Vigilance II Monitor - Edwards Vigilance II Continuous Cardiac Output/Oximetry/Volumetric (CCO/SvO2/CEDV) Monitor Model: VIG2/VIG2E Software Versions: 00.51;00.60; 00.61; 00.62; 00.63; 00.64; 00.67; 00.69; Edwards Lifesciences LLC.; Irvine; CA 92614 | Edwards Lifesciences; Llc | 2 | MAP offset errors: If users are "slaving" mean arterial pressure (MAP) from the bedside monitor into the Vigilance II Monitor for the purpose of the continuous calculation of systemic vascular resistance (SVR); at times; on some bedside monitors; the Vigilance II monitor may display a MAP offset of as much as 25mmHg. Since the MAP is used in calculations of various patient parameters such as SVR; | An Urgent Customer notification letter dated May 17, 2007 was mailed to all Vigilance II customers advising that If they are "slaving" mean arterial pressure (MAP) from their bedside monitor into the Vigilance II Monitor for the purpose of the continuous calculation of systemic vascular resistance (SVR), to be aware that at times, on some bedside monitors, the Vigilance II monitor may display a MAP offset of as much as 25 mmHg. Since the MAP is used in calculations of various patient parameters such as SVR, these may also not be accurate. A table is provided showing those bedside monitors with the problem: GE Solar, Phillips (HP) M1006A, M1006B, as well as, those bedside monitors verified to not have the problem: SpaceLabs Model 90402 and Nihon Kohden Model BSM 8000. The letter also serves to remind the customer of the Warning in the Vigilance II Directions for Use which is also provided in the letter that instructs user to verify slaved MAP values with those displayed on bedside monitor. A temporary work around for this offset problem is to manually enter the MAP based on the readings from the bedside monitor. All Vigilance II calculations from the MAP have been re-verified to be correct when the MAP is accurate. The problem is in the transfer of the MAP between the bedside monitor and the Vigilance II. Customers are also advised that if they have the GE Solar or the Phillips (HP) models listed in the letter or any monitor not listed as verified to be compatible with the Vigilance II (Space labs and Nihon Kohden models listed above) Edwards recommends verification of MAP accuracy prior to treating based on parameters from a Vigilance II in a "slave" configuration. If there is a difference, they should enter the MAP into the Vigilance II manually for a static calculation of SVR. Edwards will be supplying a software update to correct this problem.
Z-0366-2008 | Dec-29-2007 | 2007 | Cardiac Viewer - Philips Medical Systems; Cardiac Viewer or Pulmonary Viewer Application that can run on all Gemini TFs; GXLs and EBW Workstations. Philips Medical Systems; Cleveland; OH 44143. (The product is a software application that is installed in Philips Medical System CT scanners.) | Philips Medical Systems (Cleveland) Inc | 2 | Incorrect Measurement: Incorrect measurement will occur when the operator uses the distance or area measurement function; such as Graphic Tool line or ROI on zoomed and batch saved images from the Cardiac Viewer or Pulmonary Viewer and then views them in the CT Viewer or on a PACS System. | Consignees were notified of this recall by an Urgent Device Notification letter sent on 9/10/07. The letter instructed the users to not save any zoomed images to a batch file. Make all measurements of zoomed images in the CT viewer, which provides correct measurements. Do not perform measurements on any zoomed batched images that have been stored on a PACS System. Phillips in investigating the problem and will provide updates free of charge when completed.
Z-0367-2008 | Dec-29-2007 | 2007 | Pulmonary Viewer - Philips Medical Systems; Cardiac Viewer or Pulmonary Viewer Application that can run on all Brilliance CT scanners using software versions 2.0 through V 2.2.5 and all EBW workstations using software versions 2.0 through V3.5.2. Philips Medical Systems; Cleveland; OH 44143. (The product is a software application that is installed in Philips Medical System CT scanners.) | Philips Medical Systems (Cleveland) Inc | 2 | Incorrect Measurement: Incorrect measurement will occur when the operator uses the distance or area measurement function; such as Graphic Tool line or ROI on zoomed and batch saved images from the Cardiac Viewer or Pulmonary Viewer and then views them in the CT Viewer or on a PACS System. | Consignees were notified of this recall by an Urgent Device Notification letter sent on 9/10/07. The letter instructed the users to not save any zoomed images to a batch file. Make all measurements of zoomed images in the CT viewer, which provides correct measurements. Do not perform measurements on any zoomed batched images that have been stored on a PACS System. Phillips in investigating the problem and will provide updates free of charge when completed.
Z-0265-2008 | Nov-28-2007 | 2007 | GE Healthcare Advantage Workstations - GE Healthcare Advantage Workstations; Advantage Worksation version 4.2 and 4.3 with Volume Viewer 2 (version 6.4-6.8) provided with Advance Vessel Analysis (AVA) and Advantage Worksation 4.4 with Volume Viewer 2 (version 6.8 or before) or Volume Viewer 3 (version 7.2 -7.3) provided with Advanced Vessel Analysis (AVA); Model 2378698-2;GE Healthcare; Waukesha; WI. 53188 | GE Healthcare | 2 | Aortic Length over-estimated: incorrect measurements when using the Aorta analysis protocol or customized protocols. Until a new software version is installed the aorta analysis protocol or the customized protocols for an aneurysm should NOT BE USED. | An Urgent Safety Notice to customers dated 9/10/07, was issued by the recalling firm informing them of the device defect, safety issues, warnings on future uses and promised new software updates to be installed by GE representatives, free of charge, beginning Nov. 2007.
Z-0312-2008 | Dec-22-2007 | 2007 | COULTER LH750 Analyzer - COULTER LH750 Analyzer; Part Number 6605632; Beckman Coulter; Brea; CA. (Beckman Coulter; Brea; CA. (Hematology analyzer) | Beckman Coulter Inc | 2 | Patient Mis-identification can occur:1)- When a positive identifier (Sample ID or Cassette/Position) is manually edited to a positive identifier that is already in the To Do list. The workstation will accept the entry and no error message will be generated; creating 2 samples with the identical positive identifier.2)- When manually entering a Patient ID; if a blank space is entered between th | Product Corrective Action (PCA) letters were mailed on April 18,2007, to all GENS, LH500, LH750 and LH780 accounts. They are informed that a potential misidentification can occur in 2 scenarios. 1. When a positive identifier (Sample ID or CassetteIPosition) is manually edited to a positive identifier that is already in the ToDo list. The workstation will accept the entry and no error message will be generated, creating 2 samples with the identical positive identifier. 2. When manually entering a Patient ID, if a blank space is entered between the characters in the Patient ID field, the system will only accept the characters before the space.----Short term fix----- The letter explains the issues and provides the customer work around. (Letters were sent by US mail). A customer response form was provided. The root cause of the problem is still under investigation, software upgrades, when approved will be made to devices.
Z-0951-2007 | Jun-19-2007 | 2007 | ARCHITECT i2000sr - ARCHITECT i2000sr Processing Module; for in vitro diagnostics; List Number 3M74-01; v2.10; v2.11; v2.12; v2.20; v2.20DB; v2.30; v2.60 and v3.00; distributed by Abbott Laboratories; Irving; TX. | Abbott Laboratories; Inc | 2 | One or more of four software issues : 1) calculation error for Architect cSystem assays using the Spline math model. 2) The system configuration option; Run controls onboard reagents by Kit; does not function correctly for assays requiring a standard sample dilution. 3) Infrequent data may be missing from automatically or manually printed reports generated on the Architect system. 4) Clinical Chem | Notification letters sent to consignees on 04/30/07. Consignees provided with work around instructions for the software glitches and asked to respond via fax using Customer Reply form acknowledging receipt, understanding and completion of correction letter. This is the short term fix pending software updates, validation,and full root cause determination.
Z-0146-2008 | Nov-07-2007 | 2007 | ARCHITECT cSystem - ARCHITECT c8000 Processing Module; for In Vitro Diagnostics; List Number 106-01; Affected software: v2.10; v2.11; v2.12; v2.20; v2.20DB; v2.30; v2.60; v3.00 and v3.10; Distributed by Abbott Laboratories; Irving; TX. | Abbott Laboratories; Inc | 2 | Incorrect Sample/Test Identification: Due to a software timing defect with certain versions of ARCHITECT cSystem Software; test results may be incorrectly associated to a Sample Identification during specific sample-processing conditions of the Robotic Sample Handler. | Notification letters sent to consignees on 09/17/07. Consignees informed of the software defect and the conditions that would cause the software error to occur. Consignees told they would be contacted by an Abbott Field Service Rep who would install ARCHITECT System Software v3.11 and were asked to respond via fax using Customer Reply form acknowledging receipt and understanding of correction letter. Firm recalling to hospital/user level.
Z-0566-2007 | Mar-06-2007 | 2007 | COULTER LH 700 Series - Coulter LH700 Series Hematology Analyzers; Part Number 6605632 with 2B Software (2B3; 2B4 and 2B5) | Beckman Coulter Inc | 2 | There is a potential for erroneous yet credible results for Body Fluids on the LH700 series hematology analyzers when a cassette label fails to read and the subsequent sample is cycled in Body Fluid mode. | A Product Corrective Action (PCA) letter was mailed on Sep 12, 2006 to all LH 750 accounts to inform them that there is an issue with COULTER LH 700 Series analyzer software that could affect the Body Fluid analysis and that this this anomaly is present only in Software revisions of 2B. The letter instructs the customer on how to determine their software version. Immediate user preventive action was provided in the letter which is to reset the analyzer when a cassette label read error occurs. The letter advises that this information be shared with laboratory staff and be retained as part of their Quality System documentation. A customer response form was also included.
Z-0437-2007 | Feb-14-2007 | 2007 | Terumo APS 1 - Terumo Advanced Perfusion System 1; 6 inch diameter Roller Pump; Model 801041. | Terumo Cardiovascular Systems Corp | 2 | The large roller pump may experience false "overspeed' error message when starting rotation; which will prevent the pump from operating; until the condition is cleared by power cycling the roller pump. | Consignees were notified of the problem via letter dated 11/10/06 and informed that they would be contacted to schedule a service call to install software upgrades. Installation of software upgrades began at customer locations on 10/16/06.
Z-0439-2007 | Feb-14-2007 | 2007 | Terumo Advanced Perfusion System - Terumo Advanced Perfusion System 1; 6 inch diameter Roller Pump; Model 801041. | Terumo Cardiovascular Systems Corp | 2 | The roller pump will pause in response to a pressure alarm condition; as designed; but then may not resume operation as expected after the alarm condition has been cleared. | Consignees were notified of the problem via letter dated 11/10/06 and informed that they would be contacted to schedule a service call to install software upgrades. Installation of software upgrades began at customer locations on 10/16/06.
Z-0040-2008 | Oct-12-2007 | 2007 | Interstim for Urinary Control - Medtronic InterStim INS (Model 3023) with N'Vision Inter Stim-B software supplied on the Model 8870 Version MMB_01/NNB_01 with N'Vision Clinician Programmer Model 8840 or InterStim iCon Patient Programmer Model 3037. Implantable Neurostimulator for Urinary Control. | Medtronic Neuromodulation | 2 | Software issue may cause false low battery alert to be displayed when interrogating an Interstim Implantable Neurostimulator (Model 3023) with an N'Vision Clinician Programmer (Model 8840); or with an Inter Stim iCon Patient Programmer (Model 3037). | Begining on July 11, 2007 a letter to the physicians was sent informing them of the issue and patient management for the two issues. The letter includes a Cover Letter, A Reply Form, and two Recall letters explaining the issue in detail.
Z-0037-2008 | Oct-11-2007 | 2007 | Interstim for Urinary Control - Medtronic InterStim II INS Model 3058 with N'Vision InterStim-B Software supplied on the Model 8870 Version MMB_01/NNB_01 N'Vision with a Clinic an Programmer Model 8840. Implantable Neurostimulator for Urinary Control. | Medtronic Neuromodulation | 2 | A software issue may cause a lower than expected battery capacity estimate when evaluating an InterStim II Implantable Neurostimulator (model 3058) with an N'Vision Clinician Programmer(Model 8840) before implant . | Begining on July 11, 2007 a letter to the physicians was sent informing them of the issue and patient mangement of the two issues. The letter includes a Cover Letter, A Reply Form, and two Recall letters explaining the issue in detail.
Z-1132-2007 | Aug-07-2007 | 2007 | Mac-Lab - GE Medical Systems Mac-Lab System; cardiac catheterization data system. Software Version 6.5 2021423-015 Revision D GE Medical Systems; Information Technologies; Inc.; 8200 West Tower Avenue; Milwaukee; WI 53223 USA. | General Electric Medical Systems Information Technology | 2 | GE Healthcare has identified conditions that could cause a potential loss of real-time monitoring with the CardioLab/Mac-Lab/ComboLab Electrophysiology and Hemodynamic Monitoring Systems. These issues are associated with an audio driver; and the use of the application's log and imaging windows. | Consignees were sent a letter dated 2/16/07. The letter described the product and problem, made recommendations and gave information on software update.
Z-0480-2007 | Jan-11-2007 | 2007 | V.I.P Bird Gold - Viasys V.I.P. Bird Gold pediatric ventilator; Catalog Number 15653 | Viasys Respiratory Care; Inc. | 2 | The reason for this correction is due to the potential of a repeat ''failure to cycle'' event. | On November 2, 2006, a product notification letter was mailed to all customers in the United States via overnight carrier (Federal Express). Consignees are informed in the notification letter that service calls are in process of being scheduled to update domestic consignees. On November 4, 2006, a product notification letter was mailed to all international customers via registed mail. Consignees are informed in the notification letter that replacement software, replacement instructions, and installation verification forms are included in their recall package.
Z-1212-2007 | Sep-25-2007 | 2007 | The Pinnacle3 Radiation Therapy Planning System - The Pinnacle3 Radiation Therapy Planning (RTP) system and its associated software; composed of several modules including the core Pinnacle3 functionality; Syntegra (TM); P3IMRT; P3MD and AcQSim3 (TM). The Pinnacle3 RTP system is composed of a Sun UNIX workstation (or UNIX-complaint computer) running the Solaris operating system and software; and a UNIX terminal emulation package on a personal computer; which allows the user to enter patient data into the system; use that data to construct a plan for radiation therapy; and evaluate the plan- Philips Medical Systems-Model #9200-0652A-ENG; 9200-0700A-ENG; 9200-0700B-ENG; 9200-0700C-ENG (with software versions 7.6C; 8.0d and 8.0h) | Philips Nuclear Medicine; Inc. | 2 | Incorrect Absolute Marking Coordinates when used with PQ-Series scanners due to a software malfunction | All customers will receive an August 21, 2007, Customer Advisory Notification letter notifying them of the issue and the steps required to avoid the issue. Return receipts will be tracked. This is a short-term fix, a long term software fix will be issued shortly by the firm.
Z-0523-2007 | Feb-27-2007 | 2007 | Ultraview SL Command Module/Multiparameter input module - Ultraview SL Command Module with Masimo Sp02 Option; a multiparameter patient monitor. | Spacelabs Healthcare; Incorporated | 2 | Potential for modules to reset; alarms returning to default settings; invasive pressure losing its labels and zero calibration and Sp02 alarm limits being frozen. This can result in delay in treatment. | On 1/3/07 the firm issued a letter to their customers advising of the issues and provided guidance to the consignee to enable the Check Setup feature that will provide an audible prompt whenever the monitor resets. The letter advises that consignees will be contacted to schedule an update, once new software has been validated.
Z-0705-2007 | Apr-06-2007 | 2007 | ABL800 Series analyzer with FLEXQ module - ABL800 FLEX Series analyzer with FLEXQ module-equipped with software version 5.2 to 5.27 | Radiometer America Inc | 2 | In some situations pre-registered patient data is not reset when leaving the pre-registration mode in the software. This occurs when a Safe Pico blood sampler associated with a specific patient is pre-registered in the analyzer by scanning a barcode. | Radiometer America issued a recall letter to their U.S. customers dated 1/16/2007. The letter informs the customers of this software anomaly and the conditions underwhich analyzer results can be stored with incorrect Patient Demographics. The letter goes on to inform the customers that a Radiometer Service Representative will be visiting each customer location shortly in order to upgrade the affected software. In the meantime, the letter recommends that the customer(s) not pre-register their samples. The letter is accompanied by an attached FAX FORM which the customer is asked to sign and complete in order to verify receipt of the recall letter and to fax back to the recalling firm.
Z-0422-2007 | Feb-17-2007 | 2007 | iE33 Ultrasound System - iE33 Ultrasound System (System; imaging; pulsed doppler; ultrasonic and system; imaging; pulsed echo; ultrasonic) | Philips Ultrasound; Inc. | 2 | Acoustic output intensity may exceed maximum specified limits and the patient contact temperature may exceed 43 degree C (when the system is used with the S12-4 transducer in CW-mode) | On 11/2/06 the firm sent a letter dated October 31, 2006 to their customers. The letter explains the problem and states that new software is being released and will be available sometime in November. Customer service representatives will visit consignees to upgrade their units. The firm later determined they sent letters out to consignees that did not have affected units. Those consignees will be notified the week of 11/20/06 to disregard the previous letter.
Z-0507-2007 | Feb-15-2007 | 2007 | GE LightSpeed VCT System - The LightSpeed VCT System is intended for head and whole body X-ray Computed Tomography applications. | GE Healthcare | 2 | When the manual cradle release feature on the table is used to move a patient out of the gantry; a discrepancy of up to 6mm between the numerical display on the gantry and the actual position of the table could be introduced. | A Product Safety Advisory Notice, dated December 8, 2006, was issued to all customers informing them of the safety issue and production correction. Customers are advised to operate the CT cradle as intended with the system''s electronic controls. If a critical need arises to release the cradle during a radiation therapy simulation scan, they are advised to independently verify patient position prior to marking the patient for therapy to avoid the potential inaccuracy. A software update will be installed to correct this problem by a GE representative.
Z-0814-2007 | May-18-2007 | 2007 | BacT/ALERT Classic (Only) with BacT/VIEW¿ Software - BacT/VIEW¿ Software; all versions. in vitro diagnostic database software. | bioMerieux; Inc. | 2 | BacT/VIEW will be unable to run scheduled macro applications unless the user touches the BacT/VIEW touch screen or clicks a function after resynchronizing the clocks. | A customer notification was sent on/about December 19, 2006. Customers will be sent a new BacT/VIEW software version C.30b QTR 1 2007.
Z-1167-2007 | Aug-16-2007 | 2007 | VITROS¿ 5;1 FS Chemistry System - VITROS¿ 5;1 FS (Fusion Series) Chemistry System Software Version 2.2.1; only when using VITROS¿ Chemistry Products THC (Cannabinoids) or AMPH (Amphetamine) Reagents; Cat. No. 680 1375. | Ortho-Clinical Diagnostics | 2 | Software design. The firm received a customer complaint regarding a qualitative positive THC (Cannabinoids) result obtained from a THC quality control fluid that was known to be negative when using VITROS? Chemistry Products THC Reagent. A second customer complaint was received a month later. OCD's investigation concluded that under specific circumstances; incorrect qualitative results for Cannabi | On 5/17/07, OCD informed its foreign affiliates of the field correction by e-mail, and instructed them to inform their customers who received the VITROS 5,1 FS Chemistry Systems as well as the VITROS Chemistry Products THC (Cannabinoids) and/or AMPH (Amphetamine) Reagents of this issue. On 5/18/07, all US OCD direct customers and end-use customers of consignees that received the VITROS 5,1 FS Chemistry Systems as well as the VITROS Chemistry Products THC (Cannabinoids) and/or AMPH (Amphetamine) Reagents, were informed of the "field correction" by US Postal Service letter, provided with a new Assay Data Disk (ADD #5465), and were instructed to load it and recalibrate these assays, if necessary.
Z-0638-2007 | Mar-20-2007 | 2007 | Viasys Vela Adult/Pediatric Lung Ventilator - VIASYS VELA Ventilator; Catalog Number-16186; Viasys Respiratory Care; Inc. | Viasys Respiratory Care; Inc.dba Bird Products | 2 | The graphical user interface may become slow in its response and in some cases non-responsive. Under some circumstances this situation may cause the ventilator to delay the annunciation of alarms. | On January 31, 2007, a product notification letter was mailed to all customers in the United States via overnight carrier (Federal Express). Consignees are informed in the notification letter that service calls are in process of being scheduled to update domestic consignees. On January 31, 2007, a product notification kit (including a product notification letter, download software, download instruction sheets and installation verification fax-back sheets) were mailed to all international customers via register mail. Consignees are informed in the notification letter that replacement software, replacement instructions, and installation verification forms are included in their recall package. Domestic Consignees will have product corrected by VIASYS service personnel. Self address postage paid envelopes will be provided with notifications to aid in effectiveness checks.
Z-0925-2007 | Jun-09-2007 | 2007 | LIFEPAK 20 Defibrillator/monitor - LIFEPAK 20 defibrillator/monitor. The device is an AC powered defibrillator device with battery backup. The defibrillators are primarily used by healthcare providers in device user facilities such as hospitals and clinics. | Medtronic Emergency Response Systems; Inc. | 2 | White Screen Lock up: LifePak 20 defibrillators with software version 48 or version 52 may experience an intermittent timing issue during the power on sequence. This may lock up the device resulting in the monitor screen going white; and may cause a potential delay or prevention of shock treatment. | On 04/27/2007 firm sent letters via return receipt to all domestic and foreign customers. The letter recommends for customers to turn the defibrillator power off and back on to restore normal operation in the event of a white screen condition. In addition, the letter advises that a representative from the firm will contact customer to make arrangements to provide a software upgrade to correct the issue.
Z-0494-2007 | Feb-14-2007 | 2007 | Terumo APS 1 - Terumo Advanced Perfusion System 1; Air Bubble Detection Module; Catalog number 802110. | Terumo Cardiovascular Systems Corp | 2 | The air bubble detection system may falsely continue to sound the air alarm after the user has disabled it. | Consignees were notified of the problem via letter dated 11/10/06 and informed that they would be contacted to schedule a service call to install software upgrades. Installation of software upgrades began at customer locations on 10/16/06.
Z-0168-2008 | Nov-09-2007 | 2007 | MEVATRON M2/PRIMUS Mid - MEVATRON M2/PRIMUS Mid;Medical charge-particle radiation therapy system (Medical Digital Linear Accelerator);MODEL NUMBER: 1940035; MANUFACTURED BY: Siemens Medical Solutions USA; Inc; 4040 Nelson Ave; Concord; CA 94520 | Siemens Medical Solutions USA; Inc | 2 | Mistreatment of radiation therapy due to inadvertent parameter changes or miss-calibration in Service Mode. The miss-calibration can affect the beam profile and Virtual Wedge angle. This leads to patient mistreatment and alteration of future treatment plans. A second reason is an unplanned Dosimetry calibration data set can be erroneously copied during "Service Mode." | Distribution of Update Instructions TH007/07/S was sent to all affected sites on March 26, 2007. This was intended to inform users that Service Mode is only for use by authorized personnel, and that unauthorized and/or incorrect softpot parameter changes performed in Service Mode can result in the risk of patient mistreatment. Current control system software versions do not allow unplanned Dosimetry calibration data sets to be erroneously copied during "Service Mode.".
Z-0436-2007 | Feb-14-2007 | 2007 | Terumo APS 1 - Terumo Advanced Perfusion System 1; Central Control Monitor; Catalog number 802100. | Terumo Cardiovascular Systems Corp | 2 | Pump may stop without an audible tone as a result of quickly clearing the internal pump alarm condition or overpressure alert/alarm condition. | Consignees were notified of the problem via letter dated 11/10/06 and informed that they would be contacted to schedule a service call to install software upgrades. Installation of software upgrades began at customer locations on 10/16/06.
Z-0987-2007 | Jul-03-2007 | 2007 | HD11 - HD11XE Diagnostic Ultrasound System | Philips Medical Systems | 2 | Potential for erroneous measurements which may lead to incorrect diagnosis based upon the presence of an inaccurate; undetectable error in clinical measurement or calculation. If the user exits an imaging mode while Active Annotation is in use; the system can enter an undefined state where the image scale calibration may not match the scale of the on-screen measurements and graphics and possib | On 04/25/2007 firm sent letters dated 04/23/2007 to consignees informing them of the issue and that a new software release to correct the issue became available 04/19/2007. The firm's Field Service Representatives will contact consignees to install the new software at no charge to the consignee.
Z-0448-2007 | Jan-09-2007 | 2007 | Bayer Advia 70 Hematology Analyzer - Bayer Advia 70 Hematology Analyzer; manufactured by Drew Scientific; Inc.; Dallas; TX. | Drew Scientific Group PLC | 2 | Flags and alerts on the user interface will not automatically transfering to the LIS. | Firm sent 'Product Notification #22' to all customers on 08/24/06. Users advised of potential risk to patient health and provided with instructions on how to eliminate risk by disabling the auto transfer feature and manually transfering data until a software patch is made available.
Z-0113-2008 | Nov-07-2007 | 2007 | GE VOLUSON E8 ULTRASOUND SYSTEM - GE Healthcare Voluson E8 ultrasound scanner with Software version 6.2.0 and 6.2.1 and DICOM stations with these software versions; GE Healthcare 9900 West Innovation Drive; Wauwatosa; WI 53226; USA. | Ge Healthcare | 2 | When a spectrum of a Pulsed Wave Doppler shows an angle correction unequal to 0 is reloaded from the archive or the image is sent to a DICOM station; this angle correction is mistakenly disregarded resulting in an underestimation of the flow velocities. | GE Healthcare sent customer letters on July 15, 2007 to affected customers. The letter included a short term mitigation recommendation and a long-term correction. The long-term correction will be software upgrade provided free of charge to all affected customers.
Z-0180-2008 | Nov-21-2007 | 2007 | VoxelQ Workstation; Imaging Workstation - Philips; Voxel Q workstations; untilizing AcQSim or AcQPlan software version 4.2; 4.2.1; 4.9; 4.9.1; 4.9.2; 4.9.3; 5.0; 5.0.1; 5.0.2 or 5.0.3 and any Mx8000; Mx8001DT; or Brilliance 6; 10; 16; 16P; and Big Bore CT systems connected to the VoxelQ; Philips Medical Systems; Cleveland; OH 44143 | Philips Medical Systems (Cleveland) Inc | 2 | Incorrect simulation and/or beam placement: An anomaly has been identified with the Philips VoxelQ workstation utlizing AcQPlan software version 4.2; 4.2.1; 4.9; 4.9.1; 4.9.2; 4.9.3; 5.0; 5.0.1; 5.0.2 or 5.0.3 when computing absolute patient marking coordinates with data containing reconstruction offsets from the Mx800; the Mx8001DT; and the Brilliance 6; 10; 16; 16P; and Big Bore CT systems. Whe | On November 1, 2007, after telcon with FDA, the firm revised its recall strategy to notify consignees via a November 9, 2007, Urgent Device Notification letter which indicates that warning labels will be affixed to the units that alerts the user of the reconstruction offset defect by Philips employees.
Z-0432-2007 | Feb-14-2007 | 2007 | Terumo APS 1 - Terumo Advanced Perfusion System 1; 100V -120V; 15 A (circuit breaker); 50/60 Hz (20A power source required) base; Model 801763. | Terumo Cardiovascular Systems Corp | 2 | A 'battery needs service' message may be displayed; or the remaining battery time may incorrectly be reported out as zero; even when the battery is fully functional. | Consignees were sent a letter dated 11/10/06 informing them of the nature of the issues and that the firm would contact them to schedule a service call to install the software upgrade package. The software upgrade installation began at customers on 10/16/06, prior to issuance of the letter to consignees.
Z-0329-2008 | Dec-07-2007 | 2007 | Allura Xper - Allura Xper FD20 X-Ray System; Release 2.2; Philips Medical Systems; Bothell; WA 98021 | Philips Medical Systems North America Co. Phillips | 2 | Inaccurate anatomical measurements when using the optional SmartMask feature | The recalling firm sent customers an Urgent Device Correction notification letter dated 09/21/07 via certified mail - return receipt requested. This notification provided a short term work around for the device defect. For a permanent fix, the recalling firm is developing software update to check for mismatch at the start of a subtracted fluoroscopy run, and will contact customers to schedule an install. Installation is expected to begin 10/2007 and be completed 06/2008.
Z-0784-2007 | May-10-2007 | 2007 | Engstrom Carestation - GE Healthcare Engstrom Carestation Critical Care Ventilator with the High Performance Display Unit with SW version 3.x and 4.11; identification numbers 1505-9000-000 and M1057699 (upgrade). Datex-Ohmeda; Inc.; a General Electric Company; doing business as GE Healthcare; P.O. Box 7750; Madison; WI 53707-7550; USA; www.gehealthcare.com. | Datex - Ohmeda; Inc | 2 | There are two unique conditions with the Engstrom Carestation that require corrective action. These include an anomaly in the control of the safety relief valve and a situation where the display unit may reset. | Device Correction Letter to customers was mailed on 2/2/07 giving information about the correction and a short term recommendation until version 4.15 could be installed by a GE representative who would be contacting the customer.
Z-1244-2007 | Sep-26-2007 | 2007 | daVinci Surgical System Model IS1200; A4.3 SW level - Endoscopic Instrument Control System (da Vinci Surgical System) Model IS1200; A4.3 SW level; Manufactured by Intuitive Surgical; Inc; Sunnyvale; CA 94086 | Intuitive Surgical; Inc. | 2 | Under certain circumstances; the product's software may crash and require a manual override or restart before functioning again. In addition; there is an unrelated addendum to the User Manual to clarify an icon not described in the previous User Manual. | As of 9/7/2007, Intuitive Field Service Engineers have performed the upgrade via service visit on 292 of 293 U.S. systems listed on the affected sites list. A notification letter that the upgrade has been performed was left with each site after the upgrade has been completed. The firm's distributors received advance notice of the upgrades 9/13/2007.
Z-0418-2007 | Feb-08-2007 | 2007 | Cytomics FC 500 - Cytomics FC 500 with CXP software; Versions 2.0 & 2.1: Part #722645; Software Kit; CXP version 2.1 New User; Part #723113 Software Kit; CXP version 2.1 Upgrade; Part #629637 Software Kit; CXP version 2.0 New User; Part #629636 Software Kit; CXP version 2.0 Upgrade. The Cytomics FC 500 is a system for the qualitative and quantitative measurement of biological and physical properties of cells and other particles. | Beckman Coulter Inc | 2 | Two workflow scenarios associated with renaming regions may produce incorrect results. | A Product Corrective Action (PCA) letter was sent on the week of August 29, 2006 to all accounts that have an FC500 or have ordered CXP or MXP software to inform them that a) incorrect statistics are generated when a polygonal region with vertical re-entrant vertices is renamed using the Region Properties dialog box. b)when a region is copied from one plot to another, and then the new region is renamed, the statistics are not updated with the new region name. (Letter was sent by US mail). Letter includes immediate user preventive action to remedy both issues. A response form is included for consignees to return to recalling firm.
Z-0555-2007 | Sep-26-2007 | 2007 | Pyrxis Anesthesia System 3500 - Pyxis Anesthesia System 3500; Cardinal Health; Clinical Technologies and Services; San Diego; CA | Cardinal Health 303 Inc DBA Alaris Products | 2 | System lock up; mini-drawers -under certain circumstances during use; are not accessable; delaying the user from accessing critical medications. | The field action consisted of phone calls and a mailing conducted to all customers for the Pyxis Anesthesia System 3500 beginning on 08/13/2007. The customers will be notified by registered return receipt mail by Recall Notifications dated August 14, 2007. Separate letters will be sent to the Director of Pharmacy, Chief of Anesthesia, Director of Risk Management, and Director of Surgery at each facility. Recommended actions will be provided as a short-term fix and a long-term fix of the software will be issued shortly.
Z-0985-2007 | Jun-27-2007 | 2007 | Aplio - Toshiba Aplio 80 Digital Ultrasound System; SSA-770A with version 6 software | Toshiba American Med Sys Inc | 2 | When using the LV measurement function and making numerous measurements the value displayed on the image screen may on occasion not update correctly. The values that are stored in MDA (measurement display area) are correct and no issue has been reported with these values. | Notification letters were sent to customers on February 2, 2007. Customers were informed of the safety issue regarding the possibility of incorrect measurement values being displayed on the active image screen while performing multiple measurements and instructed to verify actual results against the values stored in the MDA (measurement display area). Customers were also informed that a Toshiba Customer Service Engineer will contact them to schedule a software upgrade when software becomes available. If customers have any questions they were instructed to contact their Toshiba Customer Service Engineer or call (800) 421-1968 and ask for the Regulatory Affairs Department. Customers were also instructed to distribute this information to all users of the product and also with the Bio-medical Department at their facility.
Z-0397-2007 | Feb-01-2007 | 2007 | Portrait¿ - Portrait¿ PSR3 System; an electro-surgical device used in dermatological applications | Rhytec Inc. | 2 | Inadequate Directions for Use: Software Upgrade to set maximum energy level at 2.5 Hz | Rhytec notified users by letter dated 10/2/06 of the software upgrade. A service representative will perform the update.
Z-0966-2007 | Jul-04-2007 | 2007 | Invacare Storm TDX Power Wheelchair - Invacare Storm TDX Power Wheelchair; Models: #TDX3-PS; TDX3-SE-PS; TDX4-PS. | Invacare Corporation | 2 | Software error may permit the chair to recline too far. | The recalling firm sent a recall letter dated March 19, 2007. The recalling firm also posted information regarding the recall on their website.
Z-0433-2007 | Feb-14-2007 | 2007 | Terumo APS 1 - Terumo Advanced Perfusion System 1; 220V -240V; 7A (circuit breaker); 50/60 Hz (10A power source required) base; Model 801764. (Not distributed within the United States). | Terumo Cardiovascular Systems Corp | 2 | A 'battery needs service' message may be displayed; or the remaining battery time may incorrectly become zero; even when the battery is fully functional. | Consignees were sent a letter dated 11/10/06 informing them of the nature of the issues and that the firm would contact them to schedule a service call to install the software upgrade package. The software upgrade installation began at customers on 10/16/06, prior to issuance of the letter to consignees.
Z-0121-2008 | Nov-07-2007 | 2007 | Brilliance CT Big Bore - The Brilliance CT Big Bore is a whole body computed tomography X-Ray system. The malfunctioning software application is the Tumor Localization (Tumor LOC) application. | Philips Medical Systems (Cleveland) Inc | 2 | The maximum intensity projection (MIP) images produced from respiratory gated images by the Tumor Localization (Tumor LOC) software application on Brilliance Big Bore CT scanners will be labeled incorrectly for certain patient orientations and view conventions when saved for later viewing. | Notification letters were sent to the installed customer base informing them of the hazard on 10/1/2007. A mandatory corrective software release due out mid October 2007, will be available at no cost to the customer. The software will be installed by Philips Field Service Engineers.
Z-1043-2007 | Jul-10-2007 | 2007 | GE Healthcare Innova 2100 IQ - GE Healthcare Innova 2100 IQ Digital Fluoroscopic Imaging System (Versatile cardiac and vascular). | General Electric Med Systems LLC | 2 | Caution on the non-detection of small invasive objects/devices; which project out from the skin surfaces (e.g. biopsy or other needles) while using the Innova Sense software option designed for patient contouring. Usage of Innova Sense patient contouring for such devices could potentially lead to patient injury. | A "GE Healthcare Urgent: Safety Notice" dated January 24, 2007 was sent to consignees describing the product, problem, safety instructions, product correction and contact information.
Z-0373-2007 | Jan-25-2007 | 2007 | GE Centricity PACS RA1000 Workstation - GE Centricity PACS RA1000 Workstation; for diagnostic image analysis. GE Healthcare Integrated IT Solutions; Barrington; IL 60010 | GE Healthcare Integrated IT Solutions | 2 | The Centricity PACS RA1000 Workstation Software may manifest certain mammography image display problems only if the images were acquired through GE Senographe versions 2000D; DS or Essential. | GE Healthcare IITS sent recall letters dated 1/12/07 to all customers who have the combination of GE Centricity PACS RA1000 Workstation software versions 2.1.X.X, or 3.0.X.X combined with the GE Senographe 2000D, the Senographe DS, or Senographe Essential mammography scanners, informing them that the entire breast image or a portion of the breast image may not be displayed in the monitor regions and may not be evident to the end-user. To mitigate this problem, the users were provided a Quick Reference Guide detailing how to use the Pan tool to move the image left, right, up and down in order to display any portion of the image that may not be displayed on the screen. The customers were requested to make sure that all necessary physicians and personnel at their location are made aware of the issues and understand the risk mitigation provided. A software patch to permanently resolve the issue will be installed at no charge when it becomes available. Any questions were directed to the GE Customer Care Center at 800-437-1171.
Z-0905-2007 | Jun-06-2007 | 2007 | FramelessArray software - FramelessArray software; Version 1.0. ; Medical Device firmware incorporated with the RadioCamera Extracranial System; distributed by Varian Medical System; Palo Alto; CA. | Varian Medical Systems Oncology Systems | 2 | Localization error; with planning data transferred to the Optical Guidance Platform via DICOM RT; The software incorrectly computes the center of the CT volume; resulting in a potential axial error ranging from 0.3 mm to 1.5 mm; affecting both Fractionated and Stereotactic Radiosurgery (SRS) Treatments | The recalling firm notified consignees by certified letter on 03/14/07. The notification was flagged as "Urgent Medical Device Correction" and advised end users that a localization error may occur when using the FramelessArray and affect its accuracy for Frameless treatments. Users were instructed to verify software in the devices and if running FramelessArray version 1.0, to discontinue use of the software when Stereotactic Radiosurgery plans are transfered via DICOM RT. The notification offered that firmware would be upgraded during scheduled service visit.
Z-0271-2008 | Nov-29-2007 | 2007 | ECAT EXACT 47 - Siemens ECAT EXACT 47; Catalog No. 1788772; PET/CT Scanner; Manufactured by Siemens Medical Solutions USA; Inc.; 810 Innovation Drive; Knoxville; TN 37932 USA; TEL: 865-218-2534 | Siemens Medical Solutions USA Inc | 2 | Incorrect Decay Correction Results - ECAT System's software versions 7.2.2 and 7.4 calculate the decay correction for emission data incorrectly; while performing whole-body studies with all transmissions followed by all emissions. | Siemans sent an Urgent Device Correction letter, dated June 15, 2007, to all affected consignees informing the user of the problem and requesting that they discontinue the practice of running all transmissions followed by all emissions until a corrective software patch is installed. The corrective software patch is expected to be installed by October 2007.
Z-1201-2007 | Aug-29-2007 | 2007 | ABL8O FLEX - ABL8O FLEX Software Version 1.10 included in the ABL8O Flex Analyzer; Model Number 914863 | Sendx Medical Inc | 2 | The ABL8O FLEX software version 1.10 contains errors in the derived parameter calculation for cBase(B;ox) and cBase(Ecf;ox). These incorrect calculations result in a mild to moderate lowering of the base excess value; the severity depending on the concentration of hemoglobin and the oxygen saturation in the sample. Elevated hemoglobin concentrations have the greatest impact on this calculation err | A total of eight (8) distributors and the Radiometer main office in Denmark were sent the recall package (Field Action Notice) by email on 6-6-07. The distributors were provided background information and specific instruction on the steps to perform to correct the problem. The Recall package included a Management Cover for Field Action Notes and the Field Action Notice. Radiometer has released a subsequent software upgrade (v.1.11) to correct the two problems noted in the Recall Notice.
Z-1146-2007 | Aug-08-2007 | 2007 | ISOLOC - ISOLOC 6.5 Software Part Number NT-NW-425-520. Software for radiotherapy treatment. | Northwest Medical Physics Equipment Inc | 2 | This correction was for several software bugs which caused image format problems; incorrect ruler display; software crashes; failure to update pixel spacing; display of unattainable treatment table positions; failure to update gantry angles; wrong shifts for two of three axes of treatment table movement. | Software to fix the problems was written and software users were notified of the availability by e-mail on 04/07/2006. They could download the fixed software from the firm's internet website.
Z-1145-2007 | Aug-08-2007 | 2007 | Image guided radiation treatment software. - Isoloc software Versions 6.5 and 4.5; Part Number NT-NW-425-520. Software for image guided radiation therapy. | Northwest Medical Physics Equipment Inc | 2 | Isoloc version 6.5 and 4.5 software issued a false warning message regarding Gantry angles when set to "Anatomic Landmark" setting. | Software to fix the described problem was written and customers contacted by e-mail on 01/06/2006 that a software patch, titled 01022006 was available to their customers on the firm's internet website.
Z-0273-2008 | Nov-29-2007 | 2007 | ECAT EXACT HR+ - Siemens ECAT EXACT HR+; Catalog No. 3548083; PET/CT Scanner; Manufactured by Siemens Medical Solutions USA; Inc.; 810 Innovation Drive; Knoxville; TN 37932 USA; TEL: 865-218-2534 | Siemens Medical Solutions USA Inc | 2 | Incorrect Data Results - ECAT System's software versions 7.2.2 and 7.4 may calculate the decay correction for emission data incorrectly while performing whole-body studies with all transmissions followed by all emissions. | Siemans sent an Urgent Device Correction letter, dated June 15, 2007, to all affected consignees informing the user of the problem and requesting that they discontinue the practice of running all transmissions followed by all emissions until a corrective software patch is installed. The corrective software patch is expected to be installed by October 2007.
Z-0704-2007 | Mar-30-2007 | 2007 | InstaTrak 3500 Plus system - InstaTrak 3500 Plus system with cranial application (aid to the surgeon for locating anatomical structures in the human body) | GE OEC Medical Systems;Inc | 2 | Tracking inaccuracy of 5-10 mm may result when the shunt guide accessory is used with version 5.1 software. | Consignees were visited by GE personnel to upgrade the software, beginning in April, 2004, to version 5.2.
Z-0545-2007 | Mar-01-2007 | 2007 | SMS - Sample Management System software for in vitro diagnostic testingProduct # 030102-03 | Siemens Medical Solutions Diagnostics | 2 | Under limited circumstances; an incorrect patient result could be printed on the optional chartable patient report. | A recall communication was sent on 1/15/2007 to all customers by fax.
Z-1147-2007 | Aug-08-2007 | 2007 | ISOLOC - ISOLOC 4.2 Software Part Number NT-NW-425-520. Image guided radiation treatment software. | Northwest Medical Physics Equipment Inc | 2 | When the ISOLOC (Version 4.2) profile distance unit is set to "centimeter" and the user clicked the "back" button in the digitizer window; the marker and target coordinates in the localization window erroneously display values ten times greater than the correct values. | Software to fix the described problems was written and customers contacted by e-mail on 05/31/2005, that a software patch, titled 05312005 was available on the firm's internet website.
Z-1153-2007 | Aug-08-2007 | 2007 | ISOLOC - ISOLOC Software Version 6.0; Localization software for determining treatment machine position (to position patients) for high accuracy radiotherapy. | Northwest Medical Physics Equipment Inc | 2 | When ISOLOC version 6.0 was used in the centimeter mode; the display in the couch moves window had a rounding error relative to the Portal Image Marker window of up to 0.5 mm. | Software to fix the described problem was written and customers contacted by e-mail on 04/11/2005, that a software patch, titled 04112005, was available to their customers on the firm's internet website.
Z-1152-2007 | Aug-08-2007 | 2007 | ISOLOC - ISOLOC Software Versions 6.0 and 4.2. Part Number NT-NW-425-520. Image guided radiotherapy treatment software. | Northwest Medical Physics Equipment Inc | 2 | When ISOLOC version 6.0 and 4.2 distance unit is set to centimeter; the localization coordinates were rounded to the nearest millimeter which could affect positioning by 0.5 mm. | Software to fix the described problems was written and customers contacted by e-mail on 04/28/2005, that a software patch, titled 04282005, was available on the firm's internet website.
Z-1155-2007 | Aug-09-2007 | 2007 | ISOLOC - ISOLOC Software Version 4.2; Part Number NT-NW-425-520. Image guided radiation treatment software. | Northwest Medical Physics Equipment Inc | 2 | When ISOLOC Version 4.2 unit is set to centimeter and the user clicks the Reenter button in the digitizer window; the coordinates of the markers and target in the anatomic imaging system erroneously displays values ten times greater than the correct values. | Software to fix the described problem was written and customers contacted by e-mail on 11/30/2004, that a software patch, titled 11302004, was available to their customers on the firm's internet website.
Z-0978-2007 | Jun-28-2007 | 2007 | LINAC aka FastPlan Treatment Planning Sysytem - LINAC Scapel Treatment Planning System; incorporating FastPlan software versions 4.0; 5.0; 5.0.1; 5.1; 5.1.1. | Varian Medical Systems Inc | 2 | A software anomaly in a medical device used for patient treatment could result in accidental or unintended exposure radiation exposure. The tilt rotation compensation algorithm does not adequately compensate for the tilt rotation. | The recalling firm notified medical device end users by letter dated 03/09/07 and advised of the software anomaly and provided details for user corrective action pending software upgrade by the manufacturer.
Z-1154-2007 | Aug-08-2007 | 2007 | ISOLOC - ISOLOC Software Version 6.0; Part Number NT-NW-425-520. Localization software for radiotherapy treatment system. | Northwest Medical Physics Equipment Inc | 2 | When ISOLOC version 6.0 software was used in the Anatomic Landmark mode; the couch moves display in the Couch Translator window was incorrect. | Software to fix the described problem was written and customers contacted by e-mail on 12/30/2004 that a software patch, titled 12302004 was available to their customers on the firm's internet website.
Z-1182-2007 | Sep-05-2007 | 2007 | Philips Intellivue Central Station Software - Intellivue Central Station Software for the: Intellivue Information Center; Philips Information Center; Central StationModels: M3145; M3150; M3154; M3155; M3169 | Philips Medical Systems | 2 | Incorrect Patient Assignment: the information center will in most cases be associated with the "incorrect" patient. This occurs on both standalone and networked systems running fixed mode monitoring. (This does not occur on systems that run flex mode monitoring) | Philips Medical Systems issued a letter notification dated July 2007 " Urgent Device Correction". Customers with networked information centers will be corrected through a database server. Stand-Alone units will have a software correction. Software version J.00.25 will correct the problem. Users are advised not to use th 12 Lead Analysis and Export Option in fixed or flex mode monitoring unitl the software correction is made.
Z-1174-2007 | Aug-22-2007 | 2007 | EasyLink Informatics System - EasyLink Informatics System; software version 3.0; Part Number 1000034941; DADE BEHRING INC.; Newark; DE. | Dade Behring; Inc. | 2 | Sample ID and Patient results could be associated with the wrong patient | The recalling firm issued a Field Corrective Action letter dated 7/3/07 to inform their customers of the problem and that a software patch would be sent for them to download. The recalling firm plans to telephone the customers after the patch has been released to verify that it has been installed.
Z-0652-2007 | Mar-28-2007 | 2007 | syngo Dynamics - Siemens syngo Dynamics 5.0 Workplace. Medical image report and archive system. | Siemens Medical Solutions; USA; Inc | 2 | The cardiac calculation feature may incorrectly calculate a derived V max value (reported out lower than it should be) and certain values imported into worksheets and reports may be incorrect due to a software defect. | Distributors were issued a recall letter dated 1/29/07 and asked to notify their customers by issuing their subaccounts a similar letter. When available, new software will be provided.
Z-0869-2007 | Jun-05-2007 | 2007 | GE Healthcare Innova 4100/4100 IQ - GE Healthcare Innova 4100/4100 IQ. Digital Fluoroscopic Imaging System (Vascular Angiography System) | General Electric Med Systems LLC | 2 | Measurement errors during the computation of vessel diameters while using the Stenosis Analysis software (also known as Quantitative Coronary Analysis function-QCA) using the auto-calibration function on Innova Digital (located in control room) or in-room InnovaCentral/Touch-Screen | A letter dated Jan. 18, 2007, was sent to consignees describing the problem, product ,safety instructions and contact information. Software upgrades were promised for correction.
Z-0181-2008 | Nov-10-2007 | 2007 | InstaTrak 3500 Plus - GE OEC; InstaTrak 3500 Plus FluoroTrak and Fluoro CAT Spinal Applications used with Regular and Super 9800 and 9600 Arm Systems. GE Health Medical Systems; Inc; Salt Lake City; UT 84116-2862 | GE OEC Medical Systems; Inc | 2 | Tracking Accuracy - Image intensifier supplier's change on the C-Arm could lead to patient injury; due to incorrect anatomical location of the instrument during a procedure. | Consignees were sent an Urgent Recall Notice letter, via mail dated 10/11/07, advising users to stop usage of the FluoroTrak and FluoroCAT spinal applications on the InstaTrak 3500 Plus System. Once a solution is identified, a representative will contact consignees to arrange for installation of the solution at no charge.
Z-0451-2008 | Dec-11-2007 | 2007 | Philips Medical Systems IntelliVue Information Center - Philips Medical Systems; M3290A IntelliVue Information Center Software for M3170 Intellivue Patient Link; Catalog Number: 865007 and 865015; Running Software Versions J.00.23; J.00.24; J.00.25; and J.00.26; Manufactured and Distributed by; Philips Medical Systems; Andover; MA 01810 | Philips Medical Systems | 2 | Alarm failure : If changes are made to the Caregroup Editor configuration; Care Group alarm status; including alarm text and prompt tones; may fail to be displayed or annunciated at bedside monitors within the Care Group | Philips Medical notified customers by letter dated November 2007 titled : Urgent Medical Device Correction Notice. Customers will be instructed to follow the Procedure to Mitigate Risk section * of the notice while they await the correction of their devices with a mandatory software upgrade to be performed by service representatives. * (no changes are to be made to the Caregroup Editor configuration on any system. If changes have already been made to the configuration prior to receipt of this notification, turn each bedside monitor "off " and then "on" in order to reset the system-Customer Care Contact telephone 1-800-722-9377).
Z-0148-2008 | Dec-28-2007 | 2007 | Eye Cubed Ultrasound System - Ellex; Eye Cubed; I 3 System-ABD Ultrasound Version 3.0 with Software Version 3.0 and Biometry A-scan modality; Ellex Innovative Imaging; Sacramento; CA 95827 | Ellex Inc. | 2 | Incorrect IOL calculations: software behavior with Version 3.0 Eye Cubed Ultrasound systems running Biometry A-scan enabled Version 3.0.0 software can cause incorrect IOL calculations. | A Global Customer Service Bulletin (GCSB)-Urgent Notification of Software Correction was issued on 25 September 2007, which provided short term workaround directions pending a free software undate promised in the near future.
Z-0616-2007 | Mar-13-2007 | 2007 | iSite PACS - The iSight PACS (picture archiving and communications system) a software package used with general purpose computer hardware to acquire; store; distribute; process; and display images/associated data throughout a clinical environment. | Stentor Inc A Phillips Medical Systems Co. | 2 | A defect may cause patient image orientation markers to be incorrectly labeled on Multi-Planar Reformation (MPR) images. Left and Right Anatomical orientation markers may be reversed. | The firm has issued a notification letter to its consignees on 02/14/2007. This notification letter contains preventive measures.
Z-1075-2007 | Jul-12-2007 | 2007 | AirLife nCPAP System Driver - AirLife nCPAP System Driver; a nasal continuous positive airway pressure device; which is part of a larger system that includes an nCPAP generator; fixation device and patient interface (prongs and.or mask); Cardinal Health; McGaw Park; IL 60085; cat. 006900 | Cardinal Health | 2 | Under isolated conditions; the device may experience issues with O2 Calibration; which may result in either a nuisance alarm or a delivered FiO2 of up to 5% below the specified tolerance. | Cardinal Health sent the accounts recall letters dated 6/1/07 via UPS next day delivery, which listed the catalog number and affected serial numbers of the CPAP Driver, and informed them that under certain isolated conditions, the CPAP Driver may experience issues with O2 Calibration, which may result in either a nuisance alarm or a delivered FiO2 of up to 5% below the specified tolerance. Customers have experienced this under the following conditions: 1) air inlet pressure is greater than O2 pressure by 5 to 10 psi and 2) FiO2 settings less than 45%. Separate recall letters were sent to those accounts who had filed complaints of O2 Calibration Errors, stating that their device will be replaced by the end of June, 2007; those accounts who purchased the devices, but had not reported O2 Calibration Errors, stating that a Cardinal Health representative will upgrade the firmware on the devices at their facility by the end of August, 2007; and those accounts who had had demonstration units in their possession, advising them to notify their sales representative if any of the conditions mentioned in the letter occurred. The accounts were requested to complete and return via fax at 847-689-9101, the enclosed acknowledgment form, indicating the number of affected drivers in their possession, and were advised that they would be contacted by a Cardinal Health representative to arrange for the replacement/service of their drivers. Any questions were directed to Quality Systems-Professional Services at 1-800-292-9332 or 1-800-327-7503.
Z-0771-2007 | May-05-2007 | 2007 | Nemio XG SSA-580A; Diagnostic Ultrasound System - Nemio XG; ultrasound system; SSA-590A with version 1.5 software; Toshiba American Medical Systems; Inc.; Tusin; CA. | Toshiba American Med Sys Inc | 2 | Transducer Surface Temperature : It was found during testing that when using the system in the 2DCDI + PW mode; with the Scale Graph function active; that the transducer surface temperature could exceed the 27 degrees Celsius limit as specified in the IEC safety standard. | Notification letters were sent to customers on December 19, 2206. Customers were instructed to avoid using Scale Graph function with live images, and instructed to call up the Scale Graph after images are frozen. Customers were informed that a Toshiba Customer Service Engineer will contact them to schedule a software upgrade when software becomes available. If customers have any questions they were instructed to contact their Toshiba Customer Service Engineer or call (800) 421-1968 and ask for the Regulatory Affairs Department. Customers were also instructed to distribute this information to all users of the product and also with the Bio-medical Department at their facility.
Z-0381-2007 | Feb-01-2007 | 2007 | Terumo APS 1 - Terumo Advanced Perfusion System 1 Integrated Centrifugal System Control Unit; Catalog number 801046. | Terumo Cardiovascular Systems Corp | 2 | Following installation of software version 1.30 in October 2006; the firm confirmed a complaint that the new software caused a discrepancy in the pump speed calculation and caused the pump to stop during clinical use and a ''connect motor'' message to be displayed. | Customers were notified by phone, fax, email or by a Terumo representative of these problems between 11/08/06 and 11/13/06 and instructed to not use the unit until new software had been installed. All version 1.30 software was replaced and version 1.40 installed prior to 12/4/06.
Z-0815-2007 | May-18-2007 | 2007 | EXCITE - GE Healthcare-Americas Signa 3.0T EXCITE Whole Body Magnetic Resonance System | GE Medical Systems LLC | 2 | To provide the 3.0T MRI systems users with the 8-channel brain coil the proper coil weighting software. The proper coil weighting factors for the eight-channel brain coil were not included in the software. | The corrective action is already completed and all customers have been notified via a personal visit to each site, to install a software patch beginning June, 2006.
Z-0364-2008 | Dec-12-2007 | 2007 | CyberKnife Robotic Radiosurgery System - Accuray; Multiplan Treatment Planning Software; Versions 2.0; 2.0.1 .6; 1.6.1; 1.6..1 or 1.6.3.; Accuray; Sunnyvale; CA 94089. | Accuray Inc | 2 | Software problem - Sagittal; oblique or coronal sliced MR Images imported from a non-Accuray manufactured device into the product may result in inverted image orientation. | The firm issued an Urgent Advisory Notification on March 20, 2007, notifying users to only import axial sliced MR image series for treatment planning. Users should not import sagittal, oblique or coronal images series until further notice. The firm intends to issue a software patch to correct the problem and will contact user regarding resolution. Users can contact Accuray Customer Support at 1-877-668-8667 (USA) or 1-408-716-4700 (non-USA) or customersupport@accuray.com.
Z-0293-2008 | Nov-27-2007 | 2007 | Digital Radiography System with Flat Panel Detector - Toshiba-Interventional Angiography System; Infinix i; Model DFP-8000D; equiped with Digital Radiography System with Software version 3.4er000; Toshiba America Medical Systems; Inc.; Tustine; CA 92780 | Toshiba American Med Sys Inc | 2 | System Lock-up: When performing DAS One Shot; the system may lockup due to a software timing error. | Notification letters (Urgent: Medical Device Correction-dated October 1, 2007) with an attached Customer Response form were sent to customers beginning on October 2, 2007. Customers were notified this action was being taken as a result of an investigation that resulted from a software issue that was found done by the manufacturer. Customers were informed that a Toshiba Customer Service Engineer would contact them to schedule a software upgrade when the software becomes available. If customers have any questions they were instructed to contact their or call (800) 421-1968 and ask for the Regulatory Affairs Department.
Z-0989-2007 | Jul-12-2007 | 2007 | Mobile fluoroscopic x-ray unit. - GEHC OEC 9900 ESP-15 Elite C-Arm fluoroscopic X-ray system; GE Healthcare; Salt Lake City; UT. | GE OEC Medical Systems; Inc | 2 | The Air Kerma Rate displayed value and cumulative air kerma displayed value could exceed the allowable error under certain imaging conditions. | Consignees were notified by certified letter on 07/06/2007. They were told that a service rep would contact them to provide a label and an addendum to their operator's manual. A permanent software solution is under development and will be provided when it becomes available.
Z-0014-2007 | Sep-26-2007 | 2007 | Brilliance CT - Brilliance(a) CT 6 slice; Model #4535 670 73191. Brilliance(a) CT 10 Slice; Model 4535 670 73181. Brilliance 16 Slice; Model #4535 670 78851. Brilliance CT 16 Power; Model #4535 670 23331. Brilliance CT 40 Slice; Model #4550 110 04011. Brilliance CT Big Bore; Model #4535 670 88051. Computed tomography systems. | Philips Medical Systems (Cleveland) Inc | 2 | A potential anomaly was identified on the CT subsytem; which could result in the radiation dose being different than the calculated dose reported to the operator. | The customers were sent Public Safety Notification November 3, 2006, indicating there was a problem with the software. However, since their unit was recently upgraded the problem did not impact their unit.
Z-0636-2007 | Mar-17-2007 | 2007 | Brilliance CT System; Big Bore Configuration. - Brilliance CT System; Big Bore Configuration; Model #728243. (version 2.2.1) | Philips Medical Systems (Cleveland) Inc | 2 | Two conditions have been identified with the Philips Medical System Brilliance Bore ;version 2.2.1 Tumor Localization; that may lead the user to misinterpret the displayed data. | Customer advisory notification letters dated 10/4/06 and 10/25/06 were sent to consignees providing short term work around instructions. Software modifications installed by the recalling firm will follow as the long term correction.
Z-0548-2007 | Feb-23-2007 | 2007 | Symbia S Series SPECT System with a Pinhole Collimator - Symbia S Series SPECT (Single-Photon Emission Computed Tomography) System with a Pinhole Collimator; Siemens Medical Solutions USA; Inc.; Molecular Imaging; 2501 N. Barrington Road; Hoffman Estates; IL 60195-5203 USA; part number 08717741 | Siemens Medical Solutions USA; Inc. | 2 | Symbia systems; running on e.soft version 5.5 or earlier; could allow a fully-extended patient pallet to come into contact with the pinhole collimator; resulting in system damage. | Siemens sent Customer Safety Advisory Notice/Urgent Field Correction Recall letters dated 1/30/07 to all users of Symbia systems with a pinhole collimator and e.soft software versions 5.5 or earlier. The accounts were informed of the potential for a fully-extended patient pallet to come into contact with the pinhole collimator, resulting in system damage. The customers were advised that a Siemens Customer Service Representative would be sent to their site to upgrade their system software to version syngeo MI Applications 2006A within 30 days. The upgrade will be done free of charge.
Z-0213-2008 | Nov-08-2007 | 2007 | Varian Ximatron C-Series Digital Imaging - Ximatron Digital Imaging (Ximavision); V 7.5.51 with Service Pack (SP); Oncology Information System; Varian Medical Systems | Varian Medical Systems Inc | 2 | Unexpected loss of Treatement Field Add-Ons (Wedges; Blocks; etc) during Ximatron simulation. If an unapproved treatment field is created without an MLC; but with an add-on; the add on will be deleted without warning. If this is not detected and corrected; incorrect dose delivery will occur. | The firm issued an Urgent Medical Device Correction letter to its customers by mail on June 25, 2007. It informed consignees of the anomaly in the software and workarounds for the issue. In addition, the firm also stated that they are currently validating a new software modification to repair the problem.
Z-0549-2007 | Feb-23-2007 | 2007 | Symbia T Series SPECT-CT System with a Pinhole Collimator - Symbia T Series SPECT-CT (Single-Photon Emission Computed Tomography and Computed Tomography) System with a Pinhole Collimator; Siemens Medical Solutions USA; Inc.; Molecular Imaging; 2501 N. Barrington Road; Hoffman Estates; IL 60195-5203 USA; part number 08717733 | Siemens Medical Solutions USA; Inc. | 2 | Symbia systems; running on e.soft version 5.5 or earlier; could allow a fully-extended patient pallet to come into contact with the pinhole collimator; resulting in system damage. | Siemens sent Customer Safety Advisory Notice/Urgent Field Correction Recall letters dated 1/30/07 to all users of Symbia systems with a pinhole collimator and e.soft software versions 5.5 or earlier. The accounts were informed of the potential for a fully-extended patient pallet to come into contact with the pinhole collimator, resulting in system damage. The customers were advised that a Siemens Customer Service Representative would be sent to their site to upgrade their system software to version syngeo MI Applications 2006A within 30 days. The upgrade will be done free of charge.
Z-0686-2007 | Mar-31-2007 | 2007 | CyberMED Version 2.1 Pharmacy Software-Aspyra - Aspyra; Inc. CyberMED Version 2.1 Pharmacy Software | Aspyra; Inc. | 2 | Medications prescribed for one patient were printed on the Medication Administration Record (MAR) of another patient. | Current customers were mailed a Recall Letter January 31,2007 explaining the issue and informing the customers of an upcoming CyberMED Version 2.1 maintenance software release to correct the potential issue. A fax-back Acknowledgement of Receipt form was sent with the Recall Letter.
Z-1243-2007 | Sep-25-2007 | 2007 | Cerner Millennium RadNet Auto Launch Study and Auto Launch Report - Cerner Millennium RadNet Auto Launch Study and Auto Launch Report software functionalities. | Cerner Corp | 2 | Defects in the Auto Launch functionality make it possible for a mismatch of patient data. | The recalling firm initiated recall via telephone on 8/23/07 to their customers who had the 2007 version in their production environment, explained the reason for recall, and informed them they should turn off the affected functionality. The firm posted details of the situation on their website in two different documents (referred to as a Priority Review Flash), one dated 8/29/07 and the other dated 8/31/07. The 8/31/07 document was also mailed to their customers via certified mail on 9/4/07. The corrected software (EP28822 ver.1) was made available to the firm's clients on 9/7/07 via their website. Beginning 9/18/07, workarounds were discussed via telephone with the clients who had downloaded the software correction which contained a new bug that affected the dictation file. A new software correction to fix the dication file problem was posted on the recalling firm's website on 9/28/07 for clients to download and validate. A Priority review Flash #PR07-0176-2 client notification dated 10/4/07 was also posted on their website and mailed via certified mail return receipt.
Z-0294-2008 | Nov-27-2007 | 2007 | Medtronic 0-arm Imaging System - Medtronic 0-arm Imaging System; Product Catalog Number: B1-700-00027; Medtronic Navigation; Inc.; Littleton; MA 01460(software version 3.0) | Medtronic Navigation; Inc | 2 | Reversed Instrument Position : a defect in the software version 3.0 causes an error in the position of surgical instruments and relative motion are reversed 180' on the StealthStation image. (If the StealthStation is not used; or if the patient orientation is prone or supine; the problem does not occur) | Medtronic Navigation notified customers by letter on October 17, 2007. Medtronic Field service will perform upgrade (software version 3.0.1) at user site.
Z-1010-2007 | Jul-03-2007 | 2007 | BrainLAB Target Positioner for Leksell Headring - BrainLAB Target Positioner for Leksell Headring used with the iPlan RT Dose (versions 3.0; 3.0.1 and 3.0.2) and BrainSCAN (version 5.31) Radiotherapy Treatment Planning Software; ; BrainLab AG; Ammerthalstrasse 8; 85551 Heimstetten; Germany; catalog/version number 40700-3A | Brainlab AG | 2 | Target Alignment Error : A mechanical component/software incompatibility caused by a specific combination of the BrainLAB Target Positioner version 40700-3A for Leksell headrings and BrainLAB planning software resulted in a 1.25 mm shift in target area alignment during radiosurgery treatment. | BrainLAB immediately notified all customers with the BrainSCAN or iPlan RT Dose software in combination with the BrainLAB Target Positioner for Leksell Headring, via e-mail on June 5, 2007. The Urgent Medical Device Product Notification letter dated June 4, 2007 included in the e-mail informed the customers that both the BrainSCAN and iPlan RT Dose treatment planning software calculates the isocenter position printed on the Target Positioner Overlays shifted 1.25 mm in cranial direction from the intended position, when the BrainLAB Target Positioner for Leksell Headring Revision A is used. The customers were instructed to cease using the BrainSCAN or iPlan RT Dose software in combination with the BrainLAB Target Positioner for Leksell Headring for treatment planning and patient set-up. An updated target positioner will be provided to the customers to correct the problem, estimated to be available in September 2007. Any questions were directed to their local BrainLAB Customer Support Representative or CAPA Manager Markus Hofman at 49-89-99-15-68-44 or 1-800-597-5911 in the U.S.
Z-0257-2008 | Nov-24-2007 | 2007 | ACCELERATOR APS System - Abbott ACCELERATOR APS System Input/Output Module (IOM); List Number: 7L01-01; Abbott Laboratories; Abbott Park; IL 60064 | Abbott Laboratories; Inc | 2 | Sample problems; Five issues identified related to software utilized by the ACCELERATOR APS System Input/Output Module; to include: Under three specific conditions; the post-aspiration radio-frequency identification (RF ID) read of the tube carrier is not performed and the APS work cell may not appropriately generate Sample Presentation/Sample Queue errors with APS software version 1.1.1 or earlie | Firm notified consignee via letter on 11/10/07. Consignee notified of software issues and informed they would be contacted by an Abbott representative to schedule installation of software upgrade.
Z-0580-2007 | Mar-08-2007 | 2007 | Lifeline AED-Defibtech Sentry - Defibtech Sentry Semi-Automatic External Defibrilator (AED) sold as 'Lifeline AED' brand -Model DDU-100A (version 2.002 and earlier) | Defibtech; LLC | 1 | Device Malfunction-The self-test software may allow a self-test to clear a previously detected low battery condition. If this occurs; the operator may be unaware of the low battery and the device may be "unable" to deliver a defibrillation shock; which could result in failure to resuscitate a patient. | Defibtech LLC initiated notification to customers/distributors on 2/17/07 by email. Hard copy letters followed on 2/22/07 with acknowledgement card. Users provided with instructions for 'Immediate Recommendations' until the unit''s software is upgraded. Distributors, their customers, and direct accounts instructed to locate affected units and upgrade these units in the field. Distributors, their customers, and direct accounts shall return to Defibtech documentation confirming unit software upgrade was completed..Press was issued by the firm on March 6, 2007.
Z-1091-2007 | Jul-18-2007 | 2007 | Colleague 3; 3 CX and 3 CXE Volumetric Infusion Pumps - Baxter Colleague Triple Channel Volumetric Infusion Pumps; Baxter Healthcare Corporation; Medication Delivery Division; Deerfield; IL 60015 U.S.A.; Made in Singapore; product codes 2M8153; 2M8163; 2M9163 | Baxter Healthcare Corp. | 1 | A software anomaly is causing newly upgraded Colleague Triple Channel Infusion Pumps to alarm; display an error code (16:310:867:0002) and stop the infusion. This occurs during user programming with all three channels simultaneously infusing fluids. In reported cases; the pump stopped infusing and both an audible and a visual alarm notification were activated. | Baxter telephoned the direct accounts on 6/20-21/07, informing them that Baxter has received multiple reports of newly upgraded Colleague Triple Channel Infusion Pumps alarming, displaying an error code and stopping infusion. These reported events occurred during user programming with all three channels simultaneously infusing fluids. In all cases, the pump stopped the infusions and both an audible and visual alarm notification were activated. The accounts were requested to immediately remove the affected triple channel pumps, identified by two attached hangtags entitled "Battery" and "Power On", from service. The accounts were informed that their Baxter representative would be in contact with them to coordinate the replacement of the affected pumps with upgraded single channel pumps or non-upgraded triple channel pumps, if available. Any questions were directed to Baxter's Medication Delivery Services at 1-800-843-7867. Baxter sent follow-up Urgent Product Recall letters dated 6/22/07.
Z-1238-2007 | Sep-27-2007 | 2007 | Exacta-Mix 2400 - Exacta-Mix 2400; OS v1.07; Model No. 8300-0073; Pharmacy Compounding System; Baxa Corporation; Englewood; CO 80112. | Baxa Corporation | 1 | A software defect could allow up to 50mL of extra volume being added to a TPN solution. | All customers were notified via telephone on 06/27/2007 with follow-up of a faxed/emailed Urgent Product Recall letter dated 27 June 2007, detailing the software defect. End users were given a work around procedure to use until conversion back to version 1.06 could be scheduled. An additional copy of the recall letter was sent by FedEx.
Z-2388-2008 | Sep-21-2008 | 2008 | Lifepak Defribrillator/Monitor - Physio-Control Lifepak 20 defibrillator/monitor is intended for the termination of certain fatal arrhythmias. | Physio Control; Inc. | 2 | Potential for the coin battery to drain prematurely causing the monitor clock time and date to be incorrect and the service light indicator to illuminate. | On 06/18/08 the firm, via regular postal service , mailed an "IMPORTANT MEDICAL DEVICE NOTIFICATION" letter to all their consignees. The letter instructs consignees that the coin battery may deplete prematurely and to contact Physio-Control if their Lifepack 20 defibrillator/monitor requires coin battery replacement prior to the recommended four year replacement schedule. If you have any questions, contact Technical Support at 1-800-442-1142, option 5 from 6:00 am to 4:00 pm (Pacific), Monday-Friday or visit www.physio-control-notices.com/coin battery.
Z-0162-2009 | Oct-27-2008 | 2008 | VISTA Negative Pressure Wound Therapy Charger - GlobTek; Inc. Power Supply for Medical Use Part No: TR9CE4000LCP-Y-MED MODEL: GTM21097-5012; Made in China.Indicated for patients who would benefit from a suction device particularly as the device may promote wound healing; including patients who would benefit from vacuum assisted drainage and removal of infectious materials or other fluid from wounds under the influence of continuous pressure. | Smith And Nephew; Inc. Wound Management Division | 2 | External battery chargers used with the Version 29 VISTA Negative Pressure Wound Therapy pumps are failing to properly charge the pump's battery. | Each consignee was notified by recall letter (Field Correction Notification) dated 8/7/08 via FEDEX overnight. Within five (5) working days of the date of the customer notification a Smith & Nephew representative will contact the customers to assess their inventory and schedule delivery of replacement battery chargers. Contact Smith & Nephew, Inc. (Customer Care Center) at 1-800-876-1261 for for assistance.
Z-2435-2008 | Sep-20-2008 | 2008 | GE Innova 2000 Fluroscopic x-ray system - Innova 2000; The principle system components include a C-arm; image acquisition; processing and archiving capabilities.Indicated for use in generating fluoroscopic images of human anatomy for vascular angiography; diagnostic and interventional procedures; and optionally; rotational imaging procedures. Also indicated for generating fluoroscopic images of human anatomy for cardiology; diagnostic; and interventional procedures. Intended to replace fluoroscopic images obtained through image intensifier technology. | GE Healthcare | 2 | The charger and / or battery of the PDB may fail earlier than expected with no advance warning and thus cause the system to shut down. | Consignees (Hospital Administrators; Hospital Risk Management Department; Mangers of Radiology/Cardiology; Radiologists/Cardiologists) were sent a "GE Urgent Medical Device Correction" letter dated July 2008. The letter described the affected products; the problem; and the user instructions, which state that the hospital should establish emergency workflow procedures to ensure patient safety should a component stop operating during a patient examination. Facilities should verify if their system has been impacted by checking if the catalog numbers listed are referenced on the PDB panel located in the equipment room. The letter also states the GE Healthcare Resolution Plan, includes correcting the condition, and contact Information.
Z-1749-2008 | Sep-01-2008 | 2008 | UBiT¿-IR300 Infrared Spectrophotometer - UBiT¿-IR300 Infrared Spectrophotometer; Software Versions 7.0; 7.01 and 7.11. | Otsuka America Pharmaceutical; Inc. | 2 | Medical device power supply may overheat and cause smoke to be emitted from the device. | The initial distributor in the US notified consignees 03/14/08 and 03/26/08 by letter. Consignees were provided two different letters depending on the hours of use of the spectrophotometer. The notifications advised of the firm's preventive service and upgrade program for the instrument's power supply, in which consignees would receive a temporary replacement instrument while the service is completed. Consignees with spectrophotometers having <15000 light source hours were notified similarly, and were provided filters and instructions for cleaning and replacing the filters to protect the power supply from unnecessary accumulation of debris. Contact firm at 1-888-MERETEK for assistance.
Z-0119-2009 | Oct-20-2008 | 2008 | Sonesta 6210 Fluoroscopy Procedure Table - Sonesta 6210 Fluoroscopy Procedure Table; a motorized procedure table with adjustments (via a remote control) for height; tilt; back and seat cushion position; Stille Surgical AB; Sundbybergsvagen 1A; SE-17173 Solna; Sweden | Stille AB | 2 | The table contains a battery backup; not mentioned in the User Manual; which automatically will provide power to the electrical actuators/motions in case of a power loss from the main power supply. However; if a short circuit occurs; it can result in an uncontrolled movement of the table which cannot be stopped by disconnecting the device from the main power supply; as the battery automatically w | Stille sent Field Safety Notices dated 8/27/08 to the direct accounts, informing them that the table contains a battery backup, not mentioned in the User Manual, which automatically will provide power to the electrical actuators/motions in case of a power loss from the main power supply. However, if a short circuit occurs, it can result in an uncontrolled movement of the table which cannot be stopped by disconnecting the device from the main power supply, as the battery automatically will supply continuous power to the control box. The accounts were requested to complete and return the enclosed acknowledgement card to Stille by fax or e-mail as soon as possible and no later than 9/26/08. The accounts were advised to not further use the table in conjunction with a C-arm or any other heavy object until a service technician can permanently remove the battery from the control box according to the enclosed Service Manual Amendment No. 1. Stille will make a permanent correction to the table once it has been validated and approved. Contact Stille Surgical Inc. at 1-800-655-1614 for assistance.
Z-0710-2008 | Feb-05-2008 | 2008 | Horizon Cardiology Hemo Monitoring System - Horizon Cardiology Hemo Monitoring System;Horizon Cardiology Hemo External Control Unit;~100 - 230 VAC50-50 Hz; Medcon Telemedicine Technology LTD.;Tel Aviv; 697190; Israel. | McKesson Provider Technologies | 2 | Power supply reliability: McKesson received word from their manufacturer of the medical grade power supply that some power grade supplies may not provide reliable power output. | Recalled by letter sent to the firm's customers beginning on 11-12-2007, advising them that the field correction would be completed by the recalling firm service technicians.
Z-0294-2009 | Nov-25-2008 | 2008 | Animas IR1250 Insulin Pump - Battery Caps used with Animas IR1250 Insulin Pump.The pump is used to help maintain blood glucose targets. | Animas Corporation | 1 | Unintentional rebooting: Pump products exhibit an intermittent loss of power due intermittent loss of contact between battery cap and battery canister resulting in the device resetting. The failure of the battery cap may result in failure of the device to administer insulin therapy which may result in hyperglycemia. | The recalling firm issued a notification letter to the patients with insulin pumps on 8/5/08 to inform them of the problem and that they needed to replace the battery. The recalling firm contacted OneTouch Ping users by telephone on 8/5/08 for this issue and another issue. The recalling firm issued letters to distributors informing them of the problem and notified customers to replace the batteries. In addition, a letter was issued to physicians to alert them of the situation with regard to current patients. Consumers can contact 1-800-423-4087 for additional information or questions.
Z-1510-2008 | Sep-19-2008 | 2008 | Axis Gamma Camera - Axis Gamma Camera; Scintillation (gamma) camera; Philips Medical System; Cleveland; OH 44143 | Philips Medical Systems (Cleveland) Inc | 3 | Resistor Failure: The Rotate Motion Shunt Resistor may overheat and results in the appearance of smoke; a burning smell; and an electronic stop condition; which will disable the gamma camera rotate motion and all other motions. | Consignees were sent an Urgent Field Safety Notice on 1/30/08. The letter instructs users who experience a failure to remove the system from service and contact the recalling firm for immediate correction. The recalling firm will schedule appointments with users to replace the affected part beginning in March 2008. For additional information, contact 1-800-722-9377.
Z-0728-2008 | Feb-16-2008 | 2008 | Bionime Rightest Blood Glucose Monitoring System - Invictus-- Bionime Rightest Blood Glucose Monitoring SystemMODEL: GM300; Invictus Scientific; Inc.; San Diego; CA | Invictus Scientific Inc | 3 | Abnormal Memory Recall: When the meter is fixed in 12 hour mode; the time frame display of all memory data will be shown incorrectly as AM or PM depending on when the user is recalling the memory stored in the meter. This will cause confusion as to when the last test was taken. | An Urgent: Device Recall Notification letter were sent to customers (direct accounts) beginning on November 20, 2007, via US Postal Service, with a return receipt requested. The distributors were instructed to inform consumers/users to fix the time format to the 24 hrs mode and to contact the distributors or the firm for replacement, if they prefer to use the time format in the 12 hrs mode. They were also instructed that if they have further distributed any of the recalled products, to please immediately contact those accounts and advise them of the recall situation, and have them return their outstanding recalled stocks. If customers have any questions, they were instructed to contact at 858-481-8485.
Z-0670-2008 | Jul-03-2008 | 2008 | Parks Medical Ultrasonic Doppler Flow Detector - Parks Medical Ultrasonic Doppler Flow Detector; Model 811-B (finished assembly without accompanying probes); Parks Medical Electronics; Inc.; Aloha; OR 97007 | Parks Medical Electronics Inc | 3 | Premature battery failure: Wrong acid flux used on circuit board which may lead to premature battery failure of ultrasonic non-directional doppler flowmeter. | 811-B: On 9/28/07 Parks Medical Electronics notified their related U.S. distributor, Parks Medical Electronics Sales, Inc., Las Vegas, NV by email instructing them to stop further shipments. On 10/1/07 firm followed up with phone call instructing all units be returned for rework. On 11/8/07 firm sent replacement units accompanied with a Warranty Exchange Letter to their other U.S. customers notifying them of affected units. Letters were customized for each customer with a fill-in-the-blank of the model and serial numbers. Letters and replacement units were sent UPS 2nd Day Air and informed customer about the assembly error and instructions for returning the defective units. On 11/8/07 emails were sent to Brazilian and Australian distributors informing about assembly error and affected units. Rework instructions and a schematic were provided by email for each distributor to perform their own rework. On 11/28/97 Canadian customers were sent Warranty Exchange Letters and replacement units via UPS Ground. Firm is retrieving units from US and Canadian firms and will rework them. Firm is waiting to hear whether Brazilian or Austraian distributors need to sub-recall. Canadian distributors instructed to conduct subrecall if needed. All returned units will be reworked by disassembly of unit, removal of affected printed circuit board, washing board to remove acid flux, and testing.
Z-1473-2008 | Jul-23-2008 | 2008 | Medline EndurO2 Oxygen Concentrator with Oxygen Monitor - Medline EndurO2 Oxygen Concentrator with Oxygen Monitor; Portable oxygen generator; Reorder Number: HCS02; Packaged in 1 unit per case; Product is an oxygen generator designed to provide oxygen to patient who has been prescribed oxygen therapy by a physician. | Medline Industries Inc | 2 | Overheating of Components: There is the potential for an electrical component in the oxygen concentrator to overheat; causing damage to the unit and making it nonfunctional. | Consignees were notified by a letter sent on 3/4/08. The letter informed users that the recalling firm will replace the affected units with new units or issue a credit for the returned recalled units. Distributors were requested to notify their accounts of the recall. The letter requested users to complete and return an enclosed response sheet, indicating the number of units to be replaced or returned for credit. For additional information, contact 866-359-1704.
Z-0414-2009 | Dec-10-2008 | 2008 | Fisher & Paykel Healthcare RT240 Adult Breathing Circuit Kit. - Fisher & Paykel Healthcare RT240 Adult Breathing Circuit Kit.lntended Use: The dual heated adult respiratory ventilator circuits are intended as conduits of breathing gas for ventilation of patients; and to maintain the temperature of the humidified gas; to reduce condensation. | Fisher & Paykel Healthcare Inc | 2 | The recall was initiated because certain lot dates of the RT240 Adult Breathing Circuit Kits manufactured on or before June 27; 2008 include a heated breathing circuit that may be more susceptible to damage when used in excess of the specified seven-day maximum duration of use; which may increase the risk of malfunction or fire. | Fisher & Paykel Healthcare, Inc. (FPH) issued an URGENT-PRODUCT RECALL letter (dated October 8, 2008) including a Customer recall letter for consignees/customers to Cardinal Health. Cardinal Health will, in turn, send the Customer recall letter dated October 10, 2008 to its consignees/customers. Fisher & Paykel Healthcare (FPH) also issued a similar set of recall notification letters to Norco on October 15, 2008. Recipients of the recall letter must perform the following: 1) Check all Cardinal Health stock of the product for the affected lots, 2) Immediately quarantine the affected product, 3) Complete a Device Recall Response Sheet, and 4) Fax or e-mail the completed Device Recall Response Sheet. Fisher & Paykel Healthcare, Inc.'s "Dear Customer Letter" for Cardinal Health and Norco instructs consignees to complete a customer response (i.e. fax enclosed Customer Acknowledgement Form) to FPH. The customer letter also provides instructions for product return. Fisher & Paykel Healthcare, Inc. recall communication instructions to Cardinal Health and Norco was for all affected product to be returned and quarantined by the distributors. The distributors will return any affected stock to Fisher & Paykel Healthcare, Inc. for documented destruction.
Z-1710-2008 | Jun-10-2008 | 2008 | Alaris Medley Pump Module - U19 Socketed Integrated Circuit ("U19 IC") on the logic board of Alaris Pump modules (aka. Medley Pump modules); Model 8100 built on or before September 8; 2005. CardinalHealth; San Diego; CA 92130 | Cardinal Health 303 Inc DBA Alaris Products | 2 | A channel error on an affected Alaris System module will stop an active infusion (or monitoring) with an audible and visual alarm. Channel errors due to failures of the U19 Socketed Integrated Circuits (U19 IC) on the logic board of the Alaris Pump module ("Pump module"). The user will be required to replace the affected module. | An Urgent Medical Device Recall Notification letter was sent to consignees on April 10, 2008. Four customer letters were issued. The U19 IC issue affects only the Pump module and the U9 IC issue affects only the Syringe and PCA modules. Therefore there will be one letter for customers affected by all three issues Ul9 IC, U9 IC, and the IUI, one letter for affected Ul9 IC and U9 IC customers, one letter for U19 IC-only customers and one letter for U9 IC-only customers. All mailings were conducted through consignees (distributors as well as direct customers for the affected devices). Distributors will be sent multiple copies of the Recall Notification and Customer Response Cards. The Recall letters all include Required Actions sections which give their customers specific information about the issue and what the customers are to do and includes a web site link for each specific problem. Each recall letter also has chart of Cardinal Health contacts support.
Z-1712-2008 | Jun-10-2008 | 2008 | (1) Alaris Syringe module Model 8110 & Syringe Module 8120 - U9 Socket Integrated Circuit ("U9 IC") on the display board of the following Alaris System modules built on or before October 4; 2005:(1) Alaris Syringe module (Model 8110) (a.k.a. Medley Syringe module)(2) Alaris PCA module (Model 8120) (a.k.a. Medley PCA module); CardinalHealth; San Diego; CA 92130 | Cardinal Health 303 Inc DBA Alaris Products | 2 | A channel error on an affected Alaris System module will stop an active infusion (or monitoring) with an audible and visual alarm. Channel errors due to failure of the U9 Socketed Integrated Circuits (U9 IC) on the display board of the Alaris Syringe Module ("Syringe module") and Alaris PCA module ("PCA module") for modules manufactured on or before October 4; 2005. The user will be required to re | An Urgent Medical Device Recall Notification letter was sent to consignees on April 10, 2008. Four customer letters were issued. The U19 IC issue affects only the Pump module and the U9 IC issue affects only the Syringe and PCA modules. Therefore there will be one letter for customers affected by all three issues Ul9 IC, U9 IC, and the IUI, one letter for affected Ul9 IC and U9 IC customers, one letter for U19 IC-only customers and one letter for U9 IC-only customers. All mailings were conducted through consignees (distributors as well as direct customers for the affected devices). Distributors will be sent multiple copies of the Recall Notification and Customer Response Cards. The Recall letters all include Required Actions sections which give their customers specific information about the issue and what the customers are to do and includes a web site link for each specific problem. Each recall letter also has chart of Cardinal Health contacts support.
Z-1464-2008 | Aug-13-2008 | 2008 | Pulmonetic Systems LTV 950 Ventilator - Viasys healthcare; Pulmonetic Systems LTV 950 VentilatorThe ventilator is suitable for use in institutional; home or transport settings. Fully self-contained ventilator in a compact and lightweight design. It can operate from an internal battery; external battery or AC power. Applications include long term acute care or short term / recovery room. Pulmonetic Systems; 17400 Medina Rd.; Suite 100; Minneapolis; MN 55447. | Cardinal Health 203; Inc | 2 | The failure of an electronic component could result in failure of the ventilator to breathe for the patient and/or failure of the ventilator to properly alarm to alert the caregiver. | An Urgent Medical Device Recall letter dated February 22, 2008 was sent to ventilator consignees/owners and distributors. The letter describes the issue, identifies the affected devices, states that the devices will be corrected by Cardinal Health and provides direction to verify ventilator operation and audible alarm function daily. Service Centers were provided information to identify affected printed circuit board assembly replacement kits' lot numbers and asked to return them. An "Update" was mailed 03/04/2008 to identify additional ventilator serial numbers that also need correction.
Z-0694-2008 | Feb-05-2008 | 2008 | Soletra Implantable Neurostimulator - Medtronic Soletra 7426; Quadripolar Multi-Programmable Neurostimulator for Deep Brain Stimulation; Rx only; Medtronic; Inc.; Minneapolis; MN 55432-5604 USA (The Medtronic Activa System in an implantable; multiprogrammable quadripolar system that delivers electrical stimulation to selected areas of the brain) | Medtronic Neuromodulation | 2 | Cessation of Therapy: separation of internal connections (lifted wirebonds) between the electronic circuit and battery may lead to sudden cessation of therapy. | A letter "Medtronic Urgent: Device Recall" dated October 2007 was sent to Healthcare Professionals. The letter described the device defect, product involved, root cause, probability of occurrence, recommendations and also provided information for assistance. At the time of the recall all of the product remaining in date for use (not expired) had been implanted, therefore Medtronic provided Important Patient Management Information for the implanting physicians to use in patient care and decisions regarding patient explants.
Z-0020-2009 | Oct-06-2008 | 2008 | GE Datex-Ohmeda Aisys - GE Datex-Ohmeda Aisys; Datex-Ohmeda; Inc.; a General Electric Company going to market as GE Healthcare; GE Healthcare PO Box 7550; Madison; WI. 53707-7550 Datex-Ohmeda AisysIntended to provide general inhalation anesthesia and ventilatory support to a wide range of patients. The device is intended for volume or pressure control ventilation. The Aisys is not suitable for use in a MRI environment. | Datex - Ohmeda; Inc | 2 | The electronic vaporization system in the GE Aisys Anesthesia machine contains the following components: Backpressure Valve; Inflow Check Valve and Cassette Interface Board; that can contribute to independent failures. | GE Healthcare "Urgent Medical Device Correction" letters dated 8/12/08 were sent on 8/20/08 individually to Health Administrator/Risk Manger, Head of Anesthesia and Director of Clinical Engineering. The letters described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information. The Safety Instructions state that the machines can be safely used as long as the following guidelines are followed: 1. Carry out the daily system checkout for any system in use. 2. The Aisys Anesthesia will alert the user both visually and audibly if it detects a fault condition. 3. Have an alternative method of anesthesia available at all times. Once each machine has been corrected an "Update" will be provided by the service technician responsible for the correction. Each customer/machine will receive a "GE Healthcare Urgent Medical Device Correction Update" letter which informs the customer that this correction is a follow-up activity to the letter that was sent earlier.
Z-1619-2008 | Aug-15-2008 | 2008 | BCI¿ 3180 Pulse Oximeter - BCI ¿3180 Oximeter - Patient monitoring medical device; Manufactured by Smiths Medical PM; Inc. For continuous patient monitoring of blood oxygen levels; pulse rate and pulse strength measurements. It is only used in hospital or clinical settings or emergency response environments. | Smiths Medical PM; Inc. | 2 | The Smith Medical PM; Inc. BCI ??3180 Pulse Oximeter tantalum capacitors C38; C72; C74; and C98 were installed backwards during assembly at the board manufacturer. This exposes the components to reverse polarity voltages that may lead to their pre-mature failure. As a result; the identified monitors may exhibit failure modes such as the monitor may shut down and not produce alarms or signals to | Consignees were contacted by an URGENT: Product Safety Information Voluntary Recall Notification letter dated March 11, 2008 with attached Safety Action Bulletin 08-SAB01 dated February 27, 2008 and email between March 7 and March 13, 2008. The consignees will need to return the device to be evaluated in the Smiths Medical PM, Inc.
Z-0324-2008 | Jan-17-2008 | 2008 | Actiwatch - Actiwatch 16 Part Number 198-0300-00; Actiwatch 16 AM 198-0300-AM; Actiwatch 64 Part Number 198-0301-00; Actiwatch 64 AM Part Number 198-0301-AM; Mini Mitter; a Respironics Company; Bend; Oregon | Mini-Mitter Company; Inc. | 2 | Communication Failure: the devices have a memory chip that has a limited number of read/write cycles which can trigger set up information and/or subject information anomalies or intermittent communication annomalies. If left unaddressed; the issues can lead to permanent communication failure. | A Voluntary Medical Device Recall letter dated June 26, 2007. The recall letter will instruct the customers to 1) fill out a business reply form to confirm contact information and devices serial numbers in their possession. Then FAX this info back immediately to Stericycle or Mini-Mitter. 2) Refer to the work scheduled rework date reserved for them and noted in their recall notification letter and return the device for service at that time; and 3) Package the device appropriately, insure it and return it via UPS using the provided UPS shipping label. Product returned to Mini-Mitter will be quarantined, decontaminated, and repaired. Min-Mitter will rework all the devices returned. Each will have a new memory chip installed, be re-calibrated, retested, and have a new battery installed if required. Min-Mitter will provide a new one year warranty.
Z-2337-2008 | Sep-25-2008 | 2008 | Philips HeartStarMRtX - HeartStarMRtX Defibrillator/MonitorModels: M3535A/M3536A HeartStart MRx with a native 256Mbyte internal memorycard with date code "0308. The product is used for the termination of ventricular tachycardia and ventricular fibrillation. | Philips Medical Systems | 2 | Defective internal memory cards may exhibit one or more of the following behaviors:"Device repeatedly restarts approximately every 20 seconds"Slow device start-up of approximately 15 seconds or more. During this time; the MRx displays either a blank screen or the Philips HeartStart MRx Start-up screen prior to displaying the selected mode of operation (i.e. Monitor; AED; Pacing or Manual Defib | Philips issued an Urgent Medical Device Correction notification on July 17, 2008. Customers in the US received their communication package via UPS. Letters to international customers mailed via a tracked method by local Philips representatives. Customers were informed on how to locate the serial number on their devices to determine whether or not they have affected units in their inventory. Customers are asked to follow the "Action To Be Taken By Customer/User" while they await , their correction. If users continue to use prior to correction, a back up defibrillator should be readily available. Devices that exhibit repeated restart or slow start -up should be removed from service. The correction will come in the form of an internal memory card replacement. Philips will contact customers to arrange for the correction of their devices. Contact Philips Medical at 1-978-687-1501 for assistance.
Z-0152-2009 | Oct-21-2008 | 2008 | Baxter Auto Syringe AS50 Infusion Pump - Baxter Auto Syringe AS50 Infusion Pump; product code 1M8550; Baxter Healthcare Corporation; Medication Delivery Division; Deerfield; IL 60015 U.S.A.Indicated for infusion via intravenous; intra-arterial; epidural or subcutaneous routes of administration. | Baxter Healthcare Corp. | 2 | Non-conforming Electrostatic Discharge (ESD) grounding squares may cause the pump to be susceptible to shorting-out of circuitry; resulting in a loss of audio; and/or interruption of therapy. | Urgent Product Recall letters dated 9/5/08 were sent on the same date via first class mail to the direct accounts, to the attention of the Director of Biomedical Engineering and Facility Risk Manager. If the account had device(s) that had been serviced by Baxter with a part affected by this action, a serial number list of the affected device(s) was provided with the notification. Customers were instructed to remove the affected device(s)/spare parts from service and contact Baxter to arrange for their return and repair. Alternatively, if removal of the device is not a viable alternative for the customer and they wish to perform self-service, the customers will also be provided with a Service Bulletin to provide instruction on replacement of the rear case assembly.
Z-0503-2009 | Dec-16-2008 | 2008 | Maxi Move - Maxi Move; Model NDA0200-20 Batteries Pack; BHM Medical Inc.; 2001 Tanguay; Magog; QC; Canada J1X 5Y5The equipment is intended to be used in hospitals; nursing homes; or other health care facilities by trained caregivers for the transfer of patients. | Arjo; Inc. | 2 | The battery pack NDA0200-20; lot 070826D; sold with the Maxi Move patient lift or as replacement parts; has a defective connector that could lead to the inability of the battery pack to recharge and/or short circuit; resulting in smoke emission. | Arjo sent the Urgent Device Recall Customer Notification dated 11/5/08 and Field Safety Notice (FSN) dated 10/31/08 to all affected Maxi Move customers on 11/5/08. The accounts were informed of the problem with the connector of the battery pack NDA0200-20, lot 070826D, sold with the Maxi Move patient lift (model numbers starting with KMC)or as replacement parts. They instructed the accounts to discontinue use of the battery pack if it does not recharge or has an intermittent electrical contact. The accounts were requested to complete and return to Arjo Inc., the enclosed customer response form indicating that they have received the notification package. The accounts will be contacted by an Arjo representative within a week to arrange the replacement of the battery pack and corresponding connectors (male part/plug) on the lift and battery charger.
Z-0445-2009 | Dec-29-2008 | 2008 | Terumo APS 1 - Terumo Advanced Perfusion System 1; 100V -120V; 15 A (circuit breaker); 50/60 Hz (20A power source required) base; Model 801763. | Terumo Cardiovascular Systems Corporation | 2 | The power supply may fail to charge the batteries due to various hardware malfunctions. | Terumo notified U.S. consignees via Urgent Medical Device Correction: Safety Advisory recall letter dated 4/17/08 and international consignees were notified via email on 4/24/08 informing them of the potential problem.
Z-1649-2008 | Aug-25-2008 | 2008 | Smith Medical PM; Inc.; BCI ¿ Mini-Torr Plus ¿NIBP Monitor with SpO2 Built-In Printer - BCI Non Invasive Blood Pressure MonitorProduct/Catalog number 6004; | Smiths Medical PM; Inc. | 2 | Due to a component that may be installed backwards; the device may turn off or may immediately reset. | Consignees were contacted on 4/2/08 with a Safety Action Bulletin, 08-SAB02, and Technical Worksheet 08-TW01 dated April 1, 2008, which included information regarding the potential safety issue and the proper inspection method for the monitor(s). If you have not received the Safety Action Bulletin or the Technical Worksheet contact Smiths Medical PM, Inc. at Telephone: (262) 542-3100 or Toll Free: (USA): (800) 558-2345 or email info.pm@smiths-medical.com.
Z-1693-2008 | Sep-23-2008 | 2008 | iLab Ultrasound Imaging System - iLab Ultrasound Imaging System; Models 120INS and 240INS; manufactured by Boston Scientific Corporation; Fremont; CA 94538 | Boston Scientific Corporation | 2 | Improperly terminated wires on a component of the iLab Acquisition processor power supply could cause the processor to lock-up or stop; and prolong the patients procedure. | On April 4 and 10, 2008, customer notifcations were sent to consignees. Field corrections are estimated to begin on April 25, 2008.
Z-1050-2008 | Mar-29-2008 | 2008 | LIFEPAK 20 - Physio-Control LIFEPAK 20 defibrillator/monitor; AC powered external defibrillator with battery backup. Part Numbers: 3202487-073; 3202487-079; 3202488-036; Physio Control; Inc; Redmond; WA 98052 | Physio Control; Inc. | 2 | Inoperable Device - Potential for delay in therapy or prevention of defibrillation therapy due to corrosion of the Printed Circuit Board Assemblies (PCBA). | The firm contacted Customers by phone and by Urgent-Medical Device Recall letters sent via return receipt requested to their consignees on 1/8/08. The letter advised consignees that the Printed Circuit Board Assemblies (PCBA) might not meet quality requirements. This manufacturing deviation could lead to eventual PCBA corrosion that could render the device inoperative and not ready for use in an emergency under battery power. The firm recommended keeping the unit in service until it could be replaced. Representatives will visit each consignee and replace units.
Z-2389-2008 | Sep-21-2008 | 2008 | Bravo Multi-modality system - NicoletOne photic adapter cable 085-463700 | Cardinal Health NeuroCare Division | 2 | Cardinal Health NeuroCare is implementing a voluntary field correction to replace a quantity of photic adapter cables; part number 085-463700; used to connect the NicLED Photic Strobe option; part number 698-625500; with NicoletOne EEG system. The photic adapter cable has a wiring error which delivers a reduced voltage and results in a decreased intensity of the NicLED photic. This reduced intensi | A "Cardinal Health Urgent: Medical Device Correction" letter was sent to consignees. The letter described the product and the problem. The consignees were asked to confirm completion of the cable exchange by returning an information card. If you have questions, contact Cardinal Health (formerly VIASYS) NeuroCare Technical Support Group at 1-800-356-0007 (US), 608-273-5000 (OUS) or email shane.sawall@cardinalhealth.com.
Z-0567-2008 | Jan-31-2008 | 2008 | CS 100 and CS300 Intra-Aortic Balloon Pump - Datascope CS 100 Intra-Aortic Balloon Pump and DatascopeCS 300 Intra-Aortic Balloon Pump; Datascope Corp. Mahwah; NJ 07430 | Datascope Corp | 2 | Defective Circuit Board -- A defect in the Printed Circuit Board's (PCB) in the CS 100 and CS300 Intra-Aortic Balloon Pump may cause the pump to exhibit intermittent malfunctions; which include failure to start-up or reset of the display screen during therapy. | Datascope sent a Product Field Correction notification letters to Hospital Administrators, via certified mail, on November 19, 2007. The letter informed them of a defect in the circuit board. It described the possible malfunction and provided instructions should malfunctioning occur. Datascope will be replacing the defective circuit boards in all affected IABP units. A Service Representative will be contacting the hospital to arrange for the replacement of the defective board.
Z-2452-2008 | Sep-21-2008 | 2008 | SOMATOM Emotion 6 - SOMATOM Emotion 6. Computed tomography X-ray system | Siemens Medical Solutions USA; Inc | 2 | Potential electrical shock hazard. The LCB power Switch (S1) could break causing the protective plastic cap to come off exposing the energized electrical contacts within the switch. The LCB S1 does have an inline fuse which is designed to open in the event of a short circuit condition between the internal switch contacts. | The firm issued and Safety Update in July 2008 informing users that the firm will have representatives visit each account to replace the cap per Update Instructions CT032/08/S.
Z-0484-2009 | Dec-02-2008 | 2008 | STERRAD NX Sterilizer - STERRAD NX Sterilizer; Product Code 10033Low temperature sterilizers that use hydrogen peroxide gas plasma to inactivate microorganisms on a broad range of medical and surgical instruments. | Advanced Sterilization Products | 2 | ASP has discovered a component defect in some of the UV lamp power supplies used in certain STERRAD NX Sterilizers. This defect can potentially cause the Hydrogen Peroxide Monitor to give inaccurate readings. | The Urgent: Product Correction customer notifications, dated October 6, 2008, were hand delivered by field representatives starting October 7, 2008. The letter advises that Advanced Sterilization Products has discovered a component defect in some of the UV lamp power supplies used in certain STERRAD Sterilizers. This defect can potentially cause the Hydrogen Peroxide Monitor to give inaccurate readings. A list of the affected sterilizer serial numbers is included with the letter. The letter advises that the Field Service Engineer has come to their facility to inspect and, if necessary, replace this power supply. Customers are asked to sign the letter acknowledging receipt of this letter upon completion of the work by the Field Service Engineer.
Z-1188-2008 | Jun-26-2008 | 2008 | Reusable Ohmeda Compatible Finger Clip Sensor - Reusable Ohmeda Compatible Finger Clip Sensor; Model Number: E403-02; Product Code: PS2225. To monitor oxygen saturation level of patients under doctor's care. Kentec Medical Inc.; Irvine; CA 92614 | Kentec Medical Inc | 2 | Incompatible components: The Finger Clip Sensors - with a Serial Number prefix of PA514; PA515; YE508; ZC502 and ZG512 - are missing an additional diode and higher resistor in its existing configuration making them inoperable when used to interface with those Ohmeda Tuffsat monitors previously re-configured by Ohmeda to include a detection circuit for diodes. | Consignees were notified via telephone on 5/18/07 and an "Important Voluntary Market Withdrawal Notice" letter on 5/23/07. The letter instructed consignees to segregate and return all affected devices. The letter included an "Acknowledgement of Market Withdrawal Notification" form for customer to complete and return to the recalling firm. For additional information, contact 1-800-825-5996.
Z-0986-2008 | May-14-2008 | 2008 | AVEA Ventilator - VIASYS AVEA Ventilator; Catalog numbers: 17312-00 (USA Standard with Compressor Version) | Viasys Respiratory Care; Inc.dba Bird Products | 2 | Alarm Failure -- There is a potential of alarms not being activated subsequent to specific electronic faults that may cause INOP condition. | On December 18, 2007, a product notification letter was mailed to all customers in the United States via overnight carrier (registed mail) and a product notification letter was mailed to all international customers via Federal Express. The recall letter informed the customers of the problem and also inform them that a Cardinal Health (Viasys) service representative will be contacting them immediately to arrange & service their ventilator. For futher information please contack Andrea von Muller at 760-778-7323.
Z-2209-2008 | Sep-17-2008 | 2008 | VVsky Navigation Platform - VectorVision (VV) Sky Navigation Platform (19" Computer Rack); the 19" computer rack is a component of the VVsky Vario; BrainSUITE iMRI; BrainSUITE NET and BrainSUITE iCT systems; BrainLab AG; Kapellenstrasse 12; 85622 Feldkirchen; GermanyIntended to be an intraoperative image guided localization system to enable minimally invasive surgery. Indicated for any medical condition in which the use of stereotactic surgery may be appropriate and where a reference to a rigid anatomical structure can be identified to relative to a CT; CTA; X-ray; MR; MRA and ultrasound based model of the anatomy. | Brainlab AG | 2 | Diameter of cables used for installation are to small for the applied current. If an internal short circuit is produced the medical power supply will not shut down automatically and will continue to deliver current; which could result in overheating cables. | BrainLAB sent Product Notification letters dated 11/12/07 to their VectorVision Sky end users, informing them of the possible overload of electrical cables in VVsky products with Data Billboards and/or microscope integrations and/or Device Connections Panels (Wall Panels). The users were requested to perform an inspection of the medical power supplies and cabling on a regular basis. If anything suspicious is noted, they were directed to shut down the entire system and contact BrainLAB immediately. BrainLAB contacted each account shortly after the sending of the letter to arrange for the installation of small fuses/circuit breakers according to the specific load limits per device, which will open the circuit in case of overload on faulty circuits, and to exchange the low voltage cables that directly connect the medical power supply from 1.5 mm2 to 2.5 mm2 thickness to withstand the full 20 amp current in case of a failure. Any questions were directed to BrainLAB's customer hotline at 1-800-597-5911.
Z-1051-2008 | Mar-18-2008 | 2008 | TI500 Isolette Infant Incubator - Draeger TI500. Isolette Infant Incubator. Catalog Number: MU20505. Draeger Medical; Inc.; Telford; PA 18969 | Draeger Medical; Inc. | 2 | Heating failure- The power board which controls the incubator heater may not regulate the temperature properly resulting in a high temperature alarm and a loss of temperature control within the patient compartment. | The firm notified consignees on 2/1/08, by an Important Safety Notice. The letter instructed users to remove any units from use until a Draeger Service Representative can replace the power board. A Service Representative will be in contact to schedule a service visit. For questions concerning this recall please contact 215-660-2418 or Anbu.arasu@draeger.com.
Z-1692-2008 | Sep-16-2008 | 2008 | Puritan-Bennett 800 Series Ventilator Backup Power Source Model 802 - Puritan-Bennett 800 Series Ventilator Backup Power Source (BPS) Model 802; Puritan-Bennett Corporation; A subsidiary of Mallinckrodt Inc.; Carlsbad; CA | Covidien Limited | 2 | Wiring in battery back-up power supply may short and cause thermal damage to ventilator. | All customers were notified by Federal Express letter and were contacted by phone through field service engineers. They were told to examine their inventory for any installed or spare part battery assemblies and to remove and quarantine these parts. Any ventilators with installed affected battery assemblies were to be taken out of service.
Z-1616-2008 | May-21-2008 | 2008 | 9051 Electrode Adapter - 9051 Electrode Adapter; Automated external defibrillator; Used with QUIK-COMBO System; Cardiac Science Corp; Bothell; WA 98021. | Cardiac Science Corporation | 3 | Adaptor Cable Incorrectly Manufactured: The red and white connector ends of the adapter are switched. Therefore; the cable will not connect to the electrodes as indicated in the instructions. | Consignees were notified by an Urgent Medical Device Recall letter sent on 01/23/2008. The letter instructs users to check each cable of the affected lot to ensure that it properly connects to the electrode. Users were also instructed to return any defective product to the recalling firm for a replacement. For additional information, contact 1-888-466-8686 x2041.
Z-1711-2008 | Jun-10-2008 | 2008 | Alaris Medley 8000;8100; 8110;8120;8210 and; 8220 - Grey Inter-Unit Interface Connectors ("IUI") on the following Alaris System modules built on or before March 26; 2004: (1) Alaris PC point-of-care unit ("PC Unit") (Model 8000)* (a.k.a. Medley PC Unit);(2) Alaris Pump Module (Model 8100) (a.k.a. Medley Pump Module);(3) Alaris Syringe Module (Model 8110) (a.k.a. Medley Syringe Module);(4) Alaris PCA Module (Model 8120) (a.k.a. Medley PCA Module); (5) Alaris SpO2 Module (Models 8210) (a.k.a. Medley SpO2 Module) with Nellcor technology; and(6) Alaris SpO2 Module (Models 8220) (a.k.a. Medley SpO2 Module) with Masimo technology; CardinalHealth; San Diego; CA 92130 | Cardinal Health 303 Inc DBA Alaris Products | 2 | A channel error on an affected Alaris System module will stop an active infusion (or monitoring) with an audible and visual alarm. Channel errors with the Alaris System due to the Inter-Unit Interface (IUI) connectors on all Alaris System modules (manufactured on or before March 26; 2004). Grey Inter-Unit Interface Connectors ("IUI") on the following Alaris System modules built on or before March | An Urgent Medical Device Recall Notification letter was sent to consignees on April 10, 2008. Four customer letters were issued. The U19 IC issue affects only the Pump module and the U9 IC issue affects only the Syringe and PCA modules. Therefore there will be one letter for customers affected by all three issues Ul9 IC, U9 IC, and the IUI, one letter for affected Ul9 IC and U9 IC customers, one letter for U19 IC-only customers and one letter for U9 IC-only customers. All mailings were conducted through consignees (distributors as well as direct customers for the affected devices). Distributors will be sent multiple copies of the Recall Notification and Customer Response Cards. The Recall letters all include Required Actions sections which give their customers specific information about the issue and what the customers are to do and includes a web site link for each specific problem. Each recall letter also has chart of Cardinal Health contacts support.
Z-2314-2008 | Sep-20-2008 | 2008 | Medical Integrated Multiparameter: Temperature Probe Adaptor - Spacelabs Medical Temperature Probe Adaptor PN-700-0031-00. Device is packaged in clear plastic bags with an attached insert card Insert card reads "TruLink series Dual Temperature Cable REF 700-0032-00 Spacelabs Medical Issaquah; WA USA" | Spacelabs Healthcare; Incorporated | 2 | Intermittent contact of temperature probe adaptors can cause a loss of temperature monitoring which can cause a delay in care when an alarm is missed. | Starting July 3, 2008 the firm issued a "MEDICAL DEVICE CORRECTION" letter notifying their customers of the potential hazards with the use of probe adapter, PN 700-031-00. All letters have a return receipt to verify each customer received the medical device correction letter. The firm will replace all PN 700-0031-00 adaptors with a replacement adaptor from their initial distributor. The firm initiated sending new adaptors separately, by mail, with a copy of the same Medical Device Correction letter to the firm's customers. The firm has requested their customers to throw away any adapter that is being recalled. Technical Support will be available for customer questions and technical questions and to ensure each unit is upgraded. Contact Spacelabs Healthcare at 425-657-7200 ext 5970 if you have questions.
Z-0496-2009 | Dec-22-2008 | 2008 | THORATEC HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM - Thoratec Heartmate II Left Ventricular Assist System (LVAS). Model/Catalog Number 1355 and Model/Catalog Number 102139.Product is manufactured and distributed by Thoratec Corporation;Pleasant; CA. The HM II LVAS is intended to provide hemodynamic support in patients with end-stage; refractory left ventricular heart failure; either for temporary support; such as a bridge to cardiac transplantation or myocardial recovery; or as permanent destination therapy. The HeartMate II is intended for use inside or outside the hospital. | Thoratec Corp | 2 | The firm has become aware that; over time; wear and fatigue of the percutaneous lead connecting the HeartMate II Left Ventricular Assist System (HM II LVAS) blood pump with external system controller may result in damage that has the potential to interrupt pump function and may require a reoperation to replace the pump. | An initial "Urgent: Medical Device Correction" letter dated October 21, 2008 and a follow-up letter dated October 24, 2008 were sent to affected HeartMate II Left Ventricular Assist System (HM II LVAS) customers. The letters identified the problem, immediate action and preventive action to be taken by customers and recommended that the pump be replaced as soon as possible if damage to the percutaneous lead is confirmed. If customers suspect that a HM II LVAS patient may have a damaged percutaneous lead, please contact Thoratec Technical Services (800-456-1477) for assistance. A return response (Acknowledgement Form) was included with the letter for consignees to complete and fax it to Thoratec Regulatory Affairs at 925-847-8628. Direct questions to the Thoratec Regulatory Affairs Department at 925-730-4135 (Monday-Friday, 8 a.m.-5 p.m. PST). Clinicians and patients with questions may contact the firm at 800-528-2577, or if calling outside the USA, 925-847-8600 (7 days a week, 8 a.m.-5 p.m. PST). Copies of the press release may be found on the firm's website, www.thoratec.com, under Investor Relations/Press Releases.
Z-2227-2008 | Sep-11-2008 | 2008 | Panda¿ iRes and Giraffe¿ Infant Warmers - Panda¿ iRes and Giraffe¿ Infant Warmers; labeled in part*** GE Healthcare 8880 Gorman Road Laurel; MD 20723 U.S.A. www.gehealthcare.com*** Infant radiant warmers provide infrared heat in a controlled manner to neonates who are unable to thermo-regulate based on their own physiology. | Datex Ohmeda; Inc. dba GE Healthcare | 2 | Medical device software may be subject to signal interference of the Hands Free Alarm Silence feature that could impact patient safety. | The recalling firm notified Hospital Administrators, NICU Nursing Manager/Labor & Delivery Nursing Managers and Heads of Biomedical Engineering departments at consignee locations by certified letter on 05/20/08. The notification was flagged as "Urgent Medical Device Correction" and advised consignees of safety issues and recommended actions when using the device. The notification advised that the safest way to mitigate the potential interference issue is to disable the Hands Free Alarm Silence. Users were informed that once a permanent correction action is identified, service representatives will implement the correction. For assistance, contact GE Healthcare at 1-800-345-2700.
Z-0205-2009 | Oct-31-2008 | 2008 | Terumo APS 1 - Terumo Advanced Perfusion System 1 base; 110/120V; Model 801763. For use for up to 6 hours in the extracorporeal circulation of blood for arterial perfusion; regional perfusion and cardiopulmonary bypass procedures. | Terumo Cardiovascular Systems Corporation | 2 | Local pump controls including start/stop button; select button and manual button may not function. | Consignees were notified by recall letter dated 7/18/08 and were provided with revised operating instructions telling them to revert to the central control module should these problems occur.
Z-1660-2008 | Sep-16-2008 | 2008 | LifeSync Disposable LeadWear - LS-202 Large LeadWear Combo Pack - Disposable ECG Monitoring component intended for use as a radiofrequency signal transmitter and receiver of electrocardiograph physiological signals which are displayed on the ECG monitors of various manufacturers' systems that have been validated for compatibility. | Lifesync Corporation | 2 | LifeSync Wireless ECG System - Intermittent failure when using LS-202 and LS-232 LeadWear in conjunction with LS-41245 and LS-41285 Adaptors. The out of box failure manifests as a V-Lead failure; intended for use as a radiofrequency signal transmitter and receiver of electrocardiograph physiological signals which are displayed on the ECG monitors of various manufacturers' systems that have been | Firm issued a Device Recall Letter on 12/14/2007. Users were requested to remove the product from inventory and discontinue use of the Large Sized LeadWear with the 12 Lead Adaptors. The current stock of product in inventory will be held under a stop shipment, pending destruction of non-conforming product.
Z-1526-2008 | Aug-16-2008 | 2008 | GE Datex-Ohmeda Engstrom Carestation - GE Datex-Ohmeda Neonatal Enhancement for the Engstrom Carestation; Model # 1505-3272-00; GE Healthcare P.O. Box 7550; Madison; WI 53707-7550; USA Intended to provide mechanical ventilation for adults and pediatrics weighing 5kg and above having degrees of pulmonary impairment varying from minor to severe. | Datex - Ohmeda; Inc | 2 | Neonatal Flow Sensors supplied for use with the Neonatal Enhancement for the Engstr?m Carestation may result in the sensors providing invalid values and alarm messages when exposed to high flow rates. | Consignees were sent a Medical Device Correction Notification letter on July 2, 2007. The firm's short term recommendation is that Neonatal Flow Sensors may continue to be used with added viligance around potential patient circuit disconnections and the associated alarms. Ther firm recommends that the MV exp high and Circuit leak alarms be set appropriately for each patient in order to detect a patient disconnect. The long term recommendation was to order the replacement part when available and dispose of the old part when the replacement is received.
Z-2230-2008 | Sep-11-2008 | 2008 | Siemens Ultrasound Transducer 17L5HD - Siemens brand Ultrasound Transducer 17L5HD used with ACUSON Sequoia Ultrasound System;Model Number(s): 10035724; used with System Models: 08245875; 08246951; 08255412; 08255413; 08267697; 08269627; 10038241; 10038242; 10040724; 10040725; 10041008;Product is manufactured and distributed by Siemens Medical Solutions USA; Inc. Ultrasound Division; Mountain View; CAThe 17L5 transducer is intended for breast and small body parts imaging. | Siemens Medical Solutions USA; Inc. | 2 | The product assembly can cause a failure of the transducer which results in a double and/or overlapped image which may ultimately result in misleading or false information; inability to accurately diagnose; incorrect positioning/locating/insertion of the biopsy needle and/or an interruption during the biopsy procedure. | On 6/6/08, the firm initiated the recall and its notification was via letters explaining the reason for the recall and requesting the affected product be returned. Affected transducers will be replaced. For assistance, contact Siemens Medical Solutions at 1-650-694-5993.
Z-0531-2008 | Jan-24-2008 | 2008 | Terumo Perfusion Systems Ultrasonic Air Sensor - Terumo ultrasonic air sensor; 1/4" by 1/16" (6.4 mm x 1.6 mm); Black; for use on Terumo perfusion systems 1; 8000 and 9000; Catalog No: 5785; Terumo Cardiovalvular Systems; Ann Arbor; MI 48103 | Terumo Cardiovascular Systems Corp | 2 | False Alarm: The air sensor may malfunction and trigger false alarms; and may continue to alarm thus preventing the device from being reset. | Consignees were notified of the defect/issue and what to do if it were to occur (short term work around) via an Urgent Medical Device Removal letter dated 9/14/07. As the permanent fix users were informed the sensor will be replaced when replacements become available.
Z-2154-2008 | Aug-26-2008 | 2008 | Corometrics¿ MODEL 250 Series Maternal/ Fetal Monitor; - Corometrics¿ MODEL 250 Series Maternal/ Fetal Monitor; Models 256 and 259; part number 2024489-002.Intended for monitoring fetal and maternal vital signs: fetal heart rate; optional fetal movement detection; FHR and UA alarms; maternal uterine activity; heart/pulse rate; blood pressure and %SP02. | Datex Ohmeda; Inc. dba GE Healthcare | 2 | Communication connector within monitor is defective and may cause potential loss of telemetry. | The recalling firm notified healthcare professionals by an Urgent Medical Device Correction letter dated 07/01/2008 to Healthcare Administrators, Risk Managers, Nurse Managers and Directors of Clinical Engineering. The letter advised of a defect in a communication board that may cause a potential loss of telemetry monitoring. The firm recommended that the device should not be used with the firm's 340 Series telemetry system pending replacement of the defective board by service engineers. The Healthcare Professionals are asked to contact the GE Healthcare Technical Support center at 1-800-558-7044.
Z-0126-2008 | Jan-03-2008 | 2008 | NicoletOne - Viasys Healthcare NicoletOne ICU Monitor; Modular Neurodiagnostic System (containing Axiomtek Panel PC computer system); released and installed between December 2004 and March 2007 (Axiomtek Panel PC Model Number: 2147-671) VIASYS NeuroCare Part Numbers: 842-673800; 842-673900; 982A0421 and 982A0422. Manufactured; designed; developed and marketed by VIASYS NeuroCare Inc.; Madison; WI 53711 | Nicolet Biomedical Div of Viasys Healthcare | 3 | Display lock-up: VIASYS Healthcare NicoletOne ICU Monitor freezes up during operation when using the digital video option. | Certified direct mail delivery of URGENT MEDICAL DEVICE FIELD CORRECTION (RECALL) were made to consignees on 9/11/07 by Nicolet Biomedical Div of Viasys Healthcare. The letter describes the product and problem. VIASYS requests that customers cease use of the video option of the device until the monitor can be replaced by VIASYS authorized personnel. Letter also requests consignees to fax or mail back the response form.
Z-2322-2008 | Sep-20-2008 | 2008 | iCon Patient Programmer - Medtronic iCon Patient Programmer; model 3037. The Patient programmer is a hand held; battery operated; microcontroller based device for use by patients to control and monitor external and implantable devices. The minimum amount of control involves the ability to turn therapy ON or OFF; maximal control allows the patient to select from different therapies and adjust the parameters with in the set therapy limits. | Medtronic Neuromodulation | 3 | Medtronic is retrieving four (4) patient programmers that were not properly loaded with application software. The application software is needed for a patient programmer to synchronize and bond with a neurostimulation device. Without this functionality a patient programmer is not useable and can not communicate with a neurostimulation device. | A letter was sent to each account telling them of the issue and to let them know that Medtronic is retrieving the affected devices. Medtronic Reps were to visit affected accounts to notify them of the recall, locate the affected device and return it. Rep is to complete the reply card and fax to Medtronic Neuromodulation. If you have questions, contact your Medtronic representative or Medtronic Neuromodulation Technical Services at 1-800-707-0933.
Z-1320-2008 | Aug-19-2008 | 2008 | Epic VR Model V-197 - Implantable Cardioverter DefibrillatorEPIC VR Tiered-therapy Cardioverter/Defibrillator; Model V-197 | St Jude Medical CRMD | 2 | A condition that could lead to a ventricular sensing anomaly in Epic and Atlas family of lmplantable Cardioverter Defibrillators(ICDs) has been identified. A loss of ventricular sensing would prevent the ICD from being able to detect an arrhythmia. If an arrhythmia occurred and the device was unable to detect it; then a patient would not receive potentially life saving therapy. | A January 16, 2008, letter entitled, "Important Device Upgrade Information" was sent to physicians by FedEx. In the US, sales representatives will follow-up with the physician and will upgrade all programmers. For the rest of the world, it will be monitored by various St. Jude offices and local distributors with this information being fed back to St. Jude Medical CRMD in Sylmar. In addition, St. Jude Medical issued a press release on January 18, 2008 and information concerning this advisory can be found on St. Jude's website, www.sirn.com.
Z-1604-2008 | Sep-23-2008 | 2008 | Deltec¿ 3000 Large Volume Infusion Pump; - Deltec¿ 3000 Large Volume Infusion Pump; CE 0473; RX Only; Smiths Medical MD; Inc. 1265 Grey Fox Road; St. Paul; MN 55112; USA; Made in UK. Reorder numbers 21-5301-01 & 21-5306-01 The product is used for general drug delivery infusion therapies. The product is only intended to be used by trained clinicians. | Smiths Medical MD; Inc. | 2 | The Deltec; Graseby; and 3M; Large Volume Infusion Pump - Models 3000 and 3100; may deliver an unintended bolus if the pump door is opened and then immediately closed. This problem is due to a timing delay when the pumping mechanism resets itself. | A Smiths Medical "Urgent Medical Device Correction" letter dated 2/5/08 was sent to consignees on 2/27/08. The letter was addressed to Risk/Safety Managers, Clinicians, Nursing Staff and other users of the above products. The letter also included Details on affected devices, Description of the problem , Advice on action to be taken by the user, Warning and requested the return the confirmation form. Contact Smiths Medical at 1-800-426-2448 for assistance. Please note: The 3M was the original brand for the large volume infusion pump models 3000 and 3100 marketed in the USA from 1994 through 1997. The Graseby large volume infusion pump models 3000 and 3100 were marketed from 1997 through 1999. Since 1999, the current marketed brand is the Deltec models 3000 and 3100. Smith Medical MD, Inc., included the 3M and Graseby brand large volume infusion pumps in the product field safety notice in the event these brands of infusion pumps are still in use.
Z-1656-2008 | Sep-16-2008 | 2008 | LIFEPAK 1000 - LIFEPAK 1000 defibrillator - portable battery powered semi-automated model that can be operated in either of three modes: Automated external defibrillator (AED); manual defibrillator and electrocardiographic (ECG) mode. The defibrillator is intended for indoor or outdoor use in a wide variety of hospital and pre-hospital settings including ground and air ambulances; physician's offices; aircraft; stadiums and specialty medical clinics. ECG monitoring is intended for use on conscious and unconscious patients of all ages for the purpose of ECG rhythm recognition and heart rate monitoring.Affected part numbers: 320371500021 through 320371500038; 320371500043; 320371500044; 320371500046; 320371500048; 320371500050; 320371500056; 320371500058; 320371500063; and 320371500067. | Physio Control; Inc. | 2 | Potential for the display screen to dim and eventually go blank. The risk to patient is a delay in defibrillation therapy if the defibrillator was needed in a cardiac arrest situation. Death or permanent impairment can occur if the user fails to deliver a defibrillation shock or defibrillation is delayed. | On February 27, 2008 the firm sent a URGENT MEDICAL DEVICE NOTIFICATION letter to consignees advising of the issue. The letter states the firm is performing accelerated function testing and that a follow up notification will be sent if the firm's final test results indicate further action is necessary. The letter recommends consignees "Continue to use your LIFEPAK 1000 defibrillator/monitor Set your device to Auto-Analyze Contact Physio-Control immediately if your display begins to dim or goes blank. What to do if the display goes blank while in use: 1. Continue to use the device to treat the patient. (voice prompts and tones function normally). 2. If the device is not set to Auto-Analyze use the lower right soft key to select Analyze from the display screen... 3. All other buttons are labeled outside the display screen." If you have any question contact technical support at 1-800-442-1142 - option 5 or visit the web site at www.physio-control-notices.com/LP1000.
Z-1707-2008 | Sep-29-2008 | 2008 | Aquilion - Toshiba America Medical Systems (TAMS) Aquilion 32 CT system; computed tomography x-ray system. | Toshiba American Med Sys Inc | 2 | Excessive radiation: Intermittent shifting of CT numbers causes the actual dose to the patient to be greater than expected; when Sure Exposure option is used. | Toshiba sent a letter to customers in July 2007, and sent field service representatives to update the Data Aquisition System in September/October 2007.
Z-0203-2009 | Nov-06-2008 | 2008 | Terumo APS 1 - Terumo Advanced Perfusion System 1 Base; 100/120V; Model 801763. For use for up to 6 hours in the extracorporeal circulation of blood for arterial perfusion; regional perfusion and cardiopulmonary bypass procedures. | Terumo Cardiovascular Systems Corporation | 2 | The pumps may exhibit blank or distorted displays; loss of local control and/or may have local control knobs that are difficult to use. | Consignees were notified via letter dated 7/18/08 and instructed to use the central control module should the problem occur. Contact Terumo Cardiovascular Systems Corp at 1-800-521-2818 for assistance.
Z-0338-2009 | Dec-16-2008 | 2008 | iLab Ultrasound Imaging System; models 120INS and 240INS - Boston Scientific iLab Ultrasound Imaging System; models 120INS and 240INS. Its intended use is for ultrasound examination of intravascular pathology. Manufactured by Boston Scientific Corporation; Fremont; CA. | Boston Scientific Corporation | 2 | Computer malware: Products may be infected with a worm (computer malware) which could infect a computer network to which it may be connected. | Customer letters sent on Friday, Sept 5 via FedEx, 2008 to all customers. This letter will be followed up by a customer service visit. Package delivery will be tracked, as well as Reply Verification tracking forms.For further information, please contact firm at 508-652-5594.
Z-0204-2009 | Nov-06-2008 | 2008 | Terumo APS 1 - Terumo Advanced Perfusion System 1 base; Model 801764. NOTE: This product is not distributed in the United States. For use for up to 6 hours in the extracorporeal circulation of blood for arterial perfusion; regional perfusion and cardiopulmonary bypass procedures. | Terumo Cardiovascular Systems Corporation | 2 | The units may exhibit blank or distorted displays; loss of local control and/or may have local control knobs that are difficult to use. | Consignees were notified via letter dated 7/18/08 and instructed to use the central control module should the problem occur. Contact Terumo Cardiovascular Systems Corp at 1-800-521-2818 for assistance.
Z-0283-2009 | Nov-07-2008 | 2008 | One Touch Ping Glucose Monitoring System - One Touch Ping Glucose Monitoring SystemCombines the functionality of OneTouch(R) Ping (TM) Insulin Pump and a OneTouch(R) Ping (TM) Meter Remote through radio frequency (RF) communication. Together; they provide an option to help make insulin delivery more discreet and flexible. Also; allows the most recent blood glucose results from the OneTouch(R) Ping (TM) Meter Remote to be automatically entered into bolus insulin calculations to cover carbohydrates in food or to correct for a high blood glucose level. | Animas Corporation | 2 | display old; inaccurate values for bolus amount delivered and amount of planned bolus totals | On 8/6/08, the recalling firm telephoned all users to inform them of problem and the need to return the system.
Z-0087-2009 | Oct-08-2008 | 2008 | Apollo Anesthesia Machine - Draeger Medical Apollo Anesthesia Machine; catalog number 8606500 | Draeger Medical; Inc. | 2 | Sporadic errors in various device functions; including low readings of the single gas flows for oxygen; nitrous oxide; or air delivery; or mechanical ventilation failure. Alarms and error messages function as designed. | Draeger Medical sent a Medical Device Recall letter to their customers via certified mail on 8/22/08 informing them of the problem. Until the exchange of affected printed circuit boards has been implemented, the firm is recommending that users closely monitor all device displays and error messages and if applicable act in accordance with the advisory statements and remedial information provided in the Instructions for Use as well as making provisions for alternative patient care measures.
Z-1768-2008 | Sep-04-2008 | 2008 | Coherence; Primeview; AG; Impression and Syngo - Siemens Linear Medical Accelerators equipped with COHERENCE Therapist R2.1 and Primeview 3 i R 2.1. Part numbers 5863506; 7339125; 7341410; 7341428; 7345411; 7345429; 7345437; 8139839; 8139847; 8147667; and 8147675 | Siemens Medical Solutions USA; Inc | 2 | Images may potentially shift; causing an incorrect alignment of the patient; which may result in a dose to the wrong location. It was discovered through the complaint handling process that a potential image shift issue can occur after using the interactive shift tool and filtering tool due to incorrect consideration of the flat panel alignment. | A Customer Safety Advisory Notice was issued on September 19, 2007--All affected consignees were notified certified mail (includes update instruction TH029/07/S). Customers were request to maintain a copy of the safety advisory notice with the COHERENCE Therapist R2.1/Primeview 3i manuals. A copy of the firm's corrections & removal report will be forwarded to Quality Management of Siemens Regional units responsible for reporting to local country authorities and a copy of the report is also forwarded to the firm's distributor.
Z-0214-2009 | Nov-12-2008 | 2008 | Sunquest Laboratory System. - Sunquest Laboratory System. Software Version: Sunquest Laboratory v5.4.2; v6.1; v6.2 and v6.3.The system is marketed as an automated solution for managing laboratory processes including: 1) Patient registration; 2) Order entry and order modification; 3) Specimen collection; verification; suitability; and distribution; 4) Quality assurance checking; 5) Workload recording; 6) Billing charge capture; 7) Standards and controls recording; 8) Test order result inquiry and reporting; 9) Organism susceptibility and epidemiology records; 10) Microbiology culture direct examination and observation recording; 11) Outreach clients; specimens; and results; 12) Client-specific reporting capabilities; 13) Inbound client service request tracking; and 14) Historical data. | Sunquest Information Systems; Inc. | 2 | While performing in-house testing for the Sunquest Laboratory System and under certain conditions; there have been instances in which an error that occurred at the database level or originated in the operating system (Cache)was not communicated or displayed in the Graphic UserInterface (GUI) application so the activity continued. While this error was registered in the error log in the SystemU | A notice (PSN-08-07) was issued to all customers on 08/01/2008 via e-mail. The notice advises customers of the problem and actions Sunquest Information Systems has taken to eliminate the Problem. Sunquest Client Support will contact their facility to schedule correcting the error trap setting. The notice also provides temporary procedural instructions to include monitoring error log using System Utilities (Function UTL) and suboption 2, System Error/Message Inquiry. Contact Client Support at 520-570-2252 for questions or additional information.
Z-1588-2008 | Sep-11-2008 | 2008 | VARIANT ¿-thalassemia Short Program - VARIANT ¿-thalassemia Short Program Hemoglobin Testing System; Model Number 270-0019 (250 tests); with ¿-thalassemia ROM card lot PA 62600; PIN 2700104 | Bio-Rad Laboratories Inc | 2 | Some product ROM cards have an incorrect program loaded and are unable to be updated. | The firm initiated its recall on March 18, 2008 by issuing a recall notice to customers and subsidiaries via e-mail. Not all VARIANT Â®-Thalassemia Short ROM Cards are affected. Inventory should be examined and quarantine the product subject to this Recall. Contact your regional Bio-Rad office for ROM Card replacement or if you have questions.
Z-0152-2008 | Feb-22-2008 | 2008 | da Vinci S Surgical System - Intuitive brand da Vinci S Surgical System; Endoscopic Instrument Control System; Model IS2000 with Isolation Transformer (P/N: 952012) and Vision Cart (Model: VS2000); Distributed and Manufactured by: Intuitive Surgical; Inc. 950 Kifer Road Sunnyvale; CA | Intuitive Surgical; Inc. | 2 | Under-Rated Fuses -- The Firm may have installed incorrect fuses. The daVinci Surgical System Vision Cart's isolation transformer requires a 10A fuse for 115V operation or a 5A fuse for 230V operation. Operation of this product with the use of an under-rated fuse will result in failure of the fuse and a loss of power to the vision cart; including the loss of power to any ancillary equipment that | A product recall was initiated on October 4, 2007. Direct and Distributor consignees were sent an Important Product Notice, via electronic mail on October 5, 2007. The firm plans on sending a Field Service Engineer to inspect the system and make any necessary corrections.
Z-1306-2008 | Jul-01-2008 | 2008 | IMPAX¿ Network Gateway Server; License Key - IMPAX¿ 6.2.1 Network Gateway Server; License Key. Catalog # ABC code: EXEGE000. The product is used to transfer medical imaging data. AGFA Corp.; Greenville; South Carolina 2960. | AGFA Corp. | 2 | Failure of Data Transmission: No imaging data could be transmitted by the device; because a key value needed for data transmission was incorrectly entered in the affected units by AGFA Healthcare. | On 12/11/07, AGFA Healthcare sent out an e-mail to all affected consignees informing them of the situation. The firm also arranged to install the corrected License Key on all affected units of the product.
Z-1242-2008 | Jun-27-2008 | 2008 | Agfa Heartlab Cardiovascular Results Management Product - Agfa Heartlab Cardiovascular Results Management Product: Versions 2.03 and 2.04. Heartlab; Inc.; Westerly; Rhode Island 02891 | AGFA Corp. | 2 | Incorrect Results: Erroneous echocardiographic measurement values due to mathematical formula being mis-configured. | On 7/27/2007 to 8/16/2007, AGFA Healthcare notified, via telephone, consignees of the situation. On 7/27/2007, a formal customer notification letter was sent to affected consignees, providing them with a respective listing of patients that were potentially affected. AGFA informed the consignees that the firm would correct any incorrect patient ECG data. AGFA also informed their consignees that they would provide a configuration update to correct the problem. A notification confirmation form to be return to the firm was also attached to the notification letter, which also provided the firm's customer service number: (877) 777-2432.
Z-0568-2008 | Jan-29-2008 | 2008 | Cyberonics VNS Therapy System - Cyberonics; VNS Therapy System Model 250 Handheld programming system using Dell x5 handheld and preprogrammed with software v7.1 and higher; manufactured by Cyberonics; Inc.; Houston; TX. | Cyberonics; Inc | 3 | Screen Freezes-- The Dell X5 Handheld PC screen will freeze caused due to incompatibility between the Microsoft 2002 OS and the model Dell X5 handheld computer. Once frozen; the handheld device becomes non-responsive to user input. | Cyberonics sent an initial notification (Important Safety Alert) to physicians via Certified Mail on 10/01/07. Letter notified physicians that the Dell X5 handheld computer screen will freeze following an interrogation. Once frozen the handheld device becomes non-responsive to user inputs. The firm informed the user that in the event that the computer becomes non-responsive, the user has two options, 1) Reinitialized the flash memory and, 2) Perform a software reset of the handheld computer. A second notification addressing the approval and addition of labeling was sent to physicians on 11/15/07. Physicians were asked to respond to both letters via an enclosed response card, acknowledging receipt of the notification. Additional information will be posted at www.VNStherapy.com.
Z-0958-2008 | Apr-30-2008 | 2008 | Cytomics FC 500 MPL Flow Cytometry System with MXP Software; Software Version 2.1 - Cytomics FC 500 MPL Flow Cytometry System with MXP Software; Software Version 2.1; Automated differential cell counter; Part Number: 626554; Beckman Coulter; Fullerton; CA 92834 | Beckman Coulter Inc | 3 | Data generation error: If the cytosettings is not refreshed or restarted during data importation; the output will include old and new data. | Consignees were notified by an Urgent Product Corrective Action letter sent on 6/29/07. The letter provided details on the data error and instruction on how to avoid the error. Users were asked to retain notification as part of the Quality Systems documentation and return a response form acknowledging receipt. For additional information, contact 1-800-526-7694.
Z-1766-2008 | Sep-02-2008 | 2008 | ACUSON Antares 5.0 ultrasound system - ACUSON Antares 5.0 ultrasound system at SW revision 200.0.054 or 200.0.059A; with one of the following options installed: 3-Scape Imaging; Factory; PN 05937011 or 3-Scape Real Time Imaging; UPG ACU; PN 10033688. Material number 10032746; 10037592; 10032746; 10037591 or 10038202. | Siemens Medical Solutions USA; Inc. | 3 | Software issue may result in distorted images and inaccurate measurements; which could lead to a misdiagnosis. | The firm initiated a field correction on 03/21/2008. Site visits to correct the problem and install the revised software version 200.0.063A are planned for all consignees. No other notification was made. The firm can be contacted at 650-694-5993.
Z-2180-2008 | Sep-04-2008 | 2008 | IMPAX¿ Clients 6.x - Impax¿ 6.x Clients including: 6.3 all versions below 6.3.1 SU5; 6.2 all versions below 6.2.1 SU14 and 6.0 all versions. The Impax Systems are Picture Archiving and Communications system (PACs). PAC systems are used in the acceptance; transfer; display; storage and digital processing of medical images. | AGFA Corp. | 3 | The IMPAX client failed to "refresh" the image area upon receiving a study retrieved from archive. The system also erroneously displays this study retrieved from archive as the "active" study. | An "Urgent Safety Notice" letter was sent to all the sites using potentially affected units on/about June 4, 2008. Dealers were notified via letter on June 9, 2008. The letter requests that all users be reminded that Agfa recommends the system is configured so that demographic information is not removed from the image viewing area. They will be provided with the appropriate Service Update, based on the current version of IMPAX Client 6.x in use at their site. They were advised to distribute the information within their facility to all those who need to be aware of it. A FAX Back form was requested with each letter to ensure someone at the account acknowledges the information was received and understood. If a FAX Back form is not received, the account will be contacted by telephone. Contact AGFA at 1-877-777-2432, prompts 1, 2, 1.
Z-1317-2008 | Apr-08-2008 | 2008 | OneTouch Data Management Software v1.0 - LifeScan OneTouch Data Management Software Version 1.0; Part Number: 021-722-01(for OneTouch Select and UltraMini (Download) Blood Glucose Meters); LifeScan Inc; Milpitas; CA 95037. | Lifescan Inc | 3 | Meter Temporarily Freezes-- A software compatibility issue may cause the blood glucose meter to cease operations and freeze temporarily. | The firm sent an Urgent Recall letter to Health Care Professionals on 12/18/2007. The letter advised user to uninstall software drivers and provided options on how to uninstall. For further assistance, please contact LifeScan Customer Service at 1-866-679-9250.
Z-1644-2008 | May-28-2008 | 2008 | EVMS (Enterprise Visual Medical System Software) - EVMS (Enterprise Visual Medical System); Picture archiving and communications system; Model Numbers: 5.30.4; 5.30.5; 5.30.6; 5.30.7; 5.40.0 and 7.0.0.1. Emageon; Inc.; Hartland; WI 53029. | Emageon | 3 | Miscalculation: The calculation of the standard uptake value does not include the required calibration factor/decay correction factor. The resulting value is incorrectly calculated; which may lead to incorrect interpretation of tissue density. The calculation error exists in both the software and user guide. | Consignees were notified by Urgent Device Correction letters that were sent on 1/9/08 and 1/18/08. The letters instructed users not to use the standard uptake value in EVMS for diagnostic interpretation when viewing PET studies unless they received service bulletin 445, 446 or 7.0.1. Users were also instructed to replace pages in the user's guide with the pages provided and to affix the enclosed warning sticker on all workstations. Users will be contacted by the recalling firm to schedule a software patch.
Z-1546-2008 | May-28-2008 | 2008 | The Selenia Full Field Digital Mammography System - Selenia Full Field Digital Mammography System with Software v. 3.3.1.1; Full field digital; system; x-ray; mammographic; Model Number: SEL-00002; Hologic; Danbury CT 06810 | Lorad; A Hologic; Inc. | 3 | Marker Misalignment: The new software version contained a magnification factor that shows CAD markers misaligned with the identified indications. The marked area is misaligned up to 2cm from the indication. | Consignees were notified by an Urgent Recall Notice sent on 3/21/08. The letter informed users that representatives were dispatched to all locations and they have corrected the magnification factor problem. The letter instructed users to re-evaluate any cases read using CAD from the period of installation until the date of system correction. The letter also instructed users that only images read with CAD markers are affected. The letter requests users to return a confirmation of notification. For additional information, contact 1-877-371-4372.
Z-1373-2008 | Jul-29-2008 | 2008 | Lumax Cardiac Resyncronization Therapy Defibrillator (CRT-D) - Lumax Cardiac Resyncronization Therapy Defibrillator (CRT-D); Implantable; dual-chamber defibrillator; Model Number: HF-T; Catalog Number 355263; Software Version 603.U; Catalog Model Number 359251; ICS 3000 units are composed of three Catalog/Model Numbers: 349527; 349530; and 355547. The model 349530 is the component which is loaded/updated with software. | Biotronik Inc | 3 | Timing Feature Error: An implementation error allows the usage of the V-V timing feature; which is not approved for use in the U.S. | Consignees were notified by a letter sent 1/10/07. The letter advised users to stop usage of the affected software. Users were instructed not to implant additional devices until the software program has been corrected. The recalling firm requested that all affected software be scrapped or returned to them. Users were requested update the software with a corrected version as soon as it becomes available.
Z-2369-2008 | Sep-18-2008 | 2008 | Philips - Philips Telemonitoring System Software (M3811B) used with Philips Telemonitoring System (M3810B)Revision: B.02.07Prescriptive medical device used to automatically collect and transmit medical information (weight; blood pressure; non-diagnostic ECG) over phone lines between provider and patient. | Philips Medical Systems | 3 | Software: If Weight Limits edited; values will return to default values that were deleted | Affected customers will receive the Urgent Medical Device Correction notification on 7/22/08 via certified mail. Delivery receipt will be tracked. The Urgent Medical Device Correction notification instructs customers to verify that daily and weekly weight limits are set or not set as intended while they await the correction, which will come in the form of a new version of software and will simultaneously update all systems. Philips has corrected the software defect and is in the process of validating the correction and releasing the corrected software as part of a scheduled release. If further information or support is needed, customers should contact the Telehealth Customer Service Line at 1-866-246-7316.
Z-1309-2008 | Aug-28-2008 | 2008 | HOUVA III Phototherapy System with DermaSense Technology - HOUVA III Phototherapy System with DermaSense Technology comes on one skid within 10 cardboard boxes.A therapeutic product designed for individuals who require specific Ultraviolet radiation therapy for diagnosed skin disorders. | National Biological Corp | 3 | Software allows operator to override "low line voltage" error warning and store light intensity value. | National Biological sent an "Important safety reminders for your Houva DermaSense phototherapy device" letters, dated 2/11/08 to users. It states that the user should collect a "Profile" of the light energy and record it on a daily/weekly flow chart. The chart should be in plain view so each operator can clearly determine the units output. If an error message appears in the display panel, treatment should not be continued unless the phototherapy booth operator was able to identify and correct the cause of the error. Make sure both breakers are in the "ON" position before starting up the unit each morning and before giving treatment.
Z-1571-2008 | Jul-24-2008 | 2008 | Diagnostic Ultrasound Unit - EUB-5500 VISION 5500 Diagnostic Ultrasound Scanner w/Step 5 (V05-7) Software. Imaging device intended to provide the physician with physiological and clinical information. | Hitachi Medical Systems America Inc | 3 | Miscalculation reading. A software error in the firm's EUB-5500/HI VISION diagnostic scanning system causes a miscalculation of the left ICA/CCA ratio when using the Carotid Calculation package for patient scans. | A Device Correction Letter was issue by the firm on 02/25/2008. The letter asks the affected customers not to use the value of ICA/CCA on the Measurement Report for diagnosis, but instead to follow the Immediate countermeasure detailed in the 'Note in Use' document which accompanies the Device Correction letter. The customers are asked to follow the Immediate Countermeasure until they receive a Temporary Countermeasure to correct the ICA/CCA ratio calculation. The Temporary Countermeasure will include: 1) a Disk to import a new carotid measure package; 2) an Installation booklet detailing how to import the new carotid measurement package and optional customization tools; and 3) an Operator manual addendum Upon release of the countermeasure, the firm's Applications or Servcie personnel shall perform the corrective action as identified in the Service Pack (A) Installation Instructions. Finally, upon release of the next software revision, service personnel will replace the software at each customer location.
Z-1895-2008 | Sep-16-2008 | 2008 | Essenta DR - Essenta DR - digital multifunctional x-ray system (application SW release 1.0.2 and lower). The device is intended for routine exams; including specialist areas like trauma or pediatric work; excluding mammography. | Philips Medical Systems North America Co. Phillips | 3 | Bright artifacts may appear on a patient image from a previous exposure; which might lead to a misdiagnosis. | On 4/08/08, the firm notified customers by letter, sent by certified mail. The letter explained the recall and advised customers of the following: "If direct radiation reaches the detector, the collimation should be not too tight, such that the area of direct radiation is larger than 30 mm x 30 mm. (1.2 inch by 1.2 inch). The reason is that a minimum area is needed to detect and prevent memory artifact... All possibly affected products in the field will be upgraded with FCO 71200026 to resolve the issue." Contact Sarah Baxter at 425-487-7665 or Philips Call Center at 800-722-9377, #5, #2 if you have questions.
Z-0720-2008 | Feb-15-2008 | 2008 | VARiS PRT Exchange - Varian brand VARiS 1.4g (Medical Charged Particle Radiation Therapy System; Record and Verify System); Linear Accelerator with RTP Exchange v6.2; v6.6; & v8.0; All Models; Varian Medical Systems; Palo Alto; CA | Varian Medical Systems Inc | 2 | Incorrect Software Validation- The ARC treatment plans containing a Dose Dynamic MLC for IMRT delivery that will not be recognized properly by the system. Under these circumstances; an arc plan can be administered to a patient whereby the gantry will not rotate; resulting in delivery to a single location rather than through the gantry arc. | An Urgent Medical Device Correction letter was distributed October 2, 2007, to all direct consignees, informing them that the ARC treatment plans containing a Dose Dynamic MLC for IMRT delivery will not be recognized properly by the System. The firm provided corrective instructions and informed consignees that a product modification will be developed to resolve the issue. A Varian Service Representative will be scheduling a visit to upgrade the application.
Z-0976-2008 | Apr-30-2008 | 2008 | Impression Therapist - Siemens Impression Therapist 3rd Party V&R medical charged-particle radiation therapy system; Part No: 7345429. Siemens AG; Concord; CA. | Siemens Medical Solutions USA; Inc | 2 | Image Orientation Incorrect (Image Shifts) - 1) if the reference image has been calibrated for centering and the reference image is used for patient positioning; the image will shift when the positioning tools are used. 2) This problem occurs only when non-square reference images are used; such as when images are cropped. At this point; the images will refresh; and the center of the reference im | On 10/03/2007, the firm sent Consignees a Customer Safety Advisory Notice informing them of the two potential issues with the display of reference images used for position correction. Consignees will be informed when a solution to this issue has been validated.
Z-1769-2008 | Aug-20-2008 | 2008 | Siemens - Medical charged-particle radiation therapy system. Siemens Linear Medical Accelerators equipped with COHERENCE Therapist R2.0 and Oncologist R2.0. Product update; TH 030//07/S. The COHERENCE Therapist R2.0 software numbers are: 7358299; 7360857; and 8141835 and applies only to those parts used with the following devices: Coherence AG Therapist; Coherence Therapist System; Coherence Impression Therapist; PRIMEVIEW 3i; AG Therapist 3rd party V&R; Coherence Therapist 2.0 and PRIMEVIEW 3i 2.0. | Siemens Medical Solutions USA; Inc | 2 | Software. The firm issued an Advisory Letter Update instruction TH030/07/S as a result of a complaint that images viewed in the device may shift; which could result in mistreatment. The advisory alerts all its affected customers of the image issue to prevent potential mistreatment. | Customer Safety Advisory Notice was issued on 10/03/2007 to consignees, via certified mail. Consignees are strongly cautioned that they should not further correct the image alignment at this time as doing so will result in incorrect offset values which could lead to incorrect repositioning of the patient. Consignees are instructed to check the "Filter Images on Loading" and "Saving of Filtered Images" config settings, and apply any additional filters to the image BEFORE using the interactive shift or landmark registration tools. Do not correct any perceived issues. A software fix is being validated and will be available soon.
Z-0165-2009 | Oct-22-2008 | 2008 | Roche/Hitachi Modular E Module - Roche/Hitachi Modular E Module immunoassay analyzer; Roche Diagnostics; Ind.; IN; GMMI Nos. 04998642001 and 03617505001. | Roche Diagnostics Corp. | 2 | A software bug may result in pipetting from an incorrect reagent pack and/or assigning calibration curve parameters incorrectly. | The consignees were notified of the problem via letter dated 6/24/08 and were provided instructions for a workaround.
Z-2095-2008 | Sep-17-2008 | 2008 | Coherence AG Therapist - Radiation therapy system - Coherence AG Therapist Part number 5863506 and Therapist system Part number 7339125; equipped with a Coherence Therapist RTT2.1 or Primeview 3i R2.1. This is a medical charged particle radiation therapy system; part of the firm's beam limiting device/accessory. The product provides data processing; and the SYNGO standard medical imaging platform to provide comprehensive oncology workflow solutions for radiation therapy. | Siemens Medical Solutions USA; Inc | 2 | Software - Under certain circumstances; the portal image may be overwritten; which may lead to incorrect dosage. | An Urgent Medical Device Correction - Customer Safety Advisory Notice letter (TH014/08/S) was released on May 15, 2008. A new software patch is under construction and is planned for release in June 2008. If you have any questions contact Roland Betz at (925) 246-8407.
Z-0946-2008 | Mar-07-2008 | 2008 | HomeChoice Automated Peritoneal Dialysis Systems - Baxter HomeChoice Automated Peritoneal Dialysis Systems; Catalog Numbers: 5C4471; 5C4471R; 5C4474; 5C4474R; T5C4441; and T5C4441R; Manufactured for Baxter Healthcare Corporation; Deerfield; IL 60015 U.S.A.; Made in U.S.A | Baxter Healthcare Renal Div | 2 | Unrecorded Infusion; If the system is powered down or a power failure occurs during a fill cycle; the HomeChoice system may not record the last 1-4 pump strokes infused into the patient when the power is restored. This situation may create a potential unrecorded delivery of approximately 15-60 mL of fluid being infused into the patient. | Urgent Product Recall letters dated 11/20/07 were sent via first class mail to all of the affected customers, informing them of the potential for overfill if the system is powered down or a power failure occurs during a fill cycle. A Baxter representative will contact the customer to schedule the software update which corects this anomaly. The accounts were requested to complete and return the enclosed Customer Reply Form, indicating that they have disseminated the information to their staff and home patients. Any questions were directed to Baxter Global Technical Services at 1-800-553-6898, prompt 1. Baxter sent a follow-up letter dated 3/17/08 to the consignees, accompanied by a copy of the 11/20/07 letter. The letter reiterated the information in the 11/20/07 letter, emphasizing the serious consequences of an overfill for neonates and young pediatric patients, and providing information describing the symptoms of excess fluid in the abdomen and steps to be taken if the user suspects overfill of a neonate or young pediatric patient. Baxter requested that Baxter be contacted if an overfill should occur so Baxter can perform a complete investigation into the event. The accounts were requested to complete and return the enclosed Customer Reply Form, indicating that they have disseminated the information to their staff and home patients. Any questions were directed to Baxter Global Technical Services at 1-800-553-6898, prompt 1.
Z-1905-2008 | Sep-16-2008 | 2008 | LIFEPAK¿ 20 defibrillator/monitor - Physio-Control LIFEPAK¿ 20 defibrillator/monitor is indicated for the termination of certain potentially fatal arrhythmias; such as ventricular fibrillation and symptomatic ventricular tachycardia. | Physio Control; Inc. | 2 | The LIFEPAK 20 defibrillator / monitors with software versions 048; 052 and 054 have an increase in likelihood for an incorrect Shock Advisory Algorithm (SAS) decision if the Auto Analyze setting in AED mode is On; leading to incorrectly render of shock or no shock decision. When Auto Analyze is set to On in AED mode the device initiates the SAS analysis immediately (no waiting period or warning p | On April 16, 2008, a letter was sent to the firms customers, via certified mail explaining the recall, giving product and software version information, request for customer to forward the notification to all sites and recommending the following: "Keep the defibrillator in service... Check the Auto Analyze Setup Option for each of your devices. If Auto Analyze is set to ON, change the setting to one of the other two choices: OFF or After First Shock... Follow the enclosed instructions to confirm or change the Auto Analyze Setup Option setting...If you update your defibrillator with new software, confirm the Auto Analyze Setup Option is set to OFF or After First Shock". Instructions for changing the auto analyze setup option for the LP 12 and LP20 were attached to the letter. There was no software issued in response to this action. Since LP20 customers may purchase a software program allowing them to update their defibrillator software, the letter recommends confirming Auto Analyze Setup Option is set to OFF or After First Shock if new software is installed . The default setting for the software update is Auto-Analyze ON. Contact Technical Support at 1-800-442-1142, option 5 - 6:00 a.m. to 4:00 p,m, (Pacific), Monday - Friday or visit www.Physio-Control-notices.com/sas if you have questions.
Z-1904-2008 | Sep-16-2008 | 2008 | LIFEPAK¿ 12 defibrillator/monitor - Physio-Control LIFEPAK¿ 12 defibrillator/monitor is indicated for the termination of certain potentially fatal arrhythmias; such as ventricular fibrillation and symptomatic ventricular tachycardia. | Physio Control; Inc. | 2 | The LIFEPAK 12 defibrillator / monitors with software version 130 have an increase in likelihood for an incorrect Shock Advisory Algorithm (SAS) decision if the Auto Analyze setting in AED mode is On; leading to incorrectly render of shock or no shock decision. When Auto Analyze is set to On in AED mode the device initiates the SAS analysis immediately (no waiting period or warning prior to analys | On April 16, 2008, a letter was sent to the firms customers, via certified mail explaining the recall, giving product and software version information, request for customer to forward the notification to all sites and recommending the following: "Keep the defibrillator in service... Check the Auto Analyze Setup Option for each of your devices. If Auto Analyze is set to ON, change the setting to one of the other two choices: OFF or After First Shock... Follow the enclosed instructions to confirm or change the Auto Analyze Setup Option setting...If you update your defibrillator with new software, confirm the Auto Analyze Setup Option is set to OFF or After First Shock". Instructions for changing the auto analyze setup option for the LP 12 and LP20 were attached to the letter. There was no software issued in response to this action. Since LP20 customers may purchase a software program allowing them to update their defibrillator software, the letter recommends confirming Auto Analyze Setup Option is set to OFF or After First Shock if new software is installed . The default setting for the software update is Auto-Analyze ON. Contact Technical Support at 1-800-442-1142, option 5 - 6:00 a.m. to 4:00 p,m, (Pacific), Monday - Friday or visit www.Physio-Control-notices.com/sas if you have questions.
Z-0960-2008 | Feb-27-2008 | 2008 | Precision Implantable Pulse Generators - Advanced Bionics Precision Implantable Pulse Generator Kit; Model Number: SC-1110 with firmware version prior to Revision 3.02; Advanced Bionics Corportation; Sylmar; CA 91342 | Advanced Bionics Corp | 2 | Incorrect Data -- Corruption of internal memory component results in an inability for the physician to reprogram the IPG with firmware version prior to Revision 3.02. When this occurs; the IPG will report an error code of "10h0" or "00h0" through the Remote Control. Under this condition; the IPG will cease to log in some data that could be used for informational purposes. ** It should be noted t | Advanced Bionics sent an Important Notification letter on October 5, 2007, via FED-Ex or US Postal Service Certified mail. (International notification activity is in progress.) The recall letter advises that the firm has become aware of an issue with Precision Implantable Pulse Generators (IPG model numbers SC-1100, Sc1110, and ONS-1110) manufactured prior to July 2007. The letter describes the issue in detail. The firm is requesting that customer immediately locate the affected IPGs and set them aside so that they are not used. A Boston Scientific/Advanced Bionics sales representative will contact your facility to confirm these IPGs and arrange for the exchange of these devices. The firm enclosed Reply Verification Tracking Form listing the IPGs that are at your Clinic according to the firm's records. In addition, the firm recommends that if error code 10h0 or 00h0 is encountered in Implanted IPGs, it may be corrected by upgrading to a newer version of IPG firmware (software that runs in the IPG). The new version of firmware corrects this condition and helps prevent it from occurring in the future. This firmware upgrade is performed non-invasively, through wireless RF telemetry. If any of patients encounter error code "10h0" or "00h0" on their Remote Control, consignees are advised to contact their Advanced Bionics representative or Technical Services team at (866) 566-8913 to help identify whether the condition can be corrected through a firmware upgrade and, if so, they will help arrange for the upgrade. Please note that some devices may not be upgradeable. The firm request that all consignees return the notification acknowledgement form is included with the letter.
Z-0340-2009 | Nov-21-2008 | 2008 | STERRAD 100S Sterilizer - STERRAD 100S Sterilizer (Product Code 10101). The sterilizer includes software and hardware components. The STERRAD Sterilization System is a low-temperature; general purpose sterilizer used to sterilize heat and moisture sensitive reusable medical devices. | Advanced Sterilization Products | 2 | 1) Inability of the sterilizer to detect when an injectiontakes place without hydrogen peroxide being transferred to the vaporizer bowl. This situation can result in a cycle completion with the injection of insufficient hydrogen peroxide sterilant and 2) Inability of the sterilizer to detect an obstructionin the door travel path while the door is closing. This situation can result in a hand b | Customer notifications will be sent the week of November 24, 2008. The notification letter (URGENT: Product Correction) will advise of two situations that could cause the STERRAD 100S Sterilizer to malfunction. The first situation involves the failure of the STERRAD 100S System to detect when an injection takes place without hydrogen peroxide having been transferred to the vaporizer bowl. This can occur under combinations of the following conditions: - the pneumatic pressure is low on the machine, or - the injector valve has a small leak, AND - a cassette fails to advance, or - a used cassette is re-inserted into the sterilizer. This situation can result in a cycle completion with insufficient hydrogen peroxide. The second situation involves the failure of the STERRAD 100S System to detect an obstruction in the door travel path while the door is closing. Under some conditions, the door sensors may become mechanically unsecured/disconnected and will not detect an obstruction. This situation can result in an operator's hand becoming pinched and bruised while the door is closing. To correct both situations above, Advanced Sterilization Products(ASP) has revised the software and is making physical hardware additions to the STERRAD 100S System. Contact ASP Customer Care Center at 888-783-7723, option 2 for questions or assistance.
Z-0280-2008 | Feb-01-2008 | 2008 | CIC Pro v4.1 - GE Healthcare CIC Pro Clinical Information Center; Central Station; Software Version 4.1. (CIC Pro v4.1). GE Medical Systems Information Technologies; Inc.; Milwaukee; Wisconsin (System hardware consists of a processing unit and a monitor) | General Electric Medical Systems Information Technology | 2 | Two separate issues may occur with operating The CiC Pro v4.1 System with Telemetry. 1) Experience a loss of audible alarms when operating system memory resource reaches approximately 170MG. This will occur if CIC Pro continuously runs for average of 160 day; and 2) System may continuously reboot upon start up in an attempt to correct the issue. | GE Healthcare sent a certified letter titled "Urgent Medical Device Correction" to customers on 3/13/07. As a Short Term Solution, GE recommends that user reboot of system according to instructions and precautions in Service Manual at least once every three months. Also, take the necessary precautions to ensure that system is shut down properly. This will decrease the likelihood of a continuous reboot. If system continues to reboot, GE suggest that user contact Technical Support at 1-800-588-7044 (Press Option 2) and ask for Technical Support. As for a Long Term Solution, GE will provide a software update that will rectify the issue as soon as an upgrade is available.
Z-1167-2008 | Jun-28-2008 | 2008 | Zoll E Series Defibrillator with 12 Lead ECG - Zoll E Series Defibrillator with 12 Lead ECG; Automated external defibrillator; Software Version 3.00 and higher; Zoll Medical Corporation; Chelmsford; MA 01824. | ZOLL Medical Corporation; World Wide Headquarters | 2 | Incorrect Patient Records: Patient records that have been stored; then later transmitted or printed may not be the correct record. Records that are transmitted or printed immediately following acquisition are correct. | Consignees were notified by an Urgent Device Correction letter sent on 1/14/08. The letter instructed users to complete and return a response form. Upon receipt of the response form, the recalling firm will have sent a software update. The letter also provided instructions for users to follow in order to avoid the software error before the software update was received. For additional information, contact 1-800-348-9011.
Z-1218-2008 | Jun-25-2008 | 2008 | INTELECT INTELECT XT - INTELECT INTELECT XT; Model No. 2760; Chattanooga Group; Hixson; TN 37343; Made in the USA | Chattanooga Corp Inc | 2 | Device malfunction: causing electric shock and burn. | Beginning on 12/17/2007, Chattanooga Corp. sent two recall notification letters, via UPS mail, to its consignees informing them of the situation. The letters stated that the recall was to the user level and requested that the user perform the software upgrade, which will eliminate the possibility of shock and burn.
Z-0522-2009 | Dec-31-2008 | 2008 | Philips Diagnostic Ultrasound System - Philips Diagnostic Ultrasound System; Model HD3. Product labeled HD3 Ultrasound System; Distributed by Philips Ultrasound; 22100 Bothell-Everett Highway; Bothell; WA 98021.Pulse Doppler Imaging Systems used for diagnostic ultrasound imaging or fluid analysis of human body. | Philips Medical Systems | 2 | The system has a software defect that affects the display of Estimated Fetal Weight (EFW) Growth Percentile result values. | On 8/12/08, the firm sent an Urgent - Medical Device Correction letter, dated July 2008, to all affected customers via UPS. The letter describes the reason for the recall, informs them that Philips Ultrasound is correcting the error, states that the software is under development, will be available in September 2008 and will contact the customer to upgrade their system. "In the interim, Philips recommends that you only use the Hadlock or Osaka growth tables to perform your EFW Growth Percentile calculations. You should make the appropriate selections in the Analysis Setup menu on your system.".
Z-1776-2008 | Aug-27-2008 | 2008 | Q-Logic Controller Software - Q-Logic Controller Software used on the following power wheelchairs: Q6000; Q6000XL; Q6000Z; R4000; Q600; Q600XL; Q610; Q1107; Q600E; and Dynamo 1107 | Pride Mobility Products Corp | 2 | Software - Watch dog timer feature was disabled. | The recalling firm issued an Urgent Device Correction letter on 4/11/08. This letter instructs the dealers to contact the end users so that the units can be upgraded to enable the watch dog feature.
Z-0519-2008 | Jan-17-2008 | 2008 | Pro - GE Medical Systems Information Technologies Clinical Information Center (CIC) Pro v5.0.3 & v5.0.6 systems; GE Healthcare; Clinical Systems | General Electric Medical Systems Information Technology | 2 | Potential for Loss of Audible Alarming and ECG Tab Setting Changes; from Clinical Informaiton Center (CIC) Pro v5.0.3 & 5.0.6 when used in combination with Telemetry and bedside monitors. | A "GE Urgent Medical Device Correction" letter was sent to consignees on 10/16/2007. The letter described the problem and product involved. The letter also gave short term recommendations. A long term solution is being worked on by GE and will be announced when determined.
Z-2468-2008 | Sep-29-2008 | 2008 | Sonoline G50/G60 Ultrasound System - Sonoline G50/G60 Ultrasound System; Software versions 1.0.0 to 1.5.0b. The product is a general purpose; mobile software controlled diagnostic ultrasound system. | Siemens Medical Solutions USA; Inc. | 2 | As a result of bugs and calculation errors; the system may display incorrect or fail to display mechanical Index and thermal index values. | The firm initiated its field correction on 08/05/2008. The firm's service personnel will make site visits to correct the problem and install the revised software (version 1.5.0c). Contact Siemens Medical Solutions USA, Inc. at 1-650-694-5993 for assistance.
Z-2231-2008 | Sep-11-2008 | 2008 | ARCHITECT c8000 Clinical Analyzer - ARCHITECT c8000 Processing Module; for in vitro diagnostics; List Number 1G06-01 and 1G06-11; Affected software: v2.20; v2.20DB; v2.60; v3.10; 3.11 and 3.12; distributed by Abbott Laboratories; Irving; TX.The Abbott ARCHITECT cSystem is designed to perform automated; random access; clinical chemistry analyzer which utilizes spectrophotometry (monochromatic and bichromatic modes of measurement) for photometric based determinations. | Abbott Laboratories; Inc | 2 | Software defect can allow tests ordered for one sample to be aspirated from a different sample; reporting erroneous results for the affected sample. | Firm notified all ARCHITECT c8000 customers with a correction letter and response form distributed on 5/23/2008. The letter provided actions the users could take until an updated software version was available. For assistance, contact Abbott at 1-877-4ABBOTT.
Z-1161-2008 | Jun-24-2008 | 2008 | S/5 iCentral - S/5 iCentral; Cardiovascular Monitoring Devices; Versions L-NET03; L-NETC03; L-NET05 and L-NETC05; GE Healthcare; Wauwatosa; WI 53226. | GE Healthcare | 2 | Failure to alarm: An active monitor could become disconnected from the unit without any notification or alarm. | Consignees were notified via a Medical Device Correction letter on 12/13/07. The letter explained the circumstances in which the "connection to monitor loss" alarm would not display, if a monitor was disconnected. The letter also included an attachment entitled "How to ensure a 'Connection to monitor lost' alarm on iCentral." For additional information, contact 1-800-558-7044.
Z-2404-2008 | Sep-22-2008 | 2008 | GE Voluson¿ E8 Ultrasound System - GE Voluson¿ E8 Ultrasound System with software versions 6.0.0; 6.0.1;6.1.0; 6.2.0; 6.2.1; 6.2.2; 6.2;3; 7.0.0; and 7.01.Intended for use by a qualified physician for ultrasound evaluation of Fetal?OB; abdominal (including GYN; pelvic and infertility monitoring/follicle development); Pediatric; Small Organ (breast; testes; thyroid etc.); Neonatal and Adult Cephalic; Cardiac (adult and pediatric); Musculo-skeletal Conventional and Superficial; Peripheral Vascular; Transvaginal (TV); Transrectal (TR); and Intraoperative (abdominal ; PV neurological). | GE Healthcare | 2 | Inaccurate dimension measurements in M-Mode when pan-Zoom is activated in GE Voluson? E8 Ultrasound Systems. Of particular concern are potential inaccuracies of fetal cardiac measurements that may impact patient safety. | Consignees were sent a GE "Urgent Medical device Correction" letter dated April 1, 22008. The letter was addressed to Healthcare Administrator/Risk Manger; Chief of Intensive Care and Director of Clinical Engineering. The letter described the safety issue; affected product details; safety instructions stating that use of the system may be continued and measurements may be made safely in M-Mode by using HD-Zoom rather than pan-Zoom; product correction stating that a GE Healthcare Rep will contact consignees in the future to schedule software updates to resolve the issue at no charge; and contact information.
Z-1695-2008 | Aug-20-2008 | 2008 | AMO WaveScan WaveFront System v3.65 - Excimer Laser System; WaveScan WaveFront System; Software Version 3.65; for use with the VISX STAR S4 and STAR S4 IR Excimer Laser Systems. System is a diagnostic instrument intended for the automated measurement; analysis and recording of refractive errors and wavefront aberrations to the target shape and then generates the commands for the VISX STAR laser to create the target shape on the cones. | VISX INCORPORATED; A SUBSIDIARY OF AMO INC | 2 | Firm became aware of a significant overcorrection in one eye of a patient who had received a wavefront-guided LASIK treatment at a facility in Germany. The probability of a recurrence of this error was determined to be extremely remote. To ensure that this type of incident will not happen again; a software update was developed. | VISX issued a notification letter on November 27, 2006, to all distributors and direct accounts. The letter informed the recipients that they were performing a field correction in the form of a mandatory upgrade on all VISX WaveScan WaveFront Systems and Offline Programming Modules to correct an issue. The field upgrade should be made by February 16, 2007. A Dear Doctor letter was also issued instructing surgeons to upgrade all Offline Programming Modules using an enclosed CD.
Z-2212-2008 | Sep-19-2008 | 2008 | ACUSON CV70 Ultrasound systems - Acuson CV70 Ultrasound systems Ultrasonic Pulsed echo/doppler imaging system with Software revisions less than 4.0.0b; material number 784539; 7848521; and 7848547.Is a general purpose; mobile; software-controlled; diagnostic ultrasound system with an on-screen display for thermal and mechanical indices related to potential bioeffect mechanisms. | Siemens Medical Solutions USA; Inc. | 2 | System may either display incorrect mechanical and thermal index values; or fail to display them altogether. | Recall initiated on 06/04/2008. Under the Siemens update program, all consignees will receive a visit from Siemens service personal to install the revised software version. Siemens staff will coordinate with foreign distributors to arrange for service abroad.
Z-0570-2008 | Mar-04-2008 | 2008 | MultiSET - BD Biosciences MultiSET Flow Cytometry Software Versions 2.1; 2.1.1; and 2.2 with OS X operating systems; BD Biosciences; San Jose; CA 95131. | BD Biosciences | 2 | Inaccurate results: Software error results in inaccurate display result statistics. If the user adjusts the lymph gate or attractors in the Lab Report view without using the Manual Gate function (as described in the system manual); the statistic results will not be updated to match the adjusted gate. | Becton Dickinson sent an Important Product Information letter by mail on 09/14/2007 informing customers of the need to follow the instructions for adjusting the lymph gate or attractors. The next software revision will alleviate this limitation.
Z-1525-2008 | Aug-16-2008 | 2008 | BD Biosciences FACSDiva Software Versions 4.0.1 to 5.0.2 - FACSDiva Software Versions 4.0.1 to 5.0.2; Flow Cytometry Acquisition & Analysis software. Catalog number 641609Flexible data acquistion and analysis package specifically designed for digital-based flow cytometers. | BD Biosciences | 2 | When data file containing one or no fluoresence parameters is exported; the software will automatically apply compensation to this file and all subsequently exported files. | The firm has issued a customer Important Product Information letter via mail on 04/20/2007 to inform them of the software issues. The letter states that to avoid this problem, make sure each tube or well has at least two fluorescence parameters before exporting FCS 3.0 data files. Manually exporting tubes or wells one at a time or restarting the software after the export of an FCS data file containing one or no fluorescence parameters will prevent the problem from occurring. Customers are requested to fill out and return the response form and to contact BD Customer Support Center if there are any additional questions. A software fix is being developed in version 6.0 and a free maintenance release with no product enhancments (version 5.0.3).
Z-2037-2008 | Jul-30-2008 | 2008 | Centricity Perinatal (formerly QS) System - PFILS software - GE Healthcare Centricity Perinatal (formerly QS) System - PFILS (Patient File Server) software; automatic patient data management providing clinical information at the bedside in Labor & Delivery; Mother-Baby and the Neonatal Intensive Care Unit. | GE Healthcare Integrated IT Solutions | 2 | Patient information mix-up: When the PFILS application; configured with a backup server has experienced a network interruption; PFILS verify could result in the Centricity Perinatal application recording non-identifiable patient information to the incorrect patient file. | GE Healthcare notified their Centricity Perinatal customers via letter dated 7/20/07 of the problem with the PFILS application that could result in the Centricity Perinatal application recording non-identifiable patient information to the incorrect patient file. The letter provided the users with a workaround and provided them with a Frequently Asked Questions document which listed the availability of a software patch to fix the problem. A follow-up letter dated 6/27/08 was sent on 6/30-7/2/08 to those accounts who had not requested a PFILS Hotfix CD-ROM or had not had their software recently upgraded. The letter a copy of the 7/20/07 letter, the most recent FAQ, the PFILS Hotfix CD (PFIL67_66) and Installation Procedure for the PFILS Patch instructions for installation of the software patch. The accounts were advised that they can request an analysis of existing patient records to determine whether they have been affected by the PFILS issue, by submitting a non-critical Technical Support Request on the GE Centricity Perinatal Web site: http://www.gehealthcare.com/usen/security/cis/stupgrades.html, indicating "PFILS ANALYSIS REQUEST" in the Description of Issue field. Any questions were directed to Technical Support at 800-433-2009.
Z-1906-2008 | Sep-16-2008 | 2008 | LIFEPAK¿ 20e defibrillator/monitor - Physio-Control LIFEPAK¿ 20e defibrillator/monitor; sold internationally only; is indicated for the termination of certain potentially fatal arrhythmias; such as ventricular fibrillation and symptomatic ventricular tachycardia. | Physio Control; Inc. | 2 | The LIFEPAK 20e defibrillator / monitors with software versions 058 have an increase in likelihood for an incorrect Shock Advisory Algorithm (SAS) decision if the Auto Analyze setting in AED mode is On; leading to incorrectly render of shock or no shock decision. When Auto Analyze is set to On in AED mode the device initiates the SAS analysis immediately (no waiting period or warning prior to anal | On April 16, 2008, a letter was sent to the firms customers, via certified mail explaining the recall, giving product and software version information, request for customer to forward the notification to all sites and recommending the following: "Keep the defibrillator in service... Check the Auto Analyze Setup Option for each of your devices. If Auto Analyze is set to ON, change the setting to one of the other two choices: OFF or After First Shock... Follow the enclosed instructions to confirm or change the Auto Analyze Setup Option setting...If you update your defibrillator with new software, confirm the Auto Analyze Setup Option is set to OFF or After First Shock". Instructions for changing the auto analyze setup option for the LP 12 and LP20 were attached to the letter. There was no software issued in response to this action. Since LP20 customers may purchase a software program allowing them to update their defibrillator software, the letter recommends confirming Auto Analyze Setup Option is set to OFF or After First Shock if new software is installed . The default setting for the software update is Auto-Analyze ON. Contact Technical Support at 1-800-442-1142, option 5 - 6:00 a.m. to 4:00 p,m, (Pacific), Monday - Friday or visit www.Physio-Control-notices.com/sas if you have questions.
Z-1243-2008 | Sep-11-2008 | 2008 | IMPAX - IMPAX¿ 4.1; IMPAX ¿ 4.5; IMPAX ¿ 5.0; IMPAX ¿ 5.1 IMPAX ¿ 5.2; IMPAX ¿ 6.0; IMPAX ¿ 6.2; Picture Archiving and Communications systems (PACs); used in the acceptance; transfer; display; storage and digital processing of medical images. | AGFA Corp. | 2 | Java version 1.6x (JAVA 6) is not compatible with AGFA IMPAX Administration/Service Tools. Due to a Software Mismatch data can be lost. | Agfa Healthcare contacted all consignees on 5/14/2007 and by registered mail on 06/07/2007. The registered Urgent Field Safety Notice letters requested the accounts take the following steps immediately: 1. Turn off automatic JAVA updates to avoid this update being applied in the background. 2. Do not manually update JAVA on their station-even for other applications. 3. If JAVA 1.6 has already been applied, do not use Service/Admin Tools. They were instructed to contact their IT department or Agfa support for assistance in backing out of this version to apply the required version for the Agfa software. A Mandatory Service Bulletin will inform and assist affected account in removing JAVA 1.6 if it has been applied and turning off the automatic updates via the JAVA Plug-in Control Panel.
Z-2308-2008 | Sep-17-2008 | 2008 | Eclipse V. 6.5 - Varian Medical Systems Eclipse device; model H48; Version 6.5 is used to plan photon and electron radiation therapy treatments. | Varian Medical Systems Oncology Systems | 2 | Software anomaly in which swapping IDs of planning fields can produce inconsistencies between dose matrix and field parameters. | Recall initiated June 11, 2008 when the firm issued an Urgent Device Correction recall letter. The letter was distributed to all affected users, with problem description and user corrective action steps. If you have questions, contact your local Varian Customer Support Representative.
Z-1394-2008 | Apr-02-2008 | 2008 | Mevatron/ PRIMUS Medical Linear Accelerator - Mevatron M2/PRIMUS MID Medical Linear Accelerator; Material Number: 019490035; with control console software versions 6.5/7.2; 8.0; 9.0 and 9.1; Manufactured by Siemens Medical Solutions USA; Inc; Concord; CA | Siemens Medical Solutions USA; Inc | 2 | Linear accelerator gantry may rotate in an unexpected direction and may cause patient injury on impact. | This is a mandatory software upgrade for consignees. The firm's representatives have contacted consignees via Urgent Device Notification letters dated February 14, 2008, notifying them of the new update and making arrangements to install it.
Z-2286-2008 | Sep-17-2008 | 2008 | Piccolo MetLyte 7 Reagent Disc - Piccolo Chemistry Analyzer software version 6.121 with Comprehensive Metabolic Lot #8111AA5 ; Model Number 400-1028 in polyfoil pouch. Piccolo Comprehensive Metabolic Panel is a single-use self-contained reagent disk; intended for use with the Piccolo and PiccoloXpress chemistry analyzers; for in vitro qualitative determination. | Abaxis Inc | 2 | The product will recover tCO2 at approximately 5-7 units high relative to some controls and patient samples | The firm issued a notification letter to its consignees on April 21, 2008. The firm intends to follow up with phone calls and provide a software upgrade for the products. Contact Abaxis Customer Service at 1-510-675-6500 for assistance.
Z-1307-2008 | Jul-02-2008 | 2008 | IMPAX ¿ 4.5; IMPAX ¿ 5.0 OT3000 Orthopedic Client; IMPAX¿ 5.0 Orthopedic Client; IMPAX¿ ES for Radio - Impax 4.5; Impax 5.0; Impax 5.2 and Impax 5.2x Picture Archiving Systems (PACs) with integrated Voxar 3D software; and Impax 5.0 and Impax 5.2 PACs with integrated MediCAD software. The softwares are used to provide digital X-ray images for pre-surgical orthopedic planning. AGFA Corp.; Greenville; South Carolina 29601 | AGFA Corp. | 2 | Misidentification: The orthopedic planning X-ray images for one patient are misidentified as the images for another patient. | On 2/15/07, AGFA notified their consignees of the situation by sending them "Urgent Medical Device Correction" letters, instructing them on how to work around the problem until the firm was able to install the upgraded software that corrected the problem. On 12/11/2007, the firm reported that the upgraded software had been installed at all affected sites.
Z-2203-2008 | Sep-16-2008 | 2008 | 740 Series Multiparameter Monitor - 740 Series Multiparameter Monitor Software Version 3.0Model Numbers: 740-1; 740M-1; 740-2T; 740M-2T; 740-2MS; 740M-2MS; 740-2NL; 740M-2NL; 740-2NN; 740M-2NN; 740-3MS; 740M-3MS; 740-3NL; 740M-3NL; 740-3NN; and 740M-3NNCAS Medical Systems; Inc.; Branford; Connecticut 06405Indicated for use for non-invasive monitoring of blood pressure; oxygen saturation; pulse and temperature of adults; pediatric and neonatal patients; in the care of health professionals. | CAS Medical Systems; Inc. | 2 | Software: Audible alarm may be silenced if changed from the factory setting | CASMED notified customers by a Product Correction Notification letter on 6/24/08 to advise of problem of the audible alarm and a software upgrade to correct the problem. The firm requests that the alarm volume not be adjusted from its factory set value to ensure that alarms sound correctly. Also, the letter states that if the volume needs to be adjusted, that customers needs to be sure that they only select two clicks above, or two clicks below the mid-point, and that there is a corresponding beep when you make your selection. Customers are asked to review the attached list of serial numbers that will need to be updated with new software and to then email the a contact at the firm once updates have been completed. All updates must be completed by August 31, 2008.
Z-1316-2008 | Sep-11-2008 | 2008 | Stryker Digital Capture System Ultra (SDC Ultra) - Stryker Digital Capture System Ultra (SDC Ultra); Model Number 240-050-988; Stryker Endoscopy; image and video archiving system. | Stryker Endoscopy | 2 | Patient Identification numbers and names may be applied to images from another patient due to a software issue. This failure mode can only occur when the pre-load function of the SDC ultra is used. | Stryker Endoscopy mailed Urgent Device Correction notices, dated February 14, 2008, to all consignees, and sales representatives will be supplied with a fix to distribute to the user base.
Z-0088-2009 | Oct-07-2008 | 2008 | Envision E700 Low Airloss Therapy Surface - Envision E700 Low Airloss Therapy Surface; Made By Hill-Rom Manufacturing; Inc.; Charleston; SC.The Envision E700 Low Airloss Therapy Surface helps prevent and treat stage III and stage IV pressure ulcers in patients who weigh between 70 lb and 400 lb and are between 4'11" and 6' 4" in height. | Hill-Rom Manufacturing; Inc. | 2 | A defect in the software of the device may not allow the patient bed exam alarm or the patient movement alarm to function correctly. | An internal Modification Notice dated 9/1/08 was sent via email to Service Centers (because the device is owned by the manufacturer, Hill-Rom, each Service Center in possession of the device will be the consignee). Each Service Center will perform the correction on the units affected. Each device in the affected serial number range will be tested for the software defect. If the defect is found, the circuit board that contains the defective software will be replaced. The Field Action Plan is expected to be completed on July 24, 2008. Expected completion date of the Correction will be 12 months from July, 2008.
Z-0660-2008 | Jan-31-2008 | 2008 | Ultraview Multigas Analyzer - Spacelabls Medical; Ultraview Multigas Analyzer; Model 91518;Spacelabs Healthcare; Issaquah; WA; 98027 | Spacelabs Healthcare; Incorporated | 2 | Inaccurate partial pressure readings; readings are not compensated for ambient pressure in mmHg or kPa. | The firm sent out their Urgent Medical Device correction/recall letters on 10/5/2007, and an updated versions on 11/30/07 and 12/7/2007, forwarding an attached set of work around directions for use and associated tables for calculations. This short term fix will be updated to a permanent fix when a new software revision is released and installed at no cost to each user as noted in the recall notification. Telephone contact was provided at 1-800-522-7025.
Z-2225-2008 | Sep-11-2008 | 2008 | Signa HDX MR System - GE 1.5T and 3.0T Signa HDX MR SystemModel Number(s): 2226300; 2377062-61; 2395001; 2395001-2; 5107849; and 5127452. This system is indicated for use as a diagnostic imaging device to produce axial sagittula; coronal and oblique images; spectroscopic images; and or spectra; dynamic images of the internal structure or organs of the entire body; | GE Healthcare | 2 | Possible misdiagnosis or inappropriate treatment based on the image. Signal homogeneity may be affected when BRAVO applications are set with IR Prep off. The concern is that signal inhomogeneity can result in an erroneous diagnosis. | Customer Letter and Operator Manual addendum were sent directly to customers in August 2008. When the software correction is issued, GE Field Engineer will visit each site to implement the correction. (Not conducted yet).
Z-0368-2008 | Sep-29-2008 | 2008 | OEC 9900 - GE Healthcare OEC 9900 mobile fluoroscopic x-ray system | OEC Medical Systems; Inc | 2 | Possible overexposure: Failure to apply the entrance exposure rate (EER) tube current limit calibration to the automatic exposure rate (AER) control system when the anatomical profile mode is changed from the default selection to another selection. Manual technique modes are not affected. | GE OEC sent a letter to customers December 15, 2007, and scheduled field representative visits for corrective software solution installation.
Z-1104-2008 | May-01-2008 | 2008 | GE Centricity PACS RA1000 Workstation - GE Centricity PACS RA1000 Workstation; for diagnostic image analysis. GE Healthcare Integrated IT Solutions; Barrington; IL 60010 | GE Healthcare Integrated IT Solutions | 2 | Report missing characters: Four special characters; when entered into the exam notes; are not transferred to the preview panel or hard copy printout (greater than [>]; less than [<]; quotation mark ["]; and apostrophe [']). This may not be evident to the end user; as the symbols only appear when the Exam Note window is open. | GE Healthcare IITS sent Product Safety Notification letters dated 12/28/07 to all customers who have the GE Centricity PACS RA1000 Workstation software versions 2.1.X, or 3.0.X, informing them of the potential patient safety issue involving the exam notes window, which may be used to provide various patient information including lab results. To mitigate this problem, the users were instructed not to use the following four special characters (greater than sign >, less than sign <, double quotation mark " and apostrophe ') in the exam notes field in order to prevent any adverse patient events. For example, if a lab result was entered as GFR <60, it should be typed as GFR less than 60. The customers were requested to make sure that all necessary physicians and personnel at their location are made aware of the problem and the recommended actions provided. A software patch to permanently resolve the issue will be installed at no charge when it becomes available. Any questions were directed to the GE Customer Care Center at 800-437-1171.
Z-0460-2009 | Dec-13-2008 | 2008 | GE Centricity PACS RA1000 Workstation - GE Centricity PACS RA1000 Workstation (for diagnostic image analysis). GE Healthcare Integrated IT Solutions; Barrington; IL 60010.The Centricity" PACS Workstation is intended for use as a primary diagnostic and analysis tool for diagnostic images by trained healthcare professionals. It is also intended for use as a clinical review workstation throughout the healthcare facility. The workstation interface provides the user with a means to display; manipulate; archive; print; and export images when connected with the Centricity" PACS infrastructure. | GE Healthcare Integrated IT Solutions | 2 | Software anomalies result in patient safety issues involving patient jacket content intermittently becoming unintentionally out of synchronization with the images being displayed; and involving Default Display Protocols (DDPs) which are used to layout images when displaying a study. If the user does not check the Study Date Time of the exam on the image title bar; then they may interpret the curr | GE Healthcare Integrated IT Solutions notified their Centricity PACS RA1000 Workstation customers via letter (Product Safety Notification) dated 9/25/08 of the software anomalies resulting in patient safety issues involving the patient jacket content intermittently becoming unintentionally out of synchronization with the images being displayed, and involving Default Display Protocols (DDPs) which are used to layout images when displaying a study. If the user does not check the Study Date Time of the exam on the image title bar, then they may interpret the current exam as a historical study and vice versa. The letter provided the users with safety instructions to follow until the software is updated. Direct any questions to the Remote Online Center, Field Engineer or the GE Customer Care Center at 800-437-1171. The accounts were requested to complete the enclosed customer reply form, supplying the workstation software or build number, and fax the completed form to 847-277-5240.
Z-2184-2008 | Sep-17-2008 | 2008 | Hologic Discovery Bone Densitometers - Hologic Discovery Bone Densitometers with Apex software versions 2.0; 2.1 and 2.2 is indicated for the estimate of bone mineral density; comparison of measured variables obtained from a given scan to a database of reference values; the estimation of fracture risk; vertebral deformity assessment; body composition analysis; and discrimination of bone from prosthetics. | Hologic; Inc. | 2 | Software densitometer readings for left hip and lumbar spine under certain conditions may be inaccurate. | Hologic notified customers by registered letter titled Urgent: Medical Device Recall, dated 6/27/08, providing instructions to update the software to Version 2.3 and discard version APEX 2.0, 2.1 and 2.2 manuals. If you have questions or require assistance with the software installation, contact Hologic's Help Desk at 1-800-321-4659.
Z-2016-2008 | Jul-11-2008 | 2008 | FC 500 with UPS - Cytomics FC 500 Flow Cytometry System with CXP Software Version 2.1 and 2.2 - FC 500 with UPS; Automated differential cell counter; Part Number: 626553; Beckman Coulter; Miami; FL 33196. | Beckman Coulter Inc | 2 | Incorrect Display: In certain modes; the protocol will not initially display the correct data in some plots. In addition; when running EWL files; any edited SampleID1 is not automatically updated in all tubes in the panel. | Consignees were notified of this safety recall by an Urgent Product Corrective Action letter sent on 5/3/07. The letter provided instructions on how to ensure correct data acquisition. Users were requested to return a response form acknowledging receipt of the letter. For additional information contact 1-800-526-7694.
Z-0112-2009 | Oct-15-2008 | 2008 | Centricity Perinatal (formerly Quantitative Sentinel) System - Fluid Total Precision - Centricity Perinatal (formerly Quantitative Sentinel) System - Fluid Total Precision software; automatic patient data management providing clinical information at the bedside in Labor & Delivery; Mother-Baby and the Neonatal Intensive Care Unit; GE Healthcare Integrated IT Solutions; Barrington; IL 60010 | GE Healthcare Integrated IT Solutions | 2 | On the I&O chart; the IN; OUT and NET fluid totals values will not honor the numeric precision configuration; always displaying a whole number; losing decimal point accuracy. | GE Healthcare notified their Centricity Perinatal customers via letter dated 4/20/07 of the problem with the Fluid I & O Module when fluid precision is required. The letter instructed those accounts that require fluid precision to contact GE Technical Support at 1-800-433-2009 to obtain a software patch to fix the problem. A follow-up letter dated 8/29/08 was sent on the same date to the accounts informing them that GE had re-classified the issue as a class II safety issue and that all customers need to have their software updated. The accounts were requested to contact GE Tech Support at 800-433-2009 to schedule a time to upgrade their system to 6.70.6 P02 as soon as possible. The accounts were requested to complete the enclosed customer reply form, supplying the workstation software or build number, and fax it back to 847-277-5240.
Z-1882-2008 | Sep-25-2008 | 2008 | GE Definium 8000 - GE Healthcare Definium 8000 Digital Radiographic System. | GE Healthcare | 2 | Radiation dose reading greater than actual dose: Error in the generator software on the GE Healthcare Revolution XR/d configured with HP 8200 PC consoles and all Definium 8000 systems may impact patient safety. Occasional generator software errors may cause light X-ray Images with reported mAs readings higher than was actually exposed to the patient using the overhead X-ray tube. The mAs rea | GE sent a Product Safety Notification dated March 7, 2008 to all customers. The letter described the safety issue, affected product, and stated that a GE Healthcare Field Engineer will schedule a field visit to install software updates.
Z-0207-2009 | Oct-31-2008 | 2008 | Sunquest Laboratory - Sunquest Laboratory versions 5.3; 5.4; 6.1; 6.2 and 6.3.Sunquest Laboratory Information System is comprised of several modules which are involved in aspects such as blood banking; collection management; specimen tracking; and use with microbiology and molecular diagnostics. Key capabilities include:1) Autoverify results directly from analyzers and reference laboratories2) Track tubes and containers from order to final disposition3) Monitor key metrics such as turnaround time and utilization using management reports4) Produce pathology results with Structured Reports5) Suggest microbiology workups online according to defined SOPs and rules6) All data entered into the system is accessible through audit trails and Ad Hoc Report tools7) Any change to patient results (inquiry; reports) is logged and available for review | Sunquest Information Systems; Inc. | 2 | Sunquest has identified a problem with IGO interfaces when performing an Online File Cleanup (OFC) that can lead to the test results of one patient filing to another patient when the order codes are the same. Performing an OFC prior to all results being assigned an accession number may cause the patient results to remain in a temporary file where they can be incorrectly associated to the wrong pat | Customers were sent a preliminary notice on 4/17/2008 followed by customer notice PSN-08-04 on 05/07/2008 via email. The notice advises that this notice is follow-up to the Preliminary notice distributed on 4/17/08 and applies to all clients who have an Instrument Generated Order (IGO) interface. Sunquest has identified a problem with IGO interfaces when performing an Online File Cleanup (OFC) that can lead to the test results of one patient filing to another patient when the order codes are the same. Performing an OFC prior to all results being assigned an accession number may cause the patient results to remain in a temporary file where they can be incorrectly associated to the wrong patient. When this occurs, Sunquest generated reports and inquiry contains the incorrect results. The letter provides the various steps that cause the issue to occur. Sites Potentially Affected by this Problem: Sunquest Laboratory using Instrument Generated Orders (IGO) Interface, Version 5.3 and later Action Sunquest Information Systems is taking to eliminate this Problem: Change Request 1-62VWLJ has been created to address this issue. This will be available as part of Sunquest Laboratory Instrument Build version 1.6.0023 and later. To request the software correction, customers can create a Service request via the Client support Webpage at www.sunquestinfo.com/supportweb. The notice also provides temporary procedural instructions: Effective immediately, do NOT perform Online File Cleanup (OFC) on any IGO instrument interface methods. -Manual OFC should not be performed using functions OFC, or Instrument Maintenance (IX or IXM), option 11 for IGO interfaces only. -If using function Instrument Maintenance (IX or IXM), option 1, 2, 2, Method, 2 (Automatic File Cleanup). This should be set to "N". Then stop and re-start the Instrument Results Processor. Note: Online file cleanup for non-IGO interfaces needs to be performed as usual. -Pending and Overdue Logs need to be monit.
Z-0297-2008 | Jan-09-2008 | 2008 | ACUSON Antares Ultrasound sytem and premium edition - Siemens; ACUSON Antares; Ultrasound System (Standard and Premium Edition); 5.0 Systems with software version 200.0.054; P/N : 10032747 and 10032746 standard; 10032746; 10037591; and 10038202 for premium edition; Siemens; Ultrasound Division; Mountain View; CA 94039 | Siemens Medical Solutions USA; Inc. | 2 | Image errors: A software problem results in on-screen indications (graphics) that lead the user to believe that the patient's right and left and the transducer orientation are oriented the same. This is not the case and the actual patient's right and left are reversed from the on-screen indicators. | The firm sent a recall letter, Urgent Medical Device Notification, to consignees on 07/10/2007, to be followed by a service visit by Siemens personnel for a mandatory software update to be installed. Short term the users were instructed to avoid increasing the SieClear compounding levels to an Advanced SieClear compounding setting after an L/R invert action with Biopsy guidelines on-screen.
Z-2053-2008 | Sep-16-2008 | 2008 | EasyVision RAD - EasyVision RAD - Radiological Image Processor; software versions: SW R4 2V2L2; SW R4 2V2L3; SW R4 2V2L4; SW R4 2V2L5; SW R4 2V2L6; and SW R4 2V2L7. The EasyVision RAD consists of a computer with digital interfaces; one or two (optional) monitors with keyboard and mouse; and software. The viewing; analysis; processing and printing functions are sent to the EasyVision by means of PMSNet (Philips Medical Systems communication protocol). The main image source for the PCR (Philips Computed Radiography) system is the CR-Reader. Remote database access and image transfer functions are implemented using the DICOM Query/Retrieve and Storage services. The EasyVision RAD provides the following features: receives images sent to it by remote applications (e.g. workstations or imaging modalities) and stores them in a local database; allows the operator to copy images from the local database to remote databases and vice versa; allows a remote system to query the EasyVision local database and to retrieve images from it; allows the operator to view; analyze; process and print the images stored in the local database. Operator may print images stored in the database on a DICOM printer; and is able to read and write DICOM CD-R disks. The device allows processing of and access to the stored raw data so that an image can be post-processed using different image processing algorithms. | Philips Medical Systems North America Co. Phillips | 2 | Inaccurate measurements when exporting radiography images to Picture Archiving and Communication System (PACS ). | On 03/14/08 the firmed mailed certified letters to the customers of the device. The letter advises of the problem and stated that images exported from the Easy Vision RAD workstation should not be used for measurements. The letter also states that the firm will be performing upgrades of the devices. If you have questions, contact Sarah Baxter at 425-487-7665 or the Philips Call Center at 800-722-9377, #5, #2.
Z-0690-2008 | Jan-31-2008 | 2008 | GE Innova - GE Innova 2100IQ Versatile Cardiac and Vascular Imaging System; GE Healthcare; Waukesha; WI 53188 ( The system consists of an a monoplane positioner; a vascular or cardiac table; an X-ray system and a digital detector) | GE Healthcare | 2 | Frozen Image Problems: During an acquisition (fluoro and/or record); there have been cases reported where an image became "frozen" on the DL (digital leader acquisition system) live monitor screen. In such cases; the system continued to send out X-rays without reporting an error message. The result was that the live imaging screen displayed an older "frozen" image until the operator released the | Customers were contacted by means of an Urgent Medical Device Correction Letter on November 9, 2007. The letter described the product and the safety issue. The letter also made short term recommendations and described a long term solution, which is pending as an expected hardware and/or software modification that will be made at no cost. Short Term Recommendation: Take extra care when working with motionless anatomy or devices where there is little evident motion to warn you of the frozen image occurrence. Always completely shut down and restart the system (not just reset) immediately after such an error occurs. Also, please immediately report this issue to your local GE Healthcare field service engineer.
Z-0150-2008 | Jan-17-2008 | 2008 | Panorama - Datascope; Panorama Patient Monitoring Network; Panorma Telepack 608;Software Versions 8.1.X; 8.3; 8.3.1; 8.4.1; and 8.5.X; Part Number: 0998-00-0191-04; Datascope Corp.; Mahwah; NJ | Datascope Corporation | 2 | Display problems: If an ECG cable; which has been damaged due to the ingress of liquid or by mechanical trauma; is utilized with the Telepack; it may cause the Telepack to switch between the 3-lead and 5 lead-input modes. During the switching process; the digital heart rate displayed at the Panorama Central Station will be frozen; and subsquent arrthymia alarms will not be announced. | Urgent Product Field Correction notices were sent to hospital administrators on October 9, 2007, by certified mail, return receipt requested. The letter describes the scenario that causes the problem and its remediation It also describes how the software upgrade will fix the problem and alerts hospitals that a Datascope Service Representative will be contacting them to arrange for the upgrade.
Z-2329-2008 | Sep-22-2008 | 2008 | SoftPath GUI Release Software Versions 3.17 and 4.1 - SoftPath GUI Release Software Versions 3.17 and 4.1. Used in pathology laboratory for receiving accessing specimen reports; billing and management reports for surgical; GYN; NON-GYN and Autopsy cases. | SCC Soft Computer | 2 | Loss of text misrepresented individual tissue diagnosis on the final diagnosis print out from the HIS system software. The specimen heading "J" was removed and the diagnosis appeared under the specimen heading "I" as a result the specimen with the heading "I" had two diagnosis. | Software corrections were sent as of 10/30/2003 for clients to upgrade to, the correction software patch is in versions 3.17.5.16, 3.17.6.10, 3.17.7.3 and 4.1.1.0. The firm did not inform the non-affected consignees of the failure because the firm concluded that the associated risk was low.
Z-0365-2008 | Jan-23-2008 | 2008 | WaveScan WaveFront System - Advanced Medical Optics (AMO); VISX WaveScan WaveFront System (software). Part Number 0070-1478. Advanced Medical Optics; Santa Clara; CA 95051. | Visx Inc | 2 | Erroneous Treatment Calculations - Two software caused errors; in combination; in the WaveScan System Software; will result in an erroneous treatment calculation (overcorrection) in patients. (1) Installation of software for a certain brand of computer (Dedicated Computing LLC) can cause a software registry setting for an algorithm that identifies image reflection to be erroneously set in the "off | AMO sent an Action Required notification on May 30, 2007, to AMO Certified Physicians. The letter included a CD software update (0070-1721) to be installed by users that provided correct registry parameter settings. Upon completion of installation, a confirmation form will print out. Customers are requested to fax back confirmation to AMO (408-773-7055).
Z-2196-2008 | Jun-28-2008 | 2008 | Siemens SIMIView 3000; SIMVIEW NT and MEV ASIM S systems - Update Instructions "TH003/08/S" applies to Siemens brand SIMIVIEW 3000; SIMVIEW NT; and MEV ASIM S systems;Part Number: 5488460; Product is manufactured and distributed by Siemens Medical Solutions; Oncology Care System;4040 Nelson Ave; Concord CAThe product is indicated for use in Radiation therapy simulation. | Siemens Medical Solutions USA; Inc | 2 | The firm sent an Update Instruction; "TH003/08/S" for Motion Enable Switch/holder assembly. The switch holder allows the motion enable switch to be activated if inserted partially; thereby enabling inadvertent motion and serious injury may result. | On June 25, 2008 all consignees were notified via written letter and on-site visit by the firm's sales representative, informing them that the switch holder allows the motion enable switch to be activated if inserted partially; thereby enabling inadvertent motion and serious injury may result and providing instructions on the recall. A new software version will be provided. For assistance contact Siemens at 1-925-246-8407.
Z-0195-2009 | Nov-12-2008 | 2008 | GSI AUDERA - GSI Audera (version 2.6 software) Viasys Healthcare; Inc.Part or (Catalog) Numbers: 2001-9700; 2001-9705; 2001-9715; 2001-9720 and 2001-9725. The GSI Audera is used as part of an audiometric test battery. It can be used to assess cochlear function; enable estimation of behavioral hearing thresholds by evoking auditory brainstem or steady state brainstem response and to aid in detection of lesions in the auditory pathway on neonates through adults. | Cardinal Health; Inc | 2 | GSI Audera systems requires a system software update due to issues with the Split-Screen and the Vestibular Evoked Myogenic Potential (VEMP) functions. A. Split Screen Field Correction:GSI Audera units with version 2.6 software may mislabel Auditory Evoked Potential (AEP) waveform responses with respect to the identification of the stimulus ear. When using 2.6 software in split-screen mod | Consignees were sent a Cardinal Health "Urgent: Medical Device Correction" letter dated September 16, 2008 on 9/22/08 and 9/26/08. The letter was addressed to the Customer/Distributor. The letter instructed consignees to perform installations if they have the following software: 1) For Version 2.6 software, you must install Version 2.6.5 software to correct a Split- Screen performance issue and to remove the VEMP function. 2) For Version 2.0 or 2.1 software, you must install the Protocol Update Software to remove the VEMP function. Versions 2.0 and 2.1 do not contain Split-Screen capability. "Attachment A-Verification Form" must be completed by the recipient/distributor. If you need assistance in making the software revision determination or have further questions, please contact the Cardinal Health (formerly VIASYS) Neurocare/GSI Technical Support Group at 1-800-356-0007, 1-800-700-2282, 608-273-5000, or by e-mail at gsiservice@viasyshc.com or gsiservice@cardinalhealth.com.
Z-0213-2009 | Nov-05-2008 | 2008 | Mobile Care Server used in conjunction with Web Viewer; Pocket Viewer; and Cellular Viewer WIT - GE Healthcare; Mobile Care Server used in conjunction with Web Viewer; Pocket Viewer; and Cellular Viewer WIT with L-WEB05 software.GE Healthcare 8200 West Tower Avenue; Milwaukee; WI 53223Mobile Care Server maintains and coordinates the network connections between the devices in the Datex-Ohmeda network or Unity Network and Web Viewer clients.Web Viewer can be used for viewing or otherwise processing information from several bedside monitors or other networked devices.Pocket Viewer can be used for viewing or otherwise processing of information from several bedside monitors or other network devices.Cellular Viewer can be used for viewing or otherwise processing of information from several bedside monitors or other networked devices. | GE Healthcare | 2 | Issue 1 -Possible failure of CIC Pro audible alarms: Requests for telemetry defaults from Mobile Care Server could cause excessive CIC Pro memory consumption. This excessive memory consumption may lead to potential loss of audio at the CIC or a system reset or reboot. Issue 2 - Mobile Care software/Mobile Viewers don not recognize ApexPro FH Telemetry transmitter: Mobile Care server does not ac | Consignees contacted via a GE Healthcare Urgent Medical device Correction letter dated 8/26/08 and supplemental instruction for installation of software" Mobile care server with software, L-WEB05 and iCentral with software, L-NET05 supplemental Instruction for installation". The letter was addressed to Healthcare Administrator/Risk manger, chief of Nursing and director of Biomedical engineering / IT. The letter also described the problem, affected Products,and Short Term Recommendations .
Z-1442-2008 | Aug-15-2008 | 2008 | Varian Medical Systems On-Board Imager; Versions 1.0.15 and 1.2.05 - Varian Medical Systems On-Board Imager; Versions 1.0.15 and 1.2.05 Imaging Accessory to Medical Linear Accelerator. The On-Board Imager Device is used for verification of correct patient position in relation to isocenter and verification of the treatment fields in relation to anatomical and/or fiducial landmarks. | Varian Medical Systems Oncology Systems | 2 | If used with a third party radiation therapy treatment planning software system; mistreatment may occur because of a misalignment. | The firm has upgraded all user sites through normal course of service and installation activities. Notification letters were distributed starting January 03, 2008.
Z-1518-2008 | Sep-30-2008 | 2008 | GE Healthcare Innova - GE Healthcare Innova 2100 IQ Digital Fluoroscopic Imaging System (Versatile cardiac and Vascular imaging system). The Innova systems are indicated for use in generating fluoroscopic images of human anatomy for vascular angiography; diagnostic and interventional procedures; and optionally; rotational imaging procedures. They are also indicated for generating fluoroscopic images of human anatomy of cardiology; diagnostic; and interventional procedures. They are intended to replace fluoroscopic images obtained through image intensifier technology. These devices are not intended for mammography applications. | Datex - Ohmeda; Inc | 2 | User was unable to hear the Innova 2100-IQ 5-minute fluoroscopy warning signal from the Table Control TSSC interface. Inability of the user to hear the audible signal indicating the completion of any preset exposure time may lead to incorrect overall dose management; resulting in unnecessary exposure and can be a contributing factor for skin injury. The alarm can be heard in the control room. | GE Field Engineers are visiting all affected customers to apply updated software. For assistance contact Datax-Ohmeda, Inc., at 1-608-221-1551.
Z-0661-2008 | Apr-24-2008 | 2008 | ACUSON Sequoia Diagnostic Ultrasound System - ACUSON Sequoia Diagnostic Ultrasound System; general purpose Diagnostic Ultrasound system. Siemens Medical Solutions USA; Inc.; Mountain View; CA | Siemens Medical Solutions USA; Inc. | 2 | Measurement Error affecting all Sequoia systems using the PAL video standard configuration and have a perspective PC installed: 1) The Sequoia" system is operating at a software revision below 8.0. 2) The Sequoia" system is configured to PAL video mode (normal for systems running at 240V 50 Hz.). 3) While using any of the Perspective" advanced display options listed above; an image or clip is t | Consignees were notified via a Urgent Medical Device Notification letter/customer letter sent out on 9/25/2007 under Siemens update program US013/07/S. The Siemens Service organization coordinates the implementation of the field correction with the distributors for all countries. Consignees who do not respond to the recall communication will be visit ed and/or telephoned for follow up.
Z-1107-2004 | Feb-15-2008 | 2008 | ACUSON Aspen Diagnostic Ultrasound System - Siemans- ACUSON Aspen Diagnostic Ultrasound System; general purpose Diagnostic Ultrasound system; Model Number: 8247887; Siemans; Mountain View; CA 94039 | Siemens Medical Solutions USA; Inc. | 2 | Software error: Error affects all Aspen customer systems using the PAL video standard configuration and have a Perspective PC installed. This can lead to a measurement error under this specific condition: On secondary capture images/clips stored from within any Perspective advanced display option (specifically: 3D Surface Rendering; 3D MultiPlanar Reconstruction; Color Doppler 3D MultiPlanar Reco | Siemens issued an Urgent Medical Device Notification to consignees 9/18/07, under Siemens update program US014/07/S. (Distribution on this product had ceased on 6/29/2005.) The firm has instructed the user to not take measurements or use measurements on these secondary captures under the conditions described in the notice, to place a copy of the instructions for use in the file, and to notify all users of the issue with the software. The Siemens Service organization has coordinated the implementation of the field correction with the distributors for all countries. Non-responding consignees will be visited or called by telephone for follow up.
Z-2459-2008 | Sep-23-2008 | 2008 | NicoletOne 5.30.2 Software - NicoletOne 5.30.2 Software when used with CSeries amplifiers in Fixed Room LTM Configurations with patient room kits. The product is used for extended monitoring of EEG in a hospital setting. The system provides connection between the patient room EEG acquisition system and central control room. | Cardinal Health NeuroCare Division | 2 | NicoletOne 5.30.2 Software when used with CSeries amplifiers in Fixed Room LTM Configurations with patient room kits. Cardinal Health NeuroCare has confirmed that the 3rd party room relay connected to the A2 wall plate does not always activate if the patient event button connected to the CSeries Amplifier is pressed twice within one second. | Consignees were sent a "Cardinal Health Urgent: Medical Device Correction" letter dated June 23, 2008. The letter identified the product and problem. It also provided instructions (Field Correction) to install the software patch and requested return of the verification form. Contact Cardinal Health NeuroCare Division at 1-608-271-3333 for assistance.
Z-0124-2008 | Jan-30-2008 | 2008 | Philips PageWriter Touch - Philips PageWriter Touch Cardiograph;Model: 860284; Philips Medical Systems; Andover; MA; 01810 | Philips Medical Systems | 2 | Incorrect display: When using the 16 lead Patient Interface Module (PIM); lead tracings of V3R and V4R are reversed in the extended lead modes of Pediatric and Balanced. | Philips Medical Systems notified affected customers by Customer Letter and Urgent - Device Correction Notice dated 9/12/07. The correction will be a software upgrade. In the US, the Philips Response Center will contact US customers and arrange their upgrades, either via a software download, or by sending a Field Service Engineer to the customer site to perform the upgrade. Until new software is installed, customers are being advised to use the 12 lead module, and to disable or disconnect the 16 lead connection.
Z-0136-2009 | Nov-04-2008 | 2008 | Coroskop X-Ray System - Coroskop X-Ray System; Model numbers 6258219 and 6379759. | Siemens Medical Solutions USA; Inc | 2 | Unintended movement of the system table and/or C arm. | Siemens Medical Solutions issued a letter to consignees on 8/11/2008 to inform them of the problem and a workaround should the issue occur prior to software revision.
Z-2194-2008 | Sep-04-2008 | 2008 | Aria Radiation Oncology Version 8.1 - Varis Aria Radiation Oncology Version 8.1; manufactured by Varian Medical Systems; Palo Alto; CA The device is designed to assist the operator in providing accurate treatment setups for each patient by monitoring parameters and preventing the device from commencing radiation treatment plan inappropriately. | Varian Medical Systems Oncology Systems | 2 | Dose delivery may be altered after import from a third party planning system during radiation therapy. | Recall was initiated on May 21, 2008. A product notification letter (sent via certified mail) was distributed to all affected users, with description of the problem i.e., dose delivery may be altered after import from a third party planning system during radiation therapy . The letter was also distributed to Varian Sites, Marketing and Service Organizations. A software upgrade is currently being developed. For assistance, contact Varian Medical Systems, Oncology Systems at 1-650-424-5731.
Z-1643-2008 | Aug-24-2008 | 2008 | syngo Imaging XS - Syn go Imaging XS is a picture archiving and communication system intended to display; process; read; report; communicate; distribute; store and archive digital radiological images; including digital mammography images. | Siemens Medical Solutions USA; Inc | 2 | The firm discovered a problem where the selected patient images on the device may display an additional image from another patient or study. | The recalling firm issued a Customer Safety Advisory via certified mail to affected customers per Update Instructions IM011/08/S. The letter informs customers of the potential issue and provides instructions to detect its occurrence. A software update to correct the issue is already available. A service representative will visit affected sites to install the software update. If you have questions concerning the advisory letter or you did not receive a letter contact Alexander Vogel at +49 (9131) 84-3840 or alexander.vogel@siemens,com and make reference to complaint 08-1KM-0073 date 2008-03-06.
Z-2137-2008 | Sep-18-2008 | 2008 | ISOLOC Software - Isoloc Software; Version 6.6; 6.5 and 6.0. Part Number MT-NW-425-520. Software computer hard disc reads "ISOLOC 6.6 Copyright NMPE Copyright 1997-2004 NMPE Build #10302006 CIVCO MEDICAL SOLUTIONSIs an EPID-based (electronic portal imaging device) version of localization software for determining the treatment machine position of radiotherapy targets. | Civco Medical Instruments Inc | 2 | A localization may be produced which has the incorrect moves; and therefore; could place the patient in the incorrect location. | On 06/16/08 Civco Medical Solutions issued an Urgent-Medical Device Correction letter to all its customers by certified mail. This letter states the circumstances in which the device works incorrectly, a temporary work around and the information that the new software version 6.7 will address the issue. The software will be released in August of 2008. Attached to the letter is an ISOLOC Device Correction Notification Customer Feedback form which is to be completed and returned acknowledging the receipt of the Device Correction Notification.
Z-1827-2008 | Sep-03-2008 | 2008 | MM 4.0 and MIM 4.1 - MIM 4.0 and MIM 4.1 is a software package designed for use in diagnostic imaging. It is a stand alone software package which operates on Windows 2000/XP. | MIMvista Corp | 2 | In MIM (4.0.2-4.0.5) and MIM (4.1.0-4.1.2) versions; there is an error that may occur during the calculation of statistics using the SUV Tool or Contour Statistics. It will occur only under the following specific circumstance: A fusion of a primary volume which has a Z-Axis slice separation larger than that of the secondary volume and statistics are reported on the secondary volume. | MIMvista Corp has issued this safety recall notification for MIM versions 4.0.2 through 4.1.2. An error has been discovered in the way SW and contour statistics are measured and reported under some specific circumstances. Users of these versions should immediately contact MIMvista support at (866) 421-2536 or visit httv://www.mimvista.com/recall/ to obtain complete instructions regarding this error. A new version that corrects this problem is available on the Downloads page, http://www,mimvista.com/download.
Z-1752-2008 | Sep-02-2008 | 2008 | EndoWave - PT-3B Control Unit - component of the EndoWave Infusion Catheter System. EndoWave Infusion System consists of three main components; a single use EndoWave Infusion Catheter System which is comprised of a Drug Delivery Catheter (CD) and removable Ultrasound Core (USC) and a reusable EKOS PT-3B Control System. Box is labeled EndoWave Infusion System Drug Delivery Catheter with Ultrasound Core | EKOS Corporation | 2 | Potential for unit to burn patient. | In March 2008, the firm planned to visit each consignee to install the new software version 01.05.02 and hand deliver letters, dated 02/20/2008, notifying consignees of the reason for the action. The old software version is 01.04.02. The software was upgraded because the Ultrasound Core heats very quickly and if removed from the DDC while operating it may result in burns to the patient or user. The software will turn off the Ultrasound Core automatically if removed from the DDC during operation, and will prompt to restart therapy if ultrasound is turned off during treatment. On 04/17/2008, the firm issued a new letter with the date of April 2008 informing customers of the issue. In addition the letter indicates that the corrective action, release of the new software is delayed to incorporate further changes, and that customers will be notified when the software is available for their unit. Customers are advised that the device is safe to use when operating according to the Instructions for Use. EKOS contact number 1-888-356-7435.
Z-0592-2008 | Jan-24-2008 | 2008 | Hoana LifeBed - Hoana LifeBed Patient Vigilance System; (Intelligent Medical Vigilance System);Model Number: 010135-01; Product distributed by Hoana Medical; Inc.; 828 Fort St. Mall; Suite 620;Honolulu; HI | Hoana Medical | 2 | Bed Exit Failure: Failure to recognize "Bed Exit" when the feature is active due to a software anomaly. | On 11/7/07, all consignees were notified via an untitled letter and on-site visit by the firm's sales representative, informing them of the affected product and providing instructions on the recall. The recalling firm will arrange for installation of a software fix at each consignee.
Z-2211-2008 | Sep-16-2008 | 2008 | GE Signa¿ Excite 1.5T MR System; GE Signa¿ Excite 3.0T MR System - GE Signa¿ Excite 1.5T MR System; GE Signa¿ Excite 3.0T MR System.Model Numbers 5107849; 5107849-2; M3000PT; and M3000PW . Indicated for use as a diagnostic imaging device to produce axial sagittal; coronal and oblique images; spectroscopic images; and/or spectra; dynamic images of the internal structures and organs of the entire body; including; but not limited to; head; neck; TMJ; spine; breast; heart; abdomen; pelvis; joints; prostate; blood vessels; and musculoskeletal regions of the body. | GE Healthcare | 2 | A software issue that may result in misregistration of functional and anatomical images associated with the 3.0T BrainwaveRT fMRI Application on the 3.0T Signa HDx scanner. | A "GE Product Safety Notification" Customer letter addressed to Hospital Administrators, Managers of Radiology/Cardiology and Radiologists/Cardiologists dated September 14, 2007 was was sent to consignees. The letter described the safety issue, affected product, safety instructions, product correction and contact information. The letter instructed customers to scan only from the Superior to Inferior direction, when using BrainwaveRT with 2D T1 weighted scan. A GE Field Engineer site visit was also conducted to implement the correction (FMI No, 60745 & 60746), which enables users to scan in both Inferior to Superior and Superior to Inferior directions and have no mismap between the functional and anatomical images.
Z-1825-2008 | Sep-02-2008 | 2008 | Centricity AW Suite 2.0 Software - GE Centricity AW Suite 2.0 Software running on the Centricity RA 1000 or RA600/CA1000 Workstation; for diagnostic image analysis; GE Healthcare Integrated IT Solutions; Barrington; IL 60010 | GE Healthcare Integrated IT Solutions | 2 | Incorrect measurement of the aortic length in certain cases of aneurysm during the use of the aorta protocol or customized protocol of the Centricity AW Suite 2.0 running on the Centricity RA 1000 or RA600/CA1000 workstations. | GE Healthcare IITS sent Urgent Safety Notice letters dated 5/9/08 to all customers who have the AW Suite Software Version 2.0, informing them of the incorrect measurement of the aortic length in certain cases of aneurysm during the use of the aorta analysis protocol or customized protocols. The users were instructed not to use the aorta analysis protocol or the customized protocols for an aneurysm until the new software version can be installed. A software patch to permanently resolve the issue will be installed at no charge. Any questions were directed to the GE Customer Care Center at 800-437-1171.
Z-2183-2008 | Sep-16-2008 | 2008 | GE Centricity AW Suite Software version 2.0 and 2.0.1 with Card IQ Xpress Pro or Plus - GE Centricity AW Suite software version 2.0 and 2.0.1 with Card IQ Xpress Pro or Plus Module; for diagnostic image analysis; GE Healthcare Integrated IT Solutions; Barrington; IL 60010 | GE Healthcare Integrated IT Solutions | 2 | When using any vessel analysis protocol of the AW Suite Card IQ application; the software can reload saved tracking objects (Save State) incorrectly and display an incorrect vessel label over the restored images. This failure only occurs when users create a custom label in the vessel analysis protocol and include leading or trailing blank characters (space bar) in the custom label. | GE Healthcare IITS sent Product Safety Notification letters dated 6/11/08 to all customers who have the AW Suite Software Versions 2.0 and 2.0.1, informing them that if they also have Card IQ Xpress Pro or Plus, the AW Suite Card IQ software can reload saved tracking objects (Save State) incorrectly and display an incorrect vessel label over the restored images. This failure only occurs when users create a custom label in the vessel analysis protocol and include leading or trailing blank characters (space bar) in the custom label. The users were instructed not to make any clinical decisions based on a Card IQ Save State when a custom label has been ordered, and not to include leading or trailing blank characters (space bar) in the new branch name when customizing a Card IQ vessel analysis protocol until the new software version can be installed. A software patch to permanently resolve the issue will be installed at no charge. Any questions were directed to the GE Customer Care Center at 800-437-1171.
Z-2463-2008 | Sep-22-2008 | 2008 | SoftPath ASCII Software Releases 1.2; 2.1; 2.2; and 2.3. - Product is SoftPath ASCII Software Releases 1.2; 2.1; 2.2; and 2.3. | SCC Soft Computer | 2 | In the creation of revised report and supplemental reports diagnosis; text was inserted from another case. | Consignees were notified of recall and provided with a software correction patch ICC number 3618.
Z-2346-2008 | Sep-18-2008 | 2008 | Philips Healthcare Informatics iSite PACS - Philips Healthcare Informatics iSite PACS (Picture Archiving and Communications System); version 3.5.xThe system is a software package used with general purpose computing hardware to acquire; store; distribute; process and display images and associated data throughout a clinical environment. The software performs digital imaging; measurement; communication and storage. | Philips Healthcare Informatics; Inc. | 2 | Scout line and localizer crosshair on MPR images will display in the incorrect position under certain circumstances. | A "product notification" was issued to the firm's consignees on July 3, 2008. The firm asks consignees to disable MPR until a new software patch can be issued. For assistance contact Philips at 1-877-328-2808.
Z-1141-2008 | May-08-2008 | 2008 | iSite PACS - iSite PACS image management system; Picture Archiving and Communications System; Radiology software package; Philips Healthcare Informatics; Foster City; CA 94404. | Philips Healthcare Informatics; Inc. | 2 | Incorrect display of one of the patient's images. This malfunction could cause an image to be incorrectly displayed; for example flipped left to right and anterior to posterior. The original image is still presented in the correct orientation next to the incorrect image when displayed on the monitor. | Philips Healthcare Informatics mailed customers notification: Urgent Product Correction regarding FCO#2954704-01/04/08-001-C. Firm released a corrected version of software (iSite PACS v.3.5.69) on January 21, 2008. Firm is scheduling customers to upgrade to the corrected version.
Z-2195-2008 | Sep-16-2008 | 2008 | FastPlan; Verison 5.5 or 5.5.1 - Varian Medical Systems FastPlan; versions 5.5 and 5.5.1; model number HZ3; manufactured by Varian Medical Systems; Palo Alto; CA The product is a radiation treatment planning system. | Varian Medical Systems Oncology Systems | 2 | Due to issues with image orientation verification; a patient may receive all or some of prescribed high dose to healthy tissue instead of intended area. | Recall was initiated by the firm on 05/29/2008. A product notification letter /field safety letter was distributed to all affected users, with a description of the problem and user corrective action steps. In addition, the letter will also be distributed to Varian Sales, Marketing and Service Organizations. Letters were sent to consignees via certified mail with return receipt mailing to US consignees. For foreign consignees, letters will be translated as necessary and distributed with return proof of notification. Varian Medical can be contacted at 1-650-483-3153 for assistance.
Z-1826-2008 | Sep-02-2008 | 2008 | GE Centricity PACS RA1000 Workstation - GE Centricity PACS RA1000 Workstation; for diagnostic image analysis. GE Healthcare Integrated IT Solutions; Barrington; IL 60010 | GE Healthcare Integrated IT Solutions | 2 | A software anomaly results in incorrect study date and time information being displayed in the report screen and title; which may result in a potential patient misdiagnosis. | GE Healthcare IITS sent Product Safety Notification letters dated 5/5/08 to all customers who have the GE Centricity PACS Software Versions 3.0.0.X, 3.0.1.X, 3.0.2.X and 3.0.3.X informing them of the incorrect study date and time information being displayed in the report screen and title, which may lead to a potential patient misdiagnosis. The users were instructed to refer to the image overlay for the correct date and time until the new software version can be installed. A software patch to permanently resolve the issue will be installed at no charge. Any questions were directed to the GE Customer Care Center at 800-437-1171.
Z-0249-2008 | Jan-04-2008 | 2008 | Biograph 6 - Biograph 6; Biograph PET/CT Scanner; Catalog No. 8728854; Manufactured by Siemens Medical Solutions USA; Inc.; 810 Innovation Drive; Knoxville; TN 37932 USA; TEL: 865-218-2534 | Siemens Medical Solutions USA Inc | 2 | Visual asymmetric appearance; An asymmetry can be introduced into the attenuation-corrected PET images when using iterative reconstruction in combination with a specific number of subsets and reconstruction matrix sizes. (This problem only occurs when using protocols other than the default Biograph protocols) | A certified letter-Urgent Device Correction- dated June 15, 2007 was sent to all affected consignees on July 23, 2007 informing them of the problem (PET iterative reconstruction algorithm) and requesting as a short term fix that they discontinue the practice of operating using protocols other than the default Biograph. Software upgrades as the long term fix were promised in the Urgent Device Correction Notice. Updated information received October 4, 2007: a software patch to fix PETsyngo 5.0 was released, software patch to fix PETsyngo 4.x was released, and target date for the software patch to fix PETsyngo 3.2 and 3.4 is the end of November 2007.
Z-0086-2009 | Oct-08-2008 | 2008 | Acuson/Sonovista X300 - Siemens Ultrasound; Acuson/Sonovista X300 ultrasound systems; software revisions Sonovista X300 software revision 1.0.06 and 1.0.07; Acuson X300 sofware revision 1.0.08 ultrasound system with onscreen display. | Siemens Medical Solutions USA; Inc. | 2 | Incorrect value calculations by the device may result in inaccurate aortic stenosis estimates. | Siemens initiated a Field Correction 08/19/2008. Service personnel will conduct a site visit and distribute a Customer Safety Advisory Notice, correct the problem and install a revised software version.
Z-1754-2008 | Sep-29-2008 | 2008 | Mevatron KDS-2 - Siemens Medical Solutions; Mevatron KDS-2 Digital Linear Accelerator; part number 9411588 equipped with a Digital Electron Variable Applicator (DEVA; Part Number 8485971) with available energy of 6 MeV. Medical charged-particle radiation therapy system. | Siemens Medical Solutions USA; Inc | 2 | Radiation leak: When used with field sizes of 5cm x 5cm or smaller; product may leak radiation at a distance 2 cm from the side of the applicator body up to 13%; in excess of IEC standards. | Siemens sent a Customer Safety Advisory Notice dated 12/12/2007 to DEVA customers. The letter stated that Siemens would temporarily disable energies 5MeV and below for all customers that decide to continue using the DEVA, resulting in the temporary unavailability of energies 5MeV and below for all available. Customers wishing not to have these energies blocked must return their DEVA and they would be contacted by a local Siemens rep to schedule a service appointment. The letter stated that as a preventative measure that those customers should refrain from using the DEVA with any combination of energies 5MeV and below with field sizes smaller than 6cm x 6cm. The firm suggested using the fixed field electron applicators and cutouts with these energies and field sizes. Siemens will release a new control console software in 2008 for free-of-charge, that will limit the use of the DEVA to energies 6MeV and above, while allowing the use of lower energies using the fixed field electron applicatiors.
Z-2310-2008 | Sep-17-2008 | 2008 | AMO WaveScan WaveFront System Version 3.90/3.901 - AMO WaveScan WaveFront System Version 3.90/3.901; Part Number 0070-1478. The device is used in conjunction with the VISX STAR S4 and STAR S4 IR Excimer Laser Systems; which are indicated for wavefront-guided laser assisted in situ keratomileusis (LASIK) eye surgery in certain patient populations. The WaveScan System is a diagnostic instrument intended for the automated measurement; analysis and recording of refractive errors of the eye. | VISX INCORPORATED; A SUBSIDIARY OF AMO INC | 2 | Inaccurate measurements may be generated by the device; which could result in improper treatment and deterioration of patient eyesight. | A product advisory notice was mailed to all consignees on June 9, 2008, advising consignees to closely monitor the WaveScan results to ensure that mistreatment is not given to patients. The firm has ceased distributing the updated software. ***UPDATE 08/29/2008*** The firm revised its recall strategy on 08/28/2008 because the firm was informed that a 180 day PMA supplement rather than a real time review supplement would be required for the software upgrade approval. The firm will instead revert products to the previous software version, 3.67/3.671. If you have questions, contact 1-877-AMO-4LIFE, options 1, 2, and then 2.
Z-0079-2008 | Feb-21-2008 | 2008 | daVinci S Surgical System IS2000 - Intuitive Surgical; daVinci S Surgical System; Model Number: IS2000; A5.1 P1 and A5.0 P5 software level; Surgical Endoscopic Instrument Control System; Intuitive Surgical; Sunnyvale; CA 94086. | Intuitive Surgical; Inc. | 2 | System Lock-Up: Software anomalies could cause product failure during use; or on start-up. System transitions to a safe "soft-lock" state. | Intuitive Field Service Engineers installed a software upgrade beginning on 03/30/2007, and left a notification letter that the upgrade had been performed at each site. New software level is A5.1 P2.
Z-1245-2008 | Sep-11-2008 | 2008 | IMPAX¿ 6.2.1 - IMPAX¿ 6.2.1 System; Picture Archiving and Communications system (PAC); used in the acceptance; transfer; display; storage and digital processing of medical images. | AGFA Corp. | 2 | Text and Images may not be synchronized. | AGFA notified all consignees via registered letter on/about 06/28/2007 informing them of the potential problem and requesting the accounts ensure the visual synchronization indicators in the Text and Image Areas are used to verify the synchronization status of the Text and Image Areas. A FAX Back form was requested with each letter to ensure someone at the account acknowledges the information was received and understood. If a FAX Back form is not received, the account will be contacted by telephone. A Mandatory Service Bulletin was deployed to correct the problem in the software.
Z-1891-2008 | Sep-17-2008 | 2008 | GE Signa Excite 1.5T MR System; GE Signa Excite 3.0T MR System - GE Signa Excite 1.5T MR System; GE Signa Excite 3.0T MR System The product is used as a diagnostic imaging device to produce axial sagittal; coronal and oblique images; spectroscopic images; and/or spectra; dynamic images of the internal structures and organs of the body. | GE Healthcare | 2 | GE healthcare has become aware of an artifact which could affect the diagnostic capability associated with the 3.0T Propeller T2 Flair application of the 3.0T Signa HDx scanner. | Consignees were sent a "GE Healthcare Product Safety Notification" letter dated 9/14/07. The letter was addressed to Hospital Administrators, Managers of Radiology/Cardiology and Radiologists/Cardiologists. The letter described the safety issue, affected product, safety instructions, product correction and contact information. FMI 60745 and 60746 was communicated directly to the GE Healthcare service engineers who provide the modification at the customer site. For assistance, contact GE Healthcare at 1-262-544-3894.
Z-1516-2008 | Sep-29-2008 | 2008 | GE Innova fluoroscopic imaging system with bolus chasing - GE Healthcare Innova 3100 Digital Fluoroscopic Imaging System with Bolus Chasing option. | GE Healthcare | 2 | Incorrect dose data: GE Healthcare has found that after 6 days and 4 hours without performing a system reset or a system reboot (turning it off/on); the displayed dose data on these systems (Innova 3100 / Innova 4100 with Bolus option) may be underestimated by up to 50%. In order to avoid this issue; it is necessary to perform a system reset or reboot every day. GE Healthcare is working on a c | GE sent all customers with affected units a letter on 11/15/07 titled Urgent Medical Device Correction. In the US this was sent by a traceable means (either US Postal Service with receipt acknowledgement (registered mail) or equivalent courier service such as FEDEX overnight Delivery). The letter was sent to 3 titles within the affected accounts; Hospital Administrator, Manager of Radiology/Cardiology and Radiologist/ Cardiologist. A correction letter identified a short term recommendation to perform a system reset or reboot every day and stated that a software correction will be installed at their site at no cost.
Z-1540-2008 | Aug-23-2008 | 2008 | IntelliVue Clinical Information Portfolio (ICIP) - CareVue Chart Release C ;Versions C.00.00 through C.00.06Product No. 862246. This device is a clinical information system intended for use in data collection; storage; and management with independent bedside devices; and other systems that are connected either directly or through networks. It is used when there is a need for generating patient records and computation of drug dosage. | Philips Medical Systems | 2 | Medications prescribed for one patient were printed on the record of another patient. | Philips Medical issued letter notification, dated 29 Feb 2008 to customers titled; "Urgent Medical Device Correction Notice". Customers were asked to not solely rely on the Worklist to remind them when to administer orders or interventions until the software has been upgraded. Customers should review the reminders and verify orders and interventions on the Medication Administration Record and Administration record. The correction will be in the form of a software upgrade and will be performed remotely by Philips service personnel.
Z-0324-2009 | Dec-03-2008 | 2008 | Syngo MultiModality WorkPlace - syngo MultiModality WorkPlace (MM WP). Software product(Model Number 10140720). The product is intended for use in picture archiving and communications system. | Siemens Medical Solutions USA; Inc | 2 | Values derived from Dynamic CT data sets may be incorrect. The corresponding images are not affected. | The recalling firm has issued a software update to affected customers via Update Instructions CT034/08/S. This software update corrects the values from the volume perfusion post processing application. A firm representative will visit each consignee to install the software update beginning 9/12/08.
Z-1822-2008 | Sep-25-2008 | 2008 | GE Definium 8000 - GE Healthcare Definium 8000 Digital Radiographic System. Formerly GE Healthcare Revolution XR/d with Tomosynthesis. | GE Healthcare | 2 | Loss of image data: Software anomaly in the Processing Software on the Definium 8000 system may impact patient safety when using the VolumeRAD advanced application (option). The slice visualization of VolumeRAD exams acquired on the Wallstand receptor will be offset by 17mm. This prevents the visualization of anatomy WITHIN a distance of 17mm from the wallstand patient barrier (receptor cover). | GE sent a Product Safety Notification, dated March 7, 2008, to all customers. The letter described the safety issue, products affected, and stated that a GE Healthcare Field Engineer will schedule a visit to each facility to install a software modification.
Z-0462-2008 | Jan-31-2008 | 2008 | GEMINI GXL PET/CT system - Philips Medical Systems; GEMINI GXL 6; 10; & 16 slice PET/CT Systems; diagnostic imaging systems for fixed or mobile installations that combine Positron Emission Tomography (PET) and X-ray Computed Tomography (CT) systems; Philips Medical Systems; Cleveland; OH 44143 | Philips Medical Systems (Cleveland) Inc | 2 | Artifacts: Artifacts may appear in some images from head scans on the CT subsystem of the GEMINI PET/CT system. Reading images with these artifacts may lead to misdiagnosis or improper treatment. | Philips Medical notified domestic customers via Product Safety Notification dated May 30, 2007. The product safety notification includes a Technical Bulletin notifying customers of the issue and how to avoid it. This information will be used until the software correction becomes available. The software correction will be released in October 2008. Once software is available, the firm will begin correcting the field.
Z-2204-2008 | Sep-16-2008 | 2008 | da Vinci S Surgical System IS2000 - da Vinci S Surgical System; Model number IS2000 Endoscopic Instrument Control System; manufactured by Intuitive Surgical; Sunnyvale; CA. | Intuitive Surgical; Inc. | 2 | Defective software chip may cause the system to fail and lock up. | Recall initiated June 1, 2008. Urgent Device Recall Letters were sent via e-mail, with follow up phone calls and service visits to correct the product. The firm recommended that continued use of the da Vinci S Surgical System during the time period prior to our service update to replace the software chip. The letter emphasizes the importance of always having backup equipment and instrumentation available when performing surgery with the da Vinci S Surgical System and being prepared to convert to alternative surgical techniques, if necessary. Customers are asked to fill out and return the acknowledgement form.
Z-2056-2008 | Sep-17-2008 | 2008 | CyberKnife Robotic Radiosurgery System - CyberKnife Robotic Radiosurgery System radiation therapy device with Multiplan treatment Planning Software versions 1.5.2 and higher. | Accuray Inc | 2 | Sample beam data (which should not be used to treat patients) may differ from actual radiation output of an installed product; and which may be used by users. | On April 16, 2008, the firm issued notifications via fedex to affected sites. In addition, installation and upgrade procedures will be updated to incorporate a step which decommissions sample beam data. If you have questions, contact Accuray Customer Support at 1-877-668-8667 (USA) or 1-408-716-4700 (non-USA) or send and email to customersupport@accuray.com.
Z-1521-2008 | Sep-30-2008 | 2008 | GE Innova - GE Healthcare Innova 3100 Digital Fluoroscopic Imaging (Cardiovascular Imaging System) | Datex - Ohmeda; Inc | 2 | GE Healthcare has become aware of a potential condition that may impact operator or patient safety. One customer has reported that inability to terminate X-ray exposure after releasing the hand switch control (located in the control room). | An Urgent Medical Device Correction letter, dated June 2008 was sent to customers The letter described the issue, affected products and provided short term user recommendations. GE developed a new software release that corrects this condition and will launch the corrective action to upgrade all affected systems.
Z-0369-2008 | Sep-04-2008 | 2008 | GE Precision 500D Classical Radiographic-Fluoroscopic - GE Precision 500D Classical Radiographic-FluoroscopicX-ray System. Model #2289299-2; with console computer and touch screen fullly integrated. the GE Precision 500D Classical Radiographic-Fluoroscopic X-ray System is used to perform radiographic and fluoroscopic x-ray examinations. | GE Healthcare | 2 | It was reported that when the GE Precision 500D foot pedal is repeatedly activated; a defect in the fluoro-timer will cause the dose measurement to report a higher value than was actually received by the patient. Patient will receive a lower dosage irradiation time which will result in indication of cumulative air kerma higher than that to which the patient was actually exposed. | A customer letter "GE Healthcare Product Safety Notification" dated 10/4/07 was distributed to customers informing them of the safety issue. A Field Modification Instructions (FMI 10806) was issued to the field for the Precision 500D product to install a sofware upgrade to correct the issue. This correction was released in October 2007.
Z-0274-2008 | Aug-02-2008 | 2008 | GE Healthcare; Precison 500D Classical R&F System - Precision 500D Classical R&F System with control room PC; Image-intensified fluoroscopic x-ray system; ; Model Number: 5117866-27. | GE Healthcare | 2 | Inaccurate cassette size display reading/ Inaccurate mAs reading: 1. When the device is in PBL mode and a larger film cassette is inserted into the table Bucky (replacing a smaller cassette) after a protocol is selected on the console; the collimator will not adjust to the larger cassette size. The collimator stays at the original smaller cassette size and allows the operator to take an exposure. | Consignees were notified by two "Product Safety Notification" letters on July 10, 2007. Each letter addressed one of the two safety issues with the affected product. The letters provided safety instructions on how avoid the software error and informed users that their representative would contact them to schedule a visit for a software correction installation.
Z-0575-2008 | Apr-22-2008 | 2008 | BioPlex 2200 - BioRad BioPlex 2200 Software; for BioPlex 2200 Instrument; an automated immunodiagnostic analyzer; Bio-Rad Laboratories; Inc.; Hercules; CA 94547 | Bio-Rad Laboratories Inc | 2 | Incorrect values: An error was found in the BioPlex 2200 Software; Version 2.0 that results in assignment values being utilized for the BioPlex assay Calibrator Set that may be slightly different from those printed on the Value Assignment Data Sheet. Some values may be identical between the lot assignment combinations with the ANA Screen with MDSS kit. | Bio-Rad sent a Voluntary Field Correction notice dated November 1, 2007. Consignees were notified by phone call on Nov. 13, 2007, followed up by facsimile (domestic). Sub-recall to be by e-mail message for consignees outside the US. Distribution of the product ceased on Nov. 2, 2007. Bio-Rad will provide Calibrator Set CDs with only one lot of calibrator assignment values until a new version of software is available.
Z-1612-2008 | Sep-23-2008 | 2008 | Eclipse TPS - Eclipse Treatment Planning System; Client 8.1.18 and DCF 8.1.17; model number H48 | Varian Medical Systems Oncology Systems | 2 | A software error causes the wedge accessory calculation to be ignored in the radiotherapy treatment plan. | An Urgent Medical Device Correction letter was distributed October 2, 2007, to all direct consignees, informing them that the ARC treatment plans containing a Dose Dynamic MLC for IMRT delivery will not be recognized properly by the System. The firm provided corrective instructions and informed consignees that a product modification will be developed to resolve the issue. A Varian Service Representative will be scheduling a visit to upgrade the application.
Z-0571-2008 | Mar-18-2008 | 2008 | Ventana Image Analysis System (VIAS) - Ventana Image Analysis System (VIAS) - PATHWAY anti-HER-2/neu (4B5) primary antibody with software version 2.3.0.2; Model Number: VIAS-799-10000-00; Pathology Image Analysis System; Ventana Medical Systems; Inc.; Tucson; AZ 85755 | Ventana Medical Systems Inc | 2 | Incomplete upgrade: The software upgrade from February 1; 2007 may not have beencompleted. This software upgrade adds the algorithm (Slidetype) for thePATHWAY anti-HER-2/neu (4B5) antibody to the VIAS system. | Beginning in 6/07, Ventana Medical Systems, Inc., made phonecalls to labs asking to confirm software version, date of upgrade if available and confirmation of whether site uses 4B5 and Imaging together. A phone script was provided. Customers with systems not ugraded were to upgrade before further use of system. A notification letter dated 8/17/2007 was later sent to VIAS customers advising that the software upgrade from February 1 , 2007 may not have been completed on their system. This software upgrade adds the algorithm (Slidetype) for the PATWAY anti-HER-2/neu (4B5) antibody to their VIAS system . This Slidetype is cleared for marketing and is required prior to imaging slides stained with PATHWAY anti-HER-2/neu (4B5). The letter also advises that their VIAS and NEXES systems are conneded to the internet which enables this upgrade to be performed remotely with the support of the Ventana Medical Systems Customer Support Center (CSC). customers are instructed to call the CSC to carry out the upgrade or to schedule a Technical Applications Specialist (TAS) to assist with the upgrade in their lab. Arrangements for the upgrade are to be made prior to imaging slides stained with PATHWAY anti-HER-2/neu (4B5).
Z-2239-2008 | Sep-19-2008 | 2008 | ARCHITECT i1000SR System Assay CD ROM WW (excluding US) - ARCHITECT i1000SR System Assay CD ROM WW (excluding US) Special edition Version 1.01 and Version 2.0; list 01P61-01 and 01P61-02. The device is intended to duplicate manual analytical procedures by performing automatically various steps such as pipetting; preparing filtrates; heating; and measuring color intensity. | Abbott Laboratories | 2 | When running the Automated Dilution Protocol with the assay files "CMV IgG" or "Toxo IgG" on the ARCHITECT i1000SR system; a software error [9010 (Pipettor Robotics) preventing (Process Path) from operating] is generated and the instrument stops running. No results are generated. | Abbott affiliates were e-mailed copies of the recall letter Urgent Field Safety Notice Product Correction on 6/16/08 for translation where needed and dissemination to their ARCHITECT accounts who have the affected assay CD-ROMs. Each Abbott international organization will be responsible for contacting customers and determining the method of effectiveness checks based on their country requirements. Abbott sent Product Correction, Immediate Action Required letters dated June 16, 2008 to all active non-U.S. ARCHITECT customers who have the affected assay CD-ROMs. The end users were informed when running the Automated Dilution Protocol of the ARCHITECT CMV IgG and Toxo IgG assays on the ARCHITECT i1000SR System, software error code [9010 (Pipettor Robotics) preventing (Process Path) from operating] is generated and the instrument stops. The accounts were advised to not use the Automated Dilution Protocol of either the ARCHITECT CMV IgG and Toxo IgG assays on the ARCHITECT i1000SR System and report results for samples requiring dilution as ">250.0 AU/mL" for CMV IgG and ">200.0 IU/mL" for Toxo IgG until a corrected version of the ARCHITECT i1000SR System Assay CD-ROM is available in late September 2008. The accounts were requested to keep a copy of the letter on file. The accounts were requested to complete the customer reply form, indicating that the letter was received, understood and disseminated to the laboratory manager/supervisor responsible for Architect testing, and fax it to Abbott by 7/1/08 at 1-800-777-0051. If you have question, contact your local area Customer Support.
Z-1530-2008 | Aug-08-2008 | 2008 | BD FACSDiva - BD FACSDiva Software v6.0 and 6.1 with the HP XW4400 Computer SystemModel/Catalog number 642685Software is used to simplify acquisition of flow cytometry data. | BD Biosciences | 2 | Values might not update in statistics views after certain functions are performed in a worksheet. | The firm issued an Important Product Information notification letter on 12/21/2007. The letter instructs users to reload data by closing and reopening an experiment before reporting results in order to determine if their statistics have been impacted. They are asked to fill out and return a form pertaining to this problem. BD Biosciences will release a version of BD FACSDiva software to correct this problem. The letter also stated that BD Biosciences rep will contact the users when the new software is available. User are asked to contact BD Customer Support Center at 1-877-232-8995 if they have any additional questions.
Z-0111-2009 | Oct-10-2008 | 2008 | AcusonX300 - Siemens Ultrasound; Acuson/Sonovista X300 ultrasound systems; software revisions 2.0.1 to 2.0.05; 3.0.01 and 3.0.02; manufactured by Siemens Ultrasound | Siemens Medical Solutions USA; Inc. | 2 | Thermal Index cranial (TIC) is not displayed for the Neo-Head exam type with the C8-5 transducer. This is a required display to support the alara principle when imaging cranial structure. | Field correction was initiated on 8/19/2008. Under Siemens update program US020/08/S & US021/08/S Siemens service personnel make a site visit, deliver a customer safety advisory notice to the customer and correct the problem by installing the revised Software versions. After the planned field correction is performed at the customer site the Service personnel log that information through Siemens internal web or reporting tools into the Siemens field update database.
Z-0305-2008 | Jan-09-2008 | 2008 | Aplio XG - Toshiba Aplio XG Diagnostic Ultrasound System; Model: SSA-790A with version 1.1 software; Toshiba America Medical Systems; Inc.; Tustin; CA. 92780 | Toshiba American Med Sys Inc | 2 | Panel Lock-up: When using dynamic image acquisition modes the panel may lock up; if the operator attempts to perform a still image capture. If the panel locks up the system will have to be rebooted to continue the examination which may cause a delay in completion of the patient examination. | Notification letters, Urgent Medical Device Correction, with an attached Customer Response form were sent to customers beginning on October 2, 2007. Customers were notified this action was being taken as a result of an investigation that resulted from a software issue that was found by the manufacturer. Customers were informed that a Toshiba Customer Service Engineer would contact them to schedule a software upgrade when the software becomes available. If customers have any questions they were instructed to contact their or call (800) 421-1968 and ask for the Regulatory Affairs Department.--A short term instruction included "Avoid using the Freeze Key when performing dynamic image acquisitions".
Z-1155-2008 | Jun-12-2008 | 2008 | GE LOGIQ 3 Expert - GE LOGIQ 3 Expert ultrasound scanner with software versions 4.0.5; 4.1.1 and 4.1.2. GE Healthcare P.O. Box 414; Milwaukee; Wisconsin 53201 | GE Healthcare Clinical Systems | 2 | Inaccurate guidelines: The BE9C is an optional; small radius Bi-planar convex probe intended for use with the LOGIQ 3 Expert system in Urology application. An attempt by the user to activate the biopsy guidelines while this probe is in use will cause incorrect guidelines to be displayed on the image. The incorrect guidelines are recognized by their orientation that is left/right reversed from t | GE sent consignees a "GE Urgent Medical Device Correction" letter dated 8/14/07. The letter informed the customers about the problem, offer a short-term workaround and informed them that a representative will schedule a visit to upgrade the software.
Z-0554-2008 | Jan-29-2008 | 2008 | Varian RPM System version 1.7.3 - Varian Real-Time Position Management System (RPM); A Respiratory Gating System; Version 1.7.3; All Models; Varian Medical Systems; Palo Alto; CA 94304 | Varian Medical Systems Inc | 2 | Erroneously Resetting: RPM System version 1.7.3; software failure can occur that affects the gated radiotherapy treatment when phase-based gating is used. RPM 1.7 erroneously resets the phase-based gating thresholds to system default during the treatment session; if the user accesses the Session options during the treatment session. This can lead to difference between the planned and delivered | Varian sent an Urgent Medical Device Correction letter on September 26, 2007, to all its direct consignees, informing them that the RPM version 1.7 may reset the phase thresholds if Sessions options are accessed during treatment. A Service Representative will be scheduling a visit to user site to perform the modification requested to address the problem, when a corrected software version is available.
Z-0102-2009 | Oct-15-2008 | 2008 | Syngo Imagaing - Syngo Imaging with versions VB20B; VB20D; VB20F; VB20G; VB25B; VB30A; VB30A_SP1; and VB30A_SP2. | Siemens Medical Solutions USA; Inc | 2 | Image may not be visible after merge. | The recalling firm issued Customer Safety Advisory Notice dated 8/22/08 to inform the consignees of the potential issue and provide instruction to avoid its occurrence. To avoid potential risk of this issue, users are asked to check whether the extended HL7 messages as stated in the letter are mandatory for their RIS-PACS workflow, and disable these messages if they are not. Siemens is preparing a modification under the UI IM086/08/S for the affected versions of syngo Imaging that will resolve the potential malfunction. The field modification will be immediately distributed and installed by the firm's Service as soon as it is available.
Z-0096-2009 | Oct-16-2008 | 2008 | Coherence Therapist 2.0 - Coherence Therapist 2.0 equipped with a Coherence Therapist R2.0 or R2.1; and PRIMEVIEW 3iR2.0 or R2.1; part number 8139839Medical charged-particle radiation therapy system. | Siemens Medical Solutions USA; Inc | 2 | Flat panel positioning calibration could be off by as much as 4 mm without the machine discovering detail. | Service calibration inspection/correction Update instructions TH038/07/S was issued on 11/06/2007 to all customers. The instructions identified the affected products, stated the reason for the update, and listed the steps to update the software. Customer notification was also issued. Software patch planned for release in June 2008.
Z-0100-2009 | Oct-10-2008 | 2008 | Philips Telemonitoring System - M3811B Philips Telemonitoring Clinical Review Software part of Philips Telemonitoring System; M3810A; software revision identified as Build 1.1.2.11d; B.02.07. Prescriptive medical device used to automatically collect and transmit medical information (weight; blood pressure; non-diagnostic ECG) over phone lines between provider and patient. | Philips Medical Systems | 2 | Multiprint report may contain incorrect vital data for patients | Philips contacted customers by sending an Urgent Medical Device Correction Notification on 9/11/08. Letters to customers sent via certified mail and receipt of delivery will be tracked. Customers are asked to follow the Action To Be Taken By Customer/User section of the Urgent Medical Device Correction Notification. This section instructs users to allow the upgrade of the software upon logging in to the system. Contact Philips at 1-866-246-7316 for assistance.
Z-0202-2009 | Oct-27-2008 | 2008 | Oxylog 3000 - Oxylog 3000 Emergency and Transport Ventilator Continuous Respirator with software version 1.10 pre-installed; Catalog number 2M86300. | Draeger Medical; Inc. | 2 | May experience an interruption of ventilation for approximately 5 seconds | The firm issued an Urgent - Medical Device Recall letter dated September 2008. Software version 1.10 may under specific conditions experience an interruption of ventilation for approximately 5 seconds. During that time period, the patient airway pressure may drop to zero. This phenomenon was identified by Draeger engineers and to date has not been reported by any user. The firm's investigations revealed that when ventilating in BIPAP mode with the pressure rise time (Ramp) set to the highest setting and the expiration time set to exactly 0.5 seconds, the inspiration phase can coincide with the internal auto zero function of the device. Under these circumstances a pressure spike will occur upon reaching Pmax +5mbar (cmH2O) which with software version 1.10 will lead to the inspiratory phase stopping. Ventilation will resume after approximately 5 seconds. The device will generate a Paw Measurement INOP alarm. Until the new software is available, customers are being instructed to avoid the use of the following combination of settings: 1) BIPAP mode; 2) Ramp = highest, and 3) Texp = 0.5 seconds.
Z-1038-2008 | Apr-29-2008 | 2008 | Leonardo Workstation - Leonardo Workstation Picture Archiving and Communication System; Model Number: 7129534; Software Version: VB30B; Siemens Medical Solutions; Inc.; Malvern; PA 19355 | Siemens Medical Solutions USA; Inc | 2 | Incorrect display: The orientation labels will be incorrectly displayed on the reconstructed InSpace 3-D image if the orientation was not originally HFS on the acquisition system. | Siemens Medical Solutions issued a Customer Safety Advisory Notice dated 12/21/07 to its customers to inform them of the problem and instructions to avoid the problem until the firm's representative is available to upgrade the software.
Z-1241-2008 | Jun-27-2008 | 2008 | Agfa HeartLab Cardiovascular DICOMstore - Agfa HeartLab Cardiovascular DICOMstore Version 2.04.40.00 software; Heartlab; Inc.; Westerly; Rhode Island 02891. The software is used in the Agfa HeartLab Cardiovascular System. | AGFA Corp. | 2 | MIsidentification: The electrocardiogram (ECG) data of one patient is misidentified as the data of another patient. | On 6/21-26/2007, AGFA Healthcare notified, via telephone, the affected consignees of the situation. On 7/2/2007, affected consignees were sent, via registered mail, an "Urgent Medical Device Correction" letter instructing them to: a) remove the affected software DICOMstore 2.04.40.00 and replace it with the prior software version, DICOMstore 2.04.34.02; and b) review patient ECG reports to determine if any patients were actually affected and to notify the firm. AGFA Healthcare also provided a customer service contact number (877) 777-2432.
Z-2160-2008 | Aug-25-2008 | 2008 | General X-Ray System types with Digital Spot Imaging (DSI) - General X-Ray System types with Digital Spot Imaging (DSI) software | Philips Medical Systems North America Co. Phillips | 2 | X-ray images may be stored in the wrong patient file or corrupted; images from two different patients may be combined in one patient's folder or may be combined on one image (corrupted). | On 5/20/08 the firm issued an URGENT - DEVICE CORRECTION letter to the consignees advising of the problem, requesting consignees to check the stored or printed images to ensure that it represents the data and images captured from the examination. The firm will issue a Field Change Order to their service representatives providing new software, when software is available. This should be issued by the end of June.
Z-1515-2008 | Aug-16-2008 | 2008 | MOSAIQ versions 1.3 and 1.4 SEQUENCER - SEQUENCER; MOSAIQ versions 1.3 and 1.4 with Elekta and Siemens Machines; Model number 40003-001000-RO; Manufactured by IMPAC Medical Systems; Sunnyvale; CA.Used to facilitate the delivery of defined radiotherapy treatment plans. | Impac Medical Systems Inc | 2 | Software issue may result in change to intended treatment field; potentially resulting in mistreatment. | Firm issued field advisory to its consignees on 02/14/2008. Phone calls will also be made to affected customers to ensure reception and understanding of the field notification. Firm advises users with affected product to review the jaw positions of IMRT fields to ensure that they match the RT Plan after an Apply Limits or Machine Change occurs, until software at their site can be updated to the corrected version. A software patch is under construction.
Z-1615-2008 | Aug-24-2008 | 2008 | TomoTherapy Hi-Art System - TomoTherapy Hi-Art System with software versions 3.0 and 3.1 - Radiation Therapy System. The device is intended to be used as an integrated system for the planning and delivery of intensity modulated radiation therapy. | TomoTherapy Incorporated | 2 | In certain cases of HI-ART Systems operating with 3.0 or 3.1 SW; the registration adjustment field numbers may appear as nonsense characters or values on printed pages of the Register Tab. These values may appear differently than those shown on the Operator Station (OS) display. For example; the number "8" may appear as the number "0" on the printed page.During the registration process; the H | Consignees were sent a letter "TomoTherapy Incorporated Urgent Medical Device Correction Field Safety Notice" dated February 22,2008. The letter described the Issue, Product Affected, Cause, Recommended Customer Action and Resolution. If you have not received this letter or have questions, contact the firm as outlined below: North America: 1 866 368 4807 Belgium: 0800 38783 France: 0805 631 565 Germany: 0800 000 7541 Italy: 800 986 399 Netherlands: 0800 0201364 Spain: 800 300049 Switzerland: 0800 001927 United Kingdom: 0808 238 6035 China/CNC (Northern): 10 800 712 1701 China (Southern): 10 800 120 1701 Hong Kong: 800 967912 Japan: 0044 22 132374 Singapore: 800 1204 683 South Korea: 0079 81 4800 7204 All other locations: +1 608 824 2900 or +32 2 400 44 44.
Z-0118-2009 | Oct-28-2008 | 2008 | iSite PACS software versions 3.6.28.x and 4.1.x. - iSite PACS (Picture Archiving and Communication System); software versions 3.6.28.x and 4.1.x.The product is an image management system (software package) used with general purpose computing hardware to acquire; store; distribute; process and display images and associated data throughout a clinical environment. Device is intended for use by trained professionals. | Philips Healthcare Informatics; Inc. | 2 | Two defects have been identified: 1) potential to display a patient on the canvas page that is different from the patient whose images are displayed on diagnostic monitors when using Conference Presentation States and when opening two or more studies from a folder or an exam worklist and 2) potential to miscalculate measurements when Pixel Spacing and Imager Pixel Spacing DICOM tag values are bot | A notification letter (Urgent Product Correction) dated 8/29/08 was sent via certified mail to all customers alerting them of the issue, requiring posting of notification. Philips Healthcare Informatics, Inc. strongly recommend all customers advise their iSite Radiology users not to open multiple patient exams simultaneously from a folder or worklist using Conference Presentation States. A software correction is being worked on, and should be available by end of Q3 2008 for version 4.1.45.2. A mandatory upgrade is required for all customers who are currently using iSite version 4.1.45.2 and former versions. For specific questions regarding these issues, contact the Customer Advocate (CARE) or Customer Care at Philips Healthcare Informatics, Inc. at 1-877-328-2808. In your communication with Customer Care at Philips Healthcare Informatics, Inc. , please refer to this notification and associated FCO 22945704-08292008-004.
Z-2146-2008 | Sep-26-2008 | 2008 | GE Stenographe - GE Stereotaxy Positioner; model 2405544-2; for use with Senographe DS Full Field Mammography system; models 2383168; 2383168-2; 2383168-3; 2383168-3-1; 2383168-4-1. The expected use of the Senographe DS Stereotaxy is an optional accessory for the Senographe system for mammography examinations. | GE Healthcare | 2 | GE Healthcare has recently become aware of x-ray emission beyond the edge of the detector primary barrier. This issue occurs when an exam is performed in a specific angulated view associated with use of the Stereotactic Positioner of your Senographe DS Acquisition system and could impact patient safety. If the Stereotactic Positioner is removed when the examination arm is at 33 degrees and; th | A Product Safety Notification letter, dated March 3, 2008, was sent to affected customers. The letter provided information about the safety issue, identified affected product and stated that GE Healthcare will provide a software upgrade. Contact GE Healthcare at 1-262-544-3894 for assistance.
Z-0527-2008 | Jan-26-2008 | 2008 | RenalSoft - Baxter RenalSoft v.2.0 Patient Management Software Suite - HD Module; clinical data management software; Baxter Healthcare Corporation; McGaw Park; Illinois 60085 | Baxter Healthcare Renal Div | 2 | Pre-Treatment/Post-Treatment Report errors: 1) The Heparin Bolus Value is displayed on Pre and Post Treatment Reports for patients with "No Heparin" orders (RenalSoft v2.0; Renal Software Suite v.3.0 and v.3.1); and 2) Withheld Medications are incorrectly displayed as Administered on the Pre-Treatment report (RenalSoft v.2.0 only). | Baxter Healthcare Corporation sent an Urgent Device Correction letter, dated 11/14/07, to all of the affected customers, informing them of the Heparin Bolus and the Withheld Medication issues and provided them with interim mitigation instructions that they should implement while waiting for a software correction. A Baxter representative will contact the customer to schedule the software update when the upgraded software is released. Customers were requested to complete and return the enclosed Customer Reply Form, indicating that they have disseminated the information to their staff, and indicating if they still have RenalSoft software and require an upgrade. Any questions were directed to Baxter Global Technical Services at 1-800-553-6898, option 4. A follow-up letter dated 3/6/08 was sent to the consignees on the same date, informing them that the upgraded RenalSoft Software v.2.0.1 was now available for installation. The consignees were requested to complete and fax back the customer reply form indicating that they received the letter, and were informed that they would be contacted by a Baxter representative to schedule an installation of the software upgrade at no charge.
Z-1374-2008 | Sep-16-2008 | 2008 | OmniDiagnost Eleva - Philips OmniDiagnost Eleva X-Ray System is a diagnostic imaging device intended for radiographic; fluoroscopic; angiographic; and interventional applications. | Philips Medical Systems North America Co. Phillips | 2 | Unexpected movement of the X-ray table may occur. These movements may lead to patient injury; e.g.; an uncontrolled tilt may cause the patient to slip off. | On 3/22/08, the firm mailed letters to their customers via certified mail. The letter informed the customers of the recall, advising the customers to, "Use the emergency stop button or release the initiated movement command to halt undesired table movement. If you perform digital and conventional acquisition, you can avoid the problem by closing the single examinations daily or by clearing the patient database regularly." The letter further informed the customers that a Philips representative will be installing updated software at a future date. Contact Sarah Baxter at 425-487-7665 or Philips Call Center at 800-722-9377, #5, #2 if you have questions.
Z-0551-2008 | Jan-25-2008 | 2008 | ACCELERATOR Device Manager (ADM) - Abbott; ACCELERATOR Device Manager (ADM); List Number: 08H74-01; Abbott Laboratories; Abbott Park; IL | Abbott Laboratories; Inc | 2 | Sample/Patient Mis-identification: Two issues identified related to the ACCELERATOR Decision Manager (ADM); to include: 1) Sample Identification (SID) and/or Patient Identification (PID) numbers that contain more than 12 characters are truncated to the final 12 characters by the ADM; 2) for SIDs using alpha characters that are case sensitive; SIDs with lower case characters are transmitted by the | Firm notified consignees via Product Correction, Immediate Action Required letter(s) on 10/25/07. Consignees were notified of software issues and informed they would be contacted by an Abbott representative to schedule installation of software upgrade.
Z-0676-2008 | Feb-08-2008 | 2008 | Precedence SPECT/CT System - Philips Medical Systems -- Cardiac Viewer or Pulmonary Viewer software application version 2.0 in use with Precedence SPECT/CT System and all CT EBW workstations using software version 3.5. Philips Nuclear Medicine; Milpitas; CA 95035 | Phillips Nuclear Medicine | 2 | Incorrect Measurement: The firm discovered a software defect in Cardiac and Pulmonary Viewer software application options used in the Precedence CT Sub-System. The defect causes incorrect measurements in derived images which are zoomed and saved in a batch file. | A Customer Advisory Notification was mailed to affected consignees on September 28, 2007. The firm informed consignees that incorrect measurement can occur when the Cardiac Viewer or Pulmonary Viewer is used to zoom and save derived images in batch files. The firm provided instructions for operators to prevent incorrect measurements. Phillips is currently investigating methods to correct this issue and will provide updates, free of charge, when they become available. If you have questions concerning this recall, please contact Phillips Medical at 1-800-722-9377, (option5: Diagnostic Imaging, option 5:NM).
Z-0731-2008 | Feb-23-2008 | 2008 | Kodak DirectView DR 3000 System - Kodak DirectView Model DR 3000 System; Stationary X-ray System; Catalog #s 6551360; 6551386; 6552152; 6552160; 6551345; 6552319; 6552335. The units are labeled under the following firms: EASTMAN KODAK COMPANY (now known as Carestream Heath Inc.) Rochester; NY 14608; or CARESTREAM HEALTH; INC.; Rochester; NY 14608. | Carestream Health; Inc. | 2 | Unintended Positioner Movement - The U-Arm Positioner starts to move 1) without command from the operator and/or 2) can unexpectedly start movement if the system is powered down and re-energized after a collision has occurred with the Quantum Table. | On 12/21/07, Service Bulletins (dated December 2007) were sent to consignees informing them of design modification, and providing modification instructions. The firm followed up with telephone calls to the direct user sites to schedule appointments to visit each site, in order to make the necessary modifications. The firm will be installing Safety Adaptation Board and software into all the affected DR3000 systems.
Z-0520-2008 | Jan-31-2008 | 2008 | Vital Images ViTAL Connect - Vital Images ViTALConnect ; Version 4.0 and 4.1; PET/CT Visualization option. Vital Images; Inc.; 5850 Opus Parkway; Suite 300; Minnetonka; MN 55343. ( A web-based medical diagnostic tool that allows physicians to use PCs or notebook computers to gain remote access to 2D; 3D; and 4D advanced visualization. The software enables users to measure; rotate; and analyze images) | Vital Images; Inc. | 2 | Standard Uptake Values (SUV) display/update problems; in the way Standard Uptake Values (SUV) are displayed in 3D/MPR views. The problem arises when a user draws a selected region of interest in order to calculate the SUV statistics of the region. If the user then changes the slice shown by scrolling through additional slices; the original SUV measurements will remain on the screen and do not up | Customers were notified via telephone (11/7 -11/8/2007) and issued an Urgent Medical Device recall notification letter on 11/8/07. The letter described the problem and product involved. The letter also informed customers they they will provide a free software patch within the next 6-8 weeks.
Z-1751-2008 | Sep-29-2008 | 2008 | Siemens syngo Dynamics - Siemens Medical Solutions; syngo Dynamics 6.0 Workplace; Picture archiving and communication system | Siemens Medical Solutions; USA; Inc | 2 | Incorrect data: A software bug may result in a transfer of patient demographic data to a different patient's file. | Siemens issued a customer safety advisory to users on 2/8/08 advising them of the problem and instructing them to complete and verify that the report is locked before opening and reporting on a second patient, and that a software patch will be available Spring 2008.
Z-0164-2009 | Oct-22-2008 | 2008 | Varian RV Software - Varian RV SoftwareVarian Treatment for Non-Varian Linacs Version 6.6.5042 and 6.6.5043;Model Number: H46; Product is manufactured and distributed by Varian Medical Systems Inc.; Palo Alto; CADesigned to assist the operator of a radiation therapy device in providing accurate treatment setups for each patient by monitoring set up parameters and preventing the radiation therapy from commencing irradiation while any parameter is out of conformance with the treatment plan. | Varian Medical Systems Oncology Systems | 2 | A malfunction within the Varis software might result in a misadministration (underdose). The software does not correctly interpret the number; causing field parameters having decimal values to be incorrect. | Recall initiated on June 2, 2008. A product notification letter was distributed by certified mail to all affected consignees.
Z-1138-2008 | May-20-2008 | 2008 | syngo US Workplace - syngo US Workplace Picture Archiving and Communication System; software version 2.0; Siemens Medical Solutions; Inc.; Mountain View; CA 94043 | Siemens Medical Solutions USA; Inc. | 2 | Inaccurate results: A software bug may cause inaccurate wall motion abnormality scoring results to be displayed. | Siemens Medical Solutions sent an Urgent Medical Device Notification letter, dated November 2007, to its consignees. The letter described action required by users, and that Siemens will contact the user with a new software version correcting this error. A field correction to update software is planned to be underway, designated to start on 12/21/2007, and the correction is expected to be completed within 45 days.
Z-1460-2008 | Jul-30-2008 | 2008 | Acuity with Conebeam Computed Tomography - Acuity with Conebeam Computed Tomography; Radiation therapy simulation system; Model Number: H77. | Varian Medical Systems Oncology Systems | 2 | Incorrect Image Orientation: The image orientation tag may not be set correctly if the patient has been scanned with an orientation other than Head First Spine. When such images are imported to a treatment planning system; they may appear with either a mirrored or a rotated view. | Consignees were notified by an Urgent Medical Device Correction letter issued on 02/05/2008. The letter provided recommended actions for users to follow in order to avoid the software error. The letter informed users that a service pack is being created to correct the issue and that the recalling firm will notify users when it becomes available. For additional information, contact 888-827-4265.
Z-0116-2009 | Oct-20-2008 | 2008 | Artiste MV system - Artiste MV system; Part number 8139789; medical linear accelerator equipped with COHERENCE therapist RTT4.1. Medical Linear accelerator for radiation therapy. The device utilizes a console system for selecting various treatment sequences and manages commands for all motorized movement. | Siemens Medical Solutions USA; Inc | 2 | Three issues: unexpected rotation of gantry; unexpected movement of table between beams during patient setup; and single exposure images will be overwritten by double exposure images. | A letter marked "Urgent" dated July 4, 2008 was sent to customers. The letter provided a detailed description of the issues, the conditions under which these issues may occur, as well as recommendations to avoid these issues. Customer notification TH022/08/S was released on July 11, 2008. A software fix is currently being validated. All consignees notified by certified mail. Contact Siemens Medical Solutions USA, Inc. at 1-925-246-8407 for assistance.
Z-2235-2008 | Sep-17-2008 | 2008 | NeuViz - NeuViz Dual series computed Tomography Scanner System; X-Ray SystemThe CT scanner is a whole body x-ray computed tomography scanner which features a continuously rotating tube-detector system and functions according to the fan beam principle. | Philips And Neusoft Medical Systems Co.; Ltd. | 2 | Labeling of patient position on scanned image does not match actual patient orientation. | On 12/17/08 the distributor sent a letter dated 12/13/07 to their customers explaining the problem and under what circumstances it can occur, the actions that should be taken by the customer/user in order to prevent risks for patients or users, and that the customer/user will be contacted for scheduling of a software upgrade by the distributor. Contact NEUISYS Service Support Department at 1-877-299-9052 for assistance.
Z-0449-2009 | Dec-29-2008 | 2008 | Hologic - Hologic; Inc.; 10 Year Fracture Risk Questionnaire Option for QDR X-Ray Bone Densitometers Software | Hologic; Inc. | 2 | Software error may lead to a high estimate of major fracture probability. | Hologic notified customers by telephone on 9/23 and 24/08 and instructed them to not use the software. A customer notification letter was sent via FedEx to customers on September 25, 2008.The letter requested de-installation of the software and return of the software to Hologic in a pre-paid envelope.
Z-1224-2008 | Jun-25-2008 | 2008 | INTELECT VET - INTELECT VET; Model No. 2843; Chattanooga Group; Hixson; TN 37343; Made in the USA | Chattanooga Corp Inc | 2 | Devices have the potential to electric shock and burn | Beginning on 12/17/2007, Chattanooga Corp. sent two recall notification letters, via UPS mail, to its consignees informing them of the situation. The letters stated that the recall was to the user level and requested that the user perform the software upgrade, which will eliminate the possibility of shock and burn.
Z-1811-2008 | Sep-17-2008 | 2008 | Intuitive Surgical - INTUITIVE SURGICAL DA VINCI SURGICAL SYSTEM; MODE; S Cardiac Probe Gasper Instrument; Model IS2000; Part Number 420215-02; Product is manufactured and distributed by Intuitive Surgical; Inc. Sunnyvale; CA . The IS2000 CPG is used during cardiac ablation procedures. | Intuitive Surgical; Inc. | 2 | The product has a software interface problem. When the product is connected to the da Vinci IS2000 System it will not allow the System to recognize the Instrument which makes the IS2000 CPG instrument nonfunctional at all sites. Risks associated are loss of operability of the instrument; delay in surgery; and loss of dexterity. | On April 18, 2008 an Urgent Device Recall letter was issued to customers. Consignees were notified via e-mail and were instructed to return the product. Customers were instructed to segregate the product in a secure area for Intuitive Surgical CSRs. If you have any questions contact Karen Uyesugi at (408) 523-2100.
Z-0866-2008 | Aug-13-2008 | 2008 | Roche/Hitachi Modular D2400 Module GMMI clinical chemistry analyzer - Roche/Hitachi Modular Analytics System; Modular D2400 Module GMMI clinical chemistry analyzer; Catalog number 04998618001.Intended for qualitative and quantitative in vitro determinations using a wide variety of tests.Roche Diagnostics Corporation; 9115 Hague Road; Indianapolis; IN; 46250 | Roche Diagnostics Corp. | 2 | Possible mismatch between patient and result. | Consignees were notified via Urgent Medical Device Correction letter dated 3/21/08, which instructed them to not manually order QC test selections while the system is in operation mode and that they can only manually order QC when the system is in Stand-by mode. Consignees are instructed to consult with their physicians or pathologist to determine specific clinical implications for their patients. The firm states that this issue will be resolved in the next version of software.
Z-2006-2008 | Sep-17-2008 | 2008 | Boston Scientific iReview Software - iReview" Software (Accessory To) iLab Ultrasound Imaging System; Model Number: H74900018010; Catalog Number: 1801; Product is manufactured and distributed by Boston Scientific Corporation; Fremont; CA The iReview" Software is provided on CD and is a computer application designed for reviewing cases recorded with firm's iLab "System; and for exporting stills or videos of those runs. | Boston Scientific Corporation | 2 | The firm initiated a recall of a specific iReview" Software; Ver 1.0 Kit; which was released for distribution in a non-validated format. | Customer notifications sent on May 2, 2008 indicating the need for field correction and included a replacement version of the software kit. For customer assistance, contact Boston Scientific at 1-508-652-5594.
Z-0732-2008 | Feb-23-2008 | 2008 | Syngo Imagaing - Siemens Syngo Imaging System with VB30A version of software. Model Number: 10014063; Siemans Medical Solutions; Malvern; PA 19355 | Siemens Medical Solutions USA; Inc | 2 | Inconsistent data - A buffer overflow may occur between information system and picture system if the system experiences a high load. (e.g. RIS/PACS communication is unusally high) | The recalling firm issued a Customer Safety Advisory, via certified mail. to the affected customers on 11/27/07. The letter informs the customer of the issue and also provides instruction to detect and avoid its occurrence. In addition, the firm is developing a software update to correct the issue, which will be installed on the affected systems when it is released.
Z-1137-2008 | May-28-2008 | 2008 | Philips MutliDiagnost Eleva FD - MultiDiagnost Eleva Flat Detector X-Ray System; Part Number: SW 9896 001 43361; Philips Medical Systems; Bothell; WA; 98021 | Philips Medical Systems North America Co. Phillips | 2 | Delay in Imaging: MultiDiagnost Eleva with Flat Detector x-ray system cannot perform fluoroscopy procedures while certain image file housekeeping operations are performed. Fluoroscopy cannot be performed and new exposures are not able to be acquired until the housekeeping operations are complete which may take several minutes. The loss of system availability could occur during an interventional | On 11/29/07 Philips Medical Systems sent a letter dated 11/28/07 to customers stating that if users perform certain image file housekeeping operations, fluoroscopy cannot be performed and new exposures cannot be acquired until those operations are complete. It also instructs that the unavailability of fluoroscopy and exposure can be avoided by not pressing the "Delete All Images" button until any pending Autopush operation is complete. FCO 70800067 will be issued to service representatives to install new software.
Z-0313-2009 | Dec-09-2008 | 2008 | AXIOM ARTIS - MODULAR ANGIOGRAPHIC SYSTEM - AXIOM Artis System dBA using software version VB31D with 30x40 cm2 flat detectors.¿AXIOM Artis is a angiography system developed for single and biplane diagnostic imaging and interventional procedures. | Siemens Medical Solutions USA; Inc | 2 | System may switch to emergency fluoroscopy mode. | A Customer Safety Advisory Letter dated 8/20/08 was issued with Update Instructions AX045/08/S. The letter informed the customer of the potential issue and provided instruction to avoid its occurrence. A software update (VB31E) was released to correct the issue.
Z-0454-2009 | Dec-16-2008 | 2008 | Vidiera NsP Nucleic Sample Preparation - Vidiera NsP Nucleic Sample Preparation; Part Number: A22421; Version 1.0.41 | Beckman Coulter Inc | 2 | Incorrect results: under certain circumstances; when transfer volume is not within the set range; the Vidiera NsP software does not report sample exclusion and does not flag the excluded sample in the "Run Results" report. The software error associated with the excluded sample appears as "Liquid was not successfully transferred to well#..........the sample will be excluded from further proce | The customer notifications were initiated by Beckman Coulter on September 11, 2007 when all affected accounts were contacted via telephone, followed by a Product Corrective Action (PCA) was mailed via US Postal Service to all affected customers. The PCA informed the customers of the issue and instructed them to: (1) Review the "run Summary" event log for your samples. (2) Review the attached instructions describing the steps for identifying the samples affected by this error. (3) Complete and return the enclosed response form within 10 days so that we may maintain our records. The letter informed the customers that Beckman Coulter would provide a software update within 60 days. The letter also asked their customers to share the PCA with their staff and provide a copy of the PCA to any customers that they may have forwarded the affected products to. For questions regarding the PCA, please contact firm at 714-961-3634.
Z-1349-2008 | Mar-31-2008 | 2008 | Agfa IMPAX CardioVascular Suite Results Management - Agfa IMPAX CardioVascular Suite Results Management Nuclear Reporting Module; Outbound Report; Model/Catalog No.: L7W1P; Software Versions: 2.04.17; 2.04.19 and 2.04.20; AGFA HealthCare; Westerly; RI | AGFA Corp. | 2 | Reporting Differences: The Outbound Report software produces an Outbound Report (text version) which does not contain all the clinical content elements present in the PDF version of the report; also generated on the Nuclear Cardiology Reporting Module. | All five (5) sites utilizing this product were notified about the issue directly via telephone. An "Urgent Safety Notice" was sent via FED-EX to all the sites using the product on February 22, 2008. The letter describes the potential issue and mitigation. The recalling firm will also visit each site and make configuration changes to resolve this issue.
Z-1640-2008 | Aug-24-2008 | 2008 | Logix TPN Software Solution - Logix TPN Software Solution; catalog 2M8410; consists of the Logix Order Entry (OE) Installation CD; version 1.0.28 and Logix Compounder Manager (CM) Installation CD; version 3.1.8. The software is included in a number of Baxter's pharmacy compounders. The software assists the pharmacist in management of compounding; creating and authorizing nutritional solutions intended for IV administration. | Baxter Healthcare Corp. | 2 | Software anomaly results in inaccurate information being printed out on the Delivery Report; even though the compounder performed the compounding correctly. | Baxter sent the 3/28/08 Urgent Device Correction letter to affected customers, to the attention of the Director of Pharmacy, via first class mail to advise them of the potential inaccuracy of the Delivery Report due to a software anomaly. The accounts were instructed to review each delivery report for accuracy. If the report contains zero actual volumes the user must manually reprint the report. The letter outlined the process for reprinting the Delivery Report. Any questions were directed to the Baxter Nutritional Professional Services at 1-800-422-2751. The accounts were requested complete the enclosed reply form acknowledging receipt and understanding of the letter and dissemination of the information to their staff. A software upgrade will be provided. If you have questions about this letter or you did not receive this letter contact Baxter Nutrition Professional Services at 1-800-422-2751.
Z-0472-2009 | Dec-31-2008 | 2008 | GE Centricity Ultra Laboratory System - GE Centricity Ultra Laboratory System Software; for recording; reporting and distribution of lab results. GE Healthcare Integrated IT Solutions; Barrington; IL 60010. | GE Healthcare Integrated IT Solutions | 2 | Software computer; error: The abnormal result flag is not being sent in the HL7 messages on results that are lower than the normal range. For further information; please telephone GE Healthcare Integrated IT Solutions at 800-437-1171. | GE Healthcare notified their Centricity Ultra Laboratory Information System customers via letter dated 9/19/08. The letter was entitled "PRODUCT SAFETY NOTIFICATION." The letter states that the software correction was installed or will be installed to correct the abnormal results flag not being sent in the HL7 on results that are lower than the normal range. The letter summarized the safety concern and the action taken. Any questions were directed to GE Technical Support at 888-778-3373 or 905-479-1249. The accounts were requested to complete the enclosed customer reply form, supplying the software version number, and fax it back to 847-277-5240.
Z-1443-2008 | Aug-16-2008 | 2008 | iSite PACS - Philips Medical Systems iSite PACS (Picture Archiving and Communications Systems) version 3.6.28.0 - Radiology Image management system. | Philips Healthcare Informatics; Inc. | 2 | When using the Freehand Region of Interest (ROI) tool; errors may occur in area calculation yielding incorrect results. The firm discovered that a descrepancy exists in the measurement provided by the tool. | The firm issued product notification letter February 20, 2008, to its consignees. The notification alerts customers to the behavior, requires them to post the notification and notify users not to use the Freehand ROI feature until they are upgraded to a corrected software version. Philips will deliver a software upgrade to correct the problem.
Z-1459-2008 | Aug-02-2008 | 2008 | High-Definition 120 MultiLeaf Collimator - High-Definition 120 MultiLeaf Collimator; Medical charged-particle radiation therapy system; Software Version 7.0; Product Code: H56; Accessory to Medical Linear Accelerator. | Varian Medical Systems Oncology Systems | 2 | Dose Calculation Error: Software anomaly may result in failure of a MLC leaf to reach planned position; potentially resulting in misadministration of dose to a patient. | Consignees were notified by a Medical Device Notification letter sent on 01/29/2008. The letter advised users to cease use of the device until a corrected version of the software is available. A representative will contact users to upgrade the application once the revised software has been validated and released for distribution.
Z-0134-2008 | Feb-26-2008 | 2008 | Sentry Pro Medication Safety Software - Smiths Medical CADD-Sentry Pro Medication Safety Software - electronic pump protocol programming software - AdministratorREF 21-6275-51 RX only; Smiths Medical MD; Inc. St Paul; Minnesota 55112 USA | Smiths Medical MD; Inc. | 2 | Incorrect dosage possible: A software anomaly in the CADD-Sentry Pro Administrator software (version B) where the values of a weight based protocol go to zero or round up to the next higher hundredths place; even if Administrator displayed the values correctly at data entry. | A Smiths Medical MD employee visited the consignees between August 14 and 31, 2007, to verbally notify them of the recall, collect the recall CD-ROM containing the affected software and replace with the updated software.
Z-0576-2008 | Mar-04-2008 | 2008 | Roche COBAS 6000 - Roche COBAS 6000 Clinical Chemistry analyzer; Catalog Number: 04745914001; Roche Diagnostics Corporation; Indianapolis; IN 46250. | Roche Diagnostics Corp. | 2 | Incorrect results: CRP assays may be calculated using incorrect calibration parameters which could result in falsely high or falsely low patient results being reported. | Roche Diagnostics contacted the consignee by phone on 12/4/07 and by follow-up letter dated 12/4/07. The user was instructed to discontinue use of two CRP assays on this analyzer. The letter described a workaround and that a new revision of software, due out first quarter 2008, will fix this.
Z-0456-2009 | Dec-03-2008 | 2008 | Centricity Perinatal (formerly Quantitative Sentinel) System - Alert and Reminder - GE Healthcare Centricity Perinatal (formerly Quantitative Sentinel) System - Alert and Reminder software; automatic patient data management providing clinical information at the bedside in Labor & Delivery; Mother-Baby and the Neonatal Intensive Care Unit; GE Healthcare Integrated IT Solutions; Barrington; IL 60010Intended for use as a clinical data management system. | GE Healthcare Integrated IT Solutions | 2 | Software anomalies in the Alert and Reminder feature could result in a delay of treatment. When attempting to select the last visible alert or reminder choice; the next choice on the list below the desired choice is selected; and an inconsistent color may be displayed for the same clinical element across a set of work stations. | The software was updated to version 6.80.1 on 4/22/08, however, the consignee was notified of the software defects relating to the last visible Alerts and Reminders choice and the inconsistent display for the same clinical element across a set of work stations, via Product Safety Notification letter dated 8/29/08, sent via first class mail on 9/24/08. If there were any questions, the customer was to contact Tech Support at 1-800-433-2009.
Z-1901-2008 | Jul-23-2008 | 2008 | Levitronix CentriMag - Levitronix CentriMag Primary System;Article Number: 201-30100; extracorporeal blood circulation system. | Levitronix; Llc | 1 | Stops pumping: Interruption of CentriMag System support may occur when using a Valleylab Force FX-C Electrocautery Unit | Levitronix issued an Urgent - Voluntary Device Correction letter on 3/17/08 mailed to the US distributor and requested to contact their accounts. The letter revises the warning in the CentriMag Console Manual. Levitronix is issuing revised manuals. On March 13, 2008 the firm emailed and mailed (FedEx) to the one Canadian Hospital and Health Canada was informed of the Voluntary Advisory Notice via email on Wednesday, March 19, 2008. Rest of the World: Wednesday, March 12, 2008 through Monday, March 17, 2008. The Notice was issued as a Voluntary Advisory Notice after review and follow up with the UK MHRA. On 7/24/08 the firm issued a revised Recall Letter to advise users not to use Valley Lab Force FX-C or SSE2L Electrosurgery device with the firm's CentriMag Blood Pumping System and include the significant risk to health. The firm revised their letter and Black Box Warning and reissued the letters to the hospital, surgeons, etc. on 7/24/08. On 2/5/09 the firm issued a "Follow-Up to Urgent Device Correction of 3/17/08 " to advise users of the software upgrade to be installed in the CentrigMag Primary software Version 3.0 and CentriMag Back Up console , software version 2.0. Site visits will be made by Levitronix between Feb 2009 and April 2009 , software upgrade installed and new Operating Manuals provided.
Z-2341-2008 | Sep-16-2008 | 2008 | LIFEPAK CR Plus - LIFEPAK CR Plus Automated External Defibrillator. Device is primarily deployed in Public Access Defibrillator (PAD). It is a small lightweight device intended for use by minimally trained responders to treat victims of sudden cardiac arrest... Product is capable of 30 full discharges. Labeling on the device states LIFEPAK CR Plus LIFEPAK EXPRESS Defibrillators with ADAPTIV Biphasic Technology; Medtronic Emergency Response Systems; Inc. 11811 Willows Road NE Redmond; WA. | Physio Control; Inc. | 1 | This recall is being conducted due to the device being configured with the incorrect software for semi-automatic instead of fully automatic use. When the device is needed for a cardiac arrest emergency; the device will require that the user press the shock button instead of automatically delivering a shock as per the normal operation of the fully automatic device. Fully automatic defibrillators h | On 8/28/08 the firm began contacting consignees by telephone and stating that the LIFEPAK CR Plus defibrillators were discomfiture as semi-automatic, and that a responder is unable to press the shock button to deliver therapy when instructed by the AED voice prompt because the shock button is covered and not visible. The consignee was instructed that the unit(s) would be replaced immediately. This phone contact was followed by a letter flagged "'URGENT - MEDICAL DEVICE RECALL" dated August 2008 that was f axed or e-mailed the same day. The letter repeated the information provided in the telephone contact and stated to "Immediately perform one of the following actions: Remove the affected AEDs from service or Remove and discard the shock button cover (see enclosed illustration). Enclosed with the letter is an illustration providing instructions for removal of the auto shock button cover. Customer notification was completed by 9/2/08.
Z-0021-2010 | Oct-06-2009 | 2009 | Plum LifeCare 5000 (Plum 1.6) with Dataport - Plum LifeCare 5000 (Plum 1.6) with Dataport; list #s 02507-04-11; 02507-04-51; AC cord manufactured by Electri-Cord; Westfield PA; Infusion pumps manufactured by Hospira; Morgan Hill; CA. | Hospira Inc | 2 | Fire/Shock hazard-- A component used in the device may cause sparks/flashes and poses a fire and shock hazard. | Firm initiated its recall on 08/14/2009. A nationwide press release was issued to AP, and an Urgent: Device Field Correction letter to consignees was sent via UPS. The letter recommended that customers check their infusion pumps and spare parts in their inventory to determine if the specified AC Power Cords are in use at their facility; and, to inspect the affected cords for bent or cracked prongs and bridge connection, burnt plastic, or excessive wear and tear. Also, it was recommended that customers immediately discontinue use of the affected cords exhibiting the described characteristics and contact their local Hospira representative for assistance. The letter stated that customers may continue to use and monitor cords that are not exhibiting the described characteristics until replacement cords are sent to their facility. Hospira will contact customers to schedule replacement. Customers are to complete and return the attached Reply Form. Question should be directed to Hospira Customer Service at 1-800-241-4002.
Z-0142-2010 | Nov-09-2009 | 2009 | Model 3300 Cardiac Output Computer - Power cord for Model 3300 Cardiac Output Computer; List # 41245-04-03; AC cord manufactured by Electri-Cord; Westfield PA; Infusion pumps manufactured by Hospira; Morgan Hill; CA. | Hospira Inc | 2 | Fire/Shock hazard-- The power cord used in the device may cause sparks/flashes and poses a fire and shock hazard. | Hospira initiated its recall on 08/14/2009. A nationwide press release was issued to AP, and a letter to consignees was sent via UPS. The firm will replace affected cords and effectiveness checks will be conducted through tracking of servicing.
Z-1989-2009 | Sep-04-2009 | 2009 | Polar Care Kodiak - Power Supply Accessory (model AD-0660G)Power supply accessory is used with the Polar Care Kodiak unit; a motorized cold-therapy device. The Power Supply is plugged into a standard electrical wall socket and provides low voltage power to the Kodiak unit. The power supply is not used with other Polar Care models such as PC300 and PC500. | Breg Inc; An Orthofix Company | 2 | Power supplies short-circuited at the electrical outlet upon initial use of the power supply; causing the circuit breaker for the electrical outlet to trip. | An Urgent: Medical Device Recall Notification was sent to customers on June 18, 2009. FOR USERS: Customers are advised to immediately open all unused Kodiak cold therapy units and inspect the lot code of the Power Supply included with the Kodiak unit. If the lot code is 3508, 3908, 4308, or 4808, then remove the power supply, cut the cord using scissors and discard. Call Breg Customer Service for a replacement: 1-800-321-0607. Complete and return the response form ASAP. FOR DISTRIBUTORS: Send Breg your distribution information, including contact and quantities. Breg will notify your customers of the recall directly.
Z-2276-2009 | Sep-29-2009 | 2009 | SYNCHROMED II PROGRAMMABLE DRUG INFUSION SYSTEM - Medtronic SynchroMed¿ II Programmable Pump; 8637; Rx Only; Sterilized by ethylene oxide gas; Medtronic inc.; Minneapolis; MN 55432-5604. (The SynchroMed¿ II Pump is supplied in 20 ml or 40 ml reservoir size.) Affected products are SynchroMed II Model 8637 pumps with batteries (MECC Lifetest Batteries) manufactured from April 9; 2002 through July 2; 2002 (internally defined as Time Period 1); and from August 15; 2003 to March 16; 2005 (internally defined as Time Period 3).The SvnchroMed II Programmable Pump is indicated when patient therapy requires the chronic infusion of the certain drugs and fluids. | Medtronic Neuromodulation | 2 | This recall is related to the potential for reduced battery performance in a small percentage of Medtronic Model 8637 SynchroMed? II pumps with batteries manufactured during two distinct time periods prior to April 2005.Nature of the Device Issue: As part of ongoing analysis of returned explanted product; Medtronic has confirmed that reduced battery performance resulted in eight (8) occurrence reduced battery performance (low battery reset: a critical alarm) occurrences and one occurrence of premature elective replacement indicator (ERI: a non-critical alarm). The devices affected were manufactured from April 2002 through July 2002 and from August 2003 through March 2005. Of the nine confirmed events, all patients have experienced return of underlying symptoms and/or withdrawal symptoms and all have undergone surgical revision to replace or remove their pump. No patient deaths have been associated with these events. Complications associated with drug withdrawal are possible. Patients receiving intrathecal baclofen therapy (i.e., Lioresal® Intrathecal) are at higher risk for adverse events, as baclofen withdrawal can become life threatening if not promptly and effectively treated. See Attachment # 5 for additional information. | Consignees were sent on July 7, 2009 a Medtronic " Urgent: Medical Device Correction" letter dated July 2009. The letter was addressed to Healthcare Professionals. The letter described the nature of the Device Issue, Scope, Potential Severity of the Issue, Recommendations, Ongoing patient Management Recommendations and Additional Information. Important patient management information is also available at http://www.professional.medtronic.com under the heading Product Advisories. Direct questions to your Medtronic Field Representative or Medtronic Neuromodulation Technical Services by calling at 1-800-707-0933.
Z-0567-2010 | Dec-23-2009 | 2009 | Eclipse Oxygen System - SeQual Technologies Inc. Eclipse Oxygen System; Model 1000 Indicated for the administration of supplemental oxygen. | Sequal Technologies Inc | 2 | Unit may shut down when operating from the Power Cartridge. A temporary power interruption would cause the unit to sound an alarm and shut down. If the patient does not have access to an alternate supply of oxygen or access to AC or DC power; this may lead to shortness of breath; hypoxia; or in some patients may cause serious health consequences. | The recall notification was initiated on November 12, 2009 with the firm forwarding a "Medical Device Correction" letter to all of the affected consignees via US Postal Service for US Consignees and FedEx for Foreign Consignees. The Medical Device Correction Letter contained the following: (1) Medical Device Correction Letter; (2) List of Serial Number affected that were distributed to the Provider; (3) Reply Card that the Provider will be requested to send back regarding receipt notification of the Medical Device Correction Letter and confirming that they have the affected product; (4) Additional Medical Device Correction Letters for Providers to distribute to their patients. The Medical Device Correction Letter informed the consignees of the reason for recall, the hazard involved and the actions that were required. For further details or information, consignees were instructed to contact SeQual Technical Support at 1-800-826-4610 from 8:00 AM to 5:00 PM PST.
Z-2252-2009 | Sep-21-2009 | 2009 | FLEXIFLO QUANTUM ENTERAL NUTRITION PUMP - Flexiflo Quantum Enteral Pump Model # 50596; 50598 and 59322.The Flexiflo Quantum enteral pump is a medical device that is used to deliver enteral nutrition products to tube-fed patients. The pumps are generally used in hospitals; nursing homes; health care institutions and home care. | Abbott Nutrition; Abbott Laboratories | 2 | The firm initiated this voluntary recall because some of its Flexiflo Quantum Enteral Pumps; have a 110-volt AC power cord plug that may crack or fail. This potential power cord plug failure may cause electrical shock and/or a delay in feeding that could represent a potential health and safety risk. | An "Urgent Device Recall" letter dated September 4, 2009 was issued via FedEx overnight delivery to all direct consignees. The letter described the affected product, issue and immediate actions for customers. Direct consignees were further instructed to immediately contact affected customers and notify them about the Urgent Device Recall. Abbott Nutrition (AN) Device Call Center (DCC) will telephone those customers who do not responded promptly to the "Urgent Device Recall" letter with the completed AN Account Summary of the Flexiflo Quantum Enteral Pump Shipments. AN Device Call Center will make arrangements for the service or replacement of the recalled pumps. Please contact the AN Device Call Center at 1-877-457-0249, Monday through Friday between the hours of 8:30 am and 5:00 pm Eastern Standard Time (EST). An AN representative will work with you to arrange service or replacement of the recalled enteral pumps.
Z-2328-2009 | Sep-30-2009 | 2009 | Martel Printer Accessory to i-STAT Portable Clinical Analyzer - Martel Printer Accessory to i-STAT Portable Clinical Analyzer; rechargeable battery pack.Printing results of clinical chemistry tests and test panels contained in the i-STAT test cartridges. | Abbott Point of Care Inc. | 2 | It appears that the rechargeable battery pack is overheating with or without the presence of smoke. | Urgent Recall Notice letters dated April 2009 were sent out the third week of April by Federal Express. The letter discussed the background, required actions, and additional information. Customers are to identify all affected Martel printer rechargeable battery packs, record the number of affected packs on the attached business reply cards, and return the card along with the affected batteries as instructed. Questions should be directed to Abbott Technical Support at 800-366-8020.
Z-1902-2009 | Aug-25-2009 | 2009 | CLINITRON¿ CII - CLINITRON¿ CII Air Fluidized Therapy Units.Intended to prevent and/or treat pressure ulcer development and wound deterioration in patients who have a significant risk of developing these problems and who generally also have one or more of the following conditions: immobility; poor nutrition; diminished level of consciousness; reduced subcutaneous tissue or multi-system failure. | Hill-Rom Manufacturing; Inc. | 2 | Incorrectly repaired power cord wires can potentially overheat resulting in smoldering sound foam in the base of the therapy bed. | Hill-Rom issued an "Urgent Field Safety Notice" letter dated February 2, 2009 addressed to Facility Risk Managers and Facility Administrators. The notice detailed the issue and instructed users on actions to be taken. For further information, please contact Hill-Rom Customer Service at 1-800-638-2546.
Z-1813-2009 | Aug-11-2009 | 2009 | Continuous Ventilator/Continuous Ventilator with Respiratory Mechanics - Respironics Esprit Ventilator System; Model Number V1000 with Optional External Battery; Product Number: 1001456. | Respironics California Inc | 2 | The recall was initiated because through routine product monitoring Respironics discovered that under certain conditions; when operating on AC line power; a portion of the current from the External Battery charger is erroneously diverted to the ventilator. When this occurs; the External Battery charger may not transition correctly. Prolonged operation in this condition results in overcharging and | Respironics California, Inc. issued an "Urgent - Medical Device Field Correction" notice dated April 22, 2009. Customers were informed of the affected device including corrective actions taken by the firm. A Respironics representative will contact customers to schedule service of the device. Further questions can be addressed by contacting Respironics US Customer Service at 1-800-345-6443.
Z-2240-2009 | Sep-24-2009 | 2009 | Innova¿ 2000 - GE; Innova¿ 2000 The principle system components include a C-arm; image acquisition; processing and archiving capabilities.Innova¿ 2000 The principle system components include a C-arm; image acquisition; processing and archiving capabilities. | GE Medical Systems; LLC | 2 | GE Healthcare has recently became aware of a potential power supply failure with the Innova? 2000; 3100 and 4100 systems that may impact patient safety. A potential power supply failure inside the GE equipment cabinet could occur with no advance warning and thereby prevent x-ray generation and image processing. Should this occur; the system cannot be restarted. | Consignees were sent a GE Healthcare "Urgent Medical Device Correction" letter dated June 25, 2009. The letter was addressed to Hospital Administrators/Risk Managers, Managers of Radiology/Cardiology or Radiologists/Cardiologists. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information. GE Healthcare will correct all affected systems by replacing the power supply located in the GE equipment cabinet. A GE Healthcare service representative will contact you to arrange for this correction. Direct questions about the recall to the firm's Call Center by calling: 1-800-437 -1171, option 4 (United States) or 0120 - 055 - 919 (Japan). For other countries, please contact your local GE Healthcare field service Representative if you have any questions or concerns regarding the notification.
Z-1206-2009 | Apr-01-2009 | 2009 | Zoll AED Plus - Zoll AED Plus Defibrillator | ZOLL Medical Corporation; World Wide Headquarters | 1 | Device fails to discharge the defibrillation energy. | On 3/31/09, the firm revised their recall strategy and required all users to a perform a "Mandatory Upgrade" software The software upgrade allows all potentially affected devices to monitor battery charging performance through periodic self-testing. If defective batteries are detected at any time prior to the recommended maximum of five years, users are prompted by the device to install fresh batteries. Frim issued Press Release on April 2, 2009. In the first notification Zoll notified customers by letter on 2/12/09 via certified mail. Customers instructed to replace device batteries every three years. In lieu of replacing batteries every three years, customers will also have the choice to update their devices with software that will monitor batteries for the identified defect. ZOLL will be send email notification to all consignees who have provided an email address as part of their contact information. ZOLL will also be publishing information of the notification on industry magazines. http://www.zollaedplusbatteryhelp.com/.
Z-1038-2009 | Mar-02-2009 | 2009 | Colleague; Colleague CX and CXE Triple Channel Infusion Pumps - Baxter Colleague Triple Channel Volumetric Infusion Pumps. Baxter Healthcare Corporation; Medication Delivery Division; Deerfield; IL 60015 U.S.A.; Made in Singapore. Product codes: 2M8153; 2M8163; 2M9163; DNM8153; DNM9163; BRM8153; BRM8153T; BRM8153RT; PNM8153; PNM8153T; PNM8153RT; CNM8153; GNM8153; HNM8153; TRM8153 and WNM8153.Electronic infusion pumps are indicated for continuous or intermittent delivery of solution through clinically acceptable routes of administration; such as intravenous (IV); intra-arterial (IA); subcutaneous; epidural or irrigation of fluid spaces applications. | Baxter Healthcare Corp. | 1 | Baxter has identified failure codes that lead to a delay or an interruption of therapy; hazards associated with improper cleaning; and damaged battery messages for the Colleague Infusion Pumps. | Baxter Healthcare Corporation sent an "Urgent Device Correction" letter dated January 23, 2009 to all Colleague Infusion Pump customers, to the attention of the Vice President of Nursing with copies to the Chief Operating Officer, Director of Biomedical Engineering and Director of Risk Management to advise them of failure codes that lead to an interruption of therapy, hazards associated with improper cleaning, and damaged battery messages associated with the Colleague pumps. The letters listed the failure codes involved that lead to an interruption of therapy which could lead to serious injury and/or death; advised institutions to have contingency plans to verify that back-up pumps are available; and provided steps for addressing an interruption of therapy with any failure code. The letters also provided information about the hazards associated with improper cleaning of the pumps and instructions for addressing damaged battery messages. Customers were instructed to complete the attached Customer Reply Form acknowledging receipt of the letter and fax it to 888-569-1585. Direct questions to your local Baxter Representative or to Baxter Global Technical Services by calling 1-800-843-7867, Prompt 2, between 7AM and 5 PM Central Standard Time (CST).
Z-1256-2009 | May-22-2009 | 2009 | HemoSense INRatio2 - HemoSense INRatio2 Prothrombin Time Monitor; manufactured by HemoSense Incorporated.The INRatio2 PT Monitoring System is used for the quantitative measurement of Prothrombin Time (PT) in fresh; capillary whole blood. | HemoSense Inc | 3 | INRatio Monitors that were manufactured with a specific EEPROM memory chip fail during use. The monitors report an error code and do not report a test result. | Letters for each product titled "Urgent: Medical Device Recall Notice" dated December 19, 2008 were issued explaining the problem and instructing users to discontinue use, complete the provided reply form and return the device by mail to Inverness Medical. Distributors of the violative device were instructed to contact their customers and provide the same recall information. Further questions should be addressed to HemoSense Technical Service at 1-877-441-7440.
Z-0563-2010 | Dec-23-2009 | 2009 | AquaStar - AquaStar Water Caloric Irrigator; water caloric stimulator; Manufactured for/Distributed by Micromedical Technologies; Inc.; 10 Kemp Drive; Chatham; IL 62629 USA; Instrumentation Difra; Made in Belgium.The AquaStar water caloric irrigator delivers a precisely controlled temperature and flow rate of cool or warm water for vestibular stimulation during caloric testing. The caloric tests evaluate the viability of the horizontal semicircular canal by stimulating them with warm and cool water while the patient's eyes are closed. The resulting dizziness and nystagmus is taken as an index of the viability of the organ. The eyes are then opened to evaluate the ability of the CNS to visually suppress inappropriate dizziness and nystagmus. | Micromedical Technologies; Inc | 3 | Some of the filters on both the display board and the main electronic board (PCB) become oxidized causing the LCD display to be garbled and the timer will not operate. | Micromedical Technologies, Inc. sent letters dated November 16, 2009 via e-mail (with read receipt) to each customer and their representative. If they did not have the customer's e-mail address, the sales representative was requested to forward the letter to the customer and copy Micromedical. The letters advised the customers that one or more electronic components in the AquaStar caloric irrigator may fail due to contact corrosion. In the event of this failure, unheated tap water will continue to flow through the device back to the water drain, but the LCD display will be garbled and the timer will not operate. If this problem develops, the account should stop using the irrigator and contact Micromedical Technologies, Inc. at 1-217-483-2122 or via e-mail at mmtinfo@micromedical.com to report the problem for repair of the unit.
Z-1260-2009 | Jun-17-2009 | 2009 | REMStar M-series Heated Humidifier System - Remstar Pro M-Series Heated Humidifier System. | Respironics | 2 | Malfunction of the J3 connection between the Heater Plate Main printed circuit assembly (PCA) and the heater plate may have an intermittent connection which may lead to excessive heat build-up on the humidifier PCA. | A Service Communication dated February 5, 2009 was provided to all Respironics customers that are approved to service Respironics devices. The communication described the affected device and actions required. Further questions may be addressed by phone at 1-800-345-6443, and by email at service@respironics.com or clinical@respironics.com.
Z-1870-2009 | Sep-09-2009 | 2009 | Uni-Circuit (W Connector; Adaptor; Y piece) - GE Healthcare; Uni-Circuit Kit; Adult; Disposable; 1.8 m/72 in; REF M1089028; M1087972; M1087975; M1126445; M1127508; M1127509; M1129780;M1129781; M1130230; M1137805; M1137199; M1145542; M1170604; M1151662;M1152224; M1158990; M1100486; M1122080; M1126444; M1163003; M1163255M1132533; M1133966; M1170051; M1135144; M1139485; M1118627; M1118628;M1145406; M1147298; M1137811; M1152651; M1160275. pkg of 20 pcs; Latex Free. Breathing circuit for use with anesthesia machine or ventilator. | GE Medical Systems; LLC | 2 | Breathing circuit inner tube may separate from the rigid connectors under certain conditions. A separation of the inner tube from the connector may prevent adequate gas flow (O2; N2O and anesthetic agents) to the patient. A separation can also cause CO2 re-breathing. This could lead to hypoxia and/or hypercarbia and inadequate anesthesia delivery. | GE Healthcare sent consignees 6/5/09 an "Urgent Medical Device Correction" letter dated June 5, 2009. The letter was addressed to Health Care Administrator/Risk Manager; Chief of Anesthesia; Director of Clinical Engineering; Risk Manager. The letter described the Safety Issue, Affected Product Details, Safety Instructions (included "Recall Confirmation" form fax to 973-790-4150), Product Correction and Contact Information.
Z-1182-2009 | Apr-30-2009 | 2009 | Stryker PainPump2 - Stryker PainPump2; 250 mL PainPump with Luer Lock Tubing Set; 2.5 inch ExFen and Standard Catheter; 5 per box; sterile; Stryker; Kalamazoo; MI; REF 525-125.The Stryker PainPump2 is an electromechanical pump designed to deliver a controlled amount of medication to the patient for pain management. | Stryker Instruments Div. of Stryker Corporation | 2 | An electronic failure during preparation of the unit; or during patient use; may cause the pump to shut down. | A "Medical Device Recall Notification" letter dated February 11, 2009 was issued to consignees via phone, mail and/or email. The letter described the issue and instructed customers to discontinue use of the recalled product, destroy it according to instructions provided with the notification and return the "Business Reply Form" by fax to Stryker Instruments Regulatory Department, 1-866-521-2762. Further questions should be addressed to Stryker Instruments at 1-800-800-4236 extension 4086.
Z-1389-2009 | Jun-23-2009 | 2009 | AED 10 - Welch Allyn AED 10 automated external defibrillator and MRL Jumpstart (collectively "AED 10") | Welch Allyn Protocol; Inc | 2 | Issue #1: AED 10 displayed a "Do Not Use" symbol as a result of a blown F1 fuse.Issue #2: AED 10 models that incorporate a spring connector with the speaker experienced intermittent electrical connection to the circuit board causing a loss of voice prompts. | Welch Allyn sent an Urgent Medical Device Recall letter dated February 26, 2009 to all distributors and customers, describing 5 failure modes. Customers were instructed to respond to the letter by calling 888-345-5356, and decide on replacement or discounted new purchase.
Z-1252-2009 | May-12-2009 | 2009 | SERVO VENTILATOR 300 - Servo Ventilator 300 (Model SV300) andServo Ventilator 300 (Model SV300A).The SV300 is intended for general and critical ventilatory care for use on adults; pediatric; infant and neonatal patients. The unit is designed to be used at the bedside for intra-hospital transport. | MAQUET Inc. | 2 | The expiratory valve in the SV 300 ventilator may stay in the open position if the printed circuit board fails. | An "Urgent: Device Correction" letter dated March 23, 2009 was sent by Federal Express to customers. The letter described the affected product, problem, indications, potential hazard and actions for users. Customers were instructed to complete the "Device Correction -- Customer Fax Response" form and return via fax (732-667-1903) to Maquet Inc. RA/QA. Direct questions to your local Maquet Inc. representative.
Z-0128-2010 | Nov-09-2009 | 2009 | Medtronic Virtuoso DR - Medtronic Virtuoso DR; D154AWG; Dual Chamber implantable cardioverter defibrillator with atrial and ventricular therapies. Medtronic; Inc.; 710 Medtronic Parkway; Minneapolis; MN 55432-5604 USA. This device is indicated to provide ventricular antitachycardia pacing and ventricular defibrillation for automated treatment of life-threatening ventricular arrhythmias in patients with NYHA functional Class II/III heart failure. In addition; the device is indicated for use in the above patients with atrial tachyarrhythmias; or those patients who are at significant risk of developing atrial tachyarrhythmias. | Medtronic Inc. Cardiac Rhythm Disease Managment | 2 | A subset of active implanted Concerto CRT-D and Virtuoso ICD devices worldwide (estimated at 6;300) may not meet expected device longevity due to gradually increasing current drain caused by low voltage capacitor degradation. This issue may present in the affected devices as reaching the Recommended Replacement Time (RRT) earlier than projected. This issue does not compromise device functionalit | An "Important Medical Device Information" letter, dated September 2009 was hand delivered by Medtronic Representatives beginning 09/09/09. The letter describes the issue, identifies affected devices and provides recommendations for patients with devices in the affected subset.. A Patient letter was mailed to patients beginning 09/28/09. For additional information: Doctors are advised to contact your local Medtronic Representative or Medtronic Technical Services at 800-723-4636; Device tracking managers are advised to contact your Medtronic representative, or the Patient Registration Department at 1-800-328-2518 x41832; and Patients are advised to contact Medtronic Patient Services number at 800-551-5544 ext. 41835.
Z-1421-2009 | May-27-2009 | 2009 | NeuroStar TMS Therapy System/SenStar Treatment Link - NeuroStar TMS Therapy System/SenStar Treatment Link.The SenStar Treatment Link is a single use; disposable component of the NeuroStar TMS system; a non-invasive treatment for depression using highly focused; pulsed magnetic fields to stimulate function in targeted brain regions. | Neuronetics; Inc. | 2 | SenStar Treatment Link may have an open circuit that during use with the NeuroStar TMS Therapy system may cause arcing and heat damage to the SenStar and/or treatment coil. Heating of the SenStar during use has a potential risk of minor scalp burn to patients. | Neuronetics Inc. notified consignees by telephone and sent a follow-up fax/letter dated March 19, 2009 titled "Urgent Product Recall Notification." The letter describes the issue and actions for consumers to take in response to the recall. Neuronetics will visit each customer in person to complete a "Returned Goods: SenStar Recall Inventory Form." Customer questions should be directed to Neuronetics Customer Service at 1-877-600-7555.
Z-2008-2009 | Sep-04-2009 | 2009 | enFlow Fluid/Blood Warmer - Model 100 enFlow Fluid/Blood Warmer Model Numbers: 980100; 91000103; 91000153; 91000154; 91000156 The device is a medical point-of-care device designed to warm blood and IV fluids. | Enginivity LLC | 2 | 1. Fluid contamination of the electronic circuitry can cause short circuits resulting in permanent fuse failure or localized circuit board overheating with possible smoldering and discoloration. 2. Electromagnetic (EMI) Interference associated with the Warmer may produce artifact in ECG; EEG; or EMG records (cardiac or neuro monitoring). | Vital Signs, Inc. issued Product Safety Notification on March 23, 2009 informing users of potential EMI and fluid ingress issues. All warmers in the field will be replaced beginning in May 2009.
Z-0035-2010 | Oct-29-2009 | 2009 | WarmAir 135 - WarmAir 135 Thermal Regulating System. Catalog Numbers: 86186-115 V; 06186-115 V (sold in the U.S.) and #86185-100 V ( not sold in the U.S.). Intended for warmingpatients before; during and after surgery; and other specialty procedures. | Cincinnati Sub-Zero Products Inc | 2 | The blower on the WarmAir 135 Warming Unit may overheat. Resistance in a connector between circuit board and wiring harness results in excessive heat that may melt connector or deform the outer casing of the blower unit. | The firm sent an "URGENT Medical Device Recall' notice via First Class Mail on 2/16/2009 (Note: the letter is dated 2/13/2009). The notice informs the customers of the potential for the device to overheat and of the 'small number of customer complaints received by CSZ concerning this issue. In addition, the notice asks that the customers inspect their units for signs of discoloration indicative of a faulty connector. In the notice, the customers were instructed to return the product to CSZ for repair by contacting their Customer Service division and obtaining a Return Authorization number and specific instructions concerning packaging and returning of the unit(s) for repair. The customer is also instructed to complete the attached Recall Response Form which accompanies the recall notification. The notification provides that customer with a toll-free number at (800) 989-7373ext.3234 or call collect at (513) 719-3234 to reach CSZ representative between the hours of 8:30 a.m. until 4:30 P.M. Monday through Friday EST. Once the customer ships the recalled unit to CSZ, the firm will repair the unit(s) using new connectors manufactured, after which time, the repaired unit is returned to distribution for customer use.
Z-0368-2010 | Nov-19-2009 | 2009 | Biosense NaviStar RMT Catheters - Biosense Webster NaviStar RMT Diagnostic/Ablation Steerable Tip 4mm Catheters; Part Numbers: D-1257-01 & D-1257-02 | BioSense Webster; Inc. | 2 | Biosense Webster initiated the recall because of the potential concern with one of the Printed Circuit Board (PCB) components used in the NaviStar RMT Catheters; which may cause the ablation catheter icon to shift on the map viewer when RF energy is being applied. Therefore the user; attempting a linear ablation in the location displayed on the CARTO RMT screen; may ultimately apply RF energy to u | The firm sent an "Urgent Medical Device Correction" letter to consignees on October 6, 2008 informing consignees. Users were advised to pay "close attention to the location of the ablation catheter icon on the map viewer when applying RF energy." If consignees suspect that they have a defective NAVISTAR RMT Catheter, they were instructed to notify Biosense Webster and to return the defective catheter for evaluation and refund/replacement. Customer were instructed to contact their local Biosense Webster representative for any questions or further information. No return reply was requested.
Z-0883-2009 | Jan-28-2009 | 2009 | Defibrillator PCA wihtin the monitor of the LifeVest - Monitor portion of the LifeVest device - specifically the Defibrillator printed circuit assembly (PCA) within the monitor.The LifeVest is indicated for adult patients who are at risk for sudden cardiac arrest and are not candidates for or refuse an implantable defibrillator. | Zoll Lifecor Corporation | 2 | The charging circuit will not shut down properly. | Physicians with active patients were notified of the product removal by letter via first-class mail dated 10/21/2008. The letter describes the problem and the potential risk. For additional information, contact ZOLL Lifecor Corporation at 1.800.543.3267.
Z-1621-2009 | Jul-17-2009 | 2009 | Implantable Cardioverter Defrillator - GUIDANT VITALITY¿ DS VR Model T135 VR; EL Model T127 DR; 2ICD Model T165 DR; 2ICD Model T175 VR; 2EL ICD Model T167 DR; 2 EL ICD Model T177 VR; DS Model T125 DR. Implantable Cardioverter Defibrillator; Sterile Lot. Guidant Corporation; Hamline Avenue North; St Paul; MN 551124-6798.Indicated for patients who are at risk for sudden cardiac death due to ventricular arrhythmias and may require pacing support. | Boston Scientific CRM Corp | 2 | In April 2007; Boston Scientific CRM communicated with physicians regarding the potential for reduced Elective Replacement Indicator (ERI) to Battery End of Life (EOL) time in a subset of implantable defibrillators due to degradation of a low-voltage capacitor. Since that time; the April 2007 advisory population has not experienced any clinically significant changes to either the rate of occurrenc | Boston Scientific CRM notified Doctors via sales representatives of an "Urgent Medical Device Information" letter dated March 4, 2009. The letter described the problem and included Description of Issue, Rate of Occurrence, Patient Management Recommendations, Devices Affected, Warranty Program & Further Information. A Boston Scientific " Product Advisory" letter dated March 4, 2009, was also made available to physicians to distribute to patients. The letter informed the patient that their defibrillator was included in a group of devices that might experience a performance issue. For further questions, Contact Boston Scientific CRM at 1-800-227-3422.
Z-0670-2009 | Jan-29-2009 | 2009 | Terumo APS 1 - Terumo Advanced Perfusion System 1 Air Bubble Detection Module; Terumo Cardiovascular Systems; Ann Arbor; MI; Catalog number 802110.For use for up to 6 hours in the extracorporeal circulation of blood for arterial perfusion; regional perfusion and cardiopulmonary bypass procedures. | Terumo Cardiovascular Systems Corporation | 2 | The system may malfunction due to t-filter corrosion on an internal circuit board. | U.S. consignees were notified of the problem by letter on 11/16/07. International affiliates were informed via email on 11/21/07. Consignees were informed that Terumo would schedule appointments to install new pressure modules and reconfigure the systems beginning in Feb. 2008 or that parts would be shipped to consignees wishing to perform the work themselves. For additional information, contact Terumo CVS Customer Service at 1-800-521-2818.
Z-1570-2009 | Jun-30-2009 | 2009 | Ethylene oxide gas sterilizer - Sterrad 50 Sterilizer; Product Code 10050.Low-temperature general-purpose sterilizer used to sterilize heat and moisture sensitive reusable medical devices. | Advanced Sterilization Products | 2 | The capacitor in the vacuum pump component within the Sterrad 50 System could overheat and potentially create a hazardous condition inside the system; resulting in smoke and/or fire. | Advanced Sterilization Products (ASP) sent an "Urgent: Product Correction" letter dated January 12, 2009 to each consignee informing them of the issue and instructed them to report smoke or a burning odors in or around their STERRAD 50 System. ASP strongly recommended for consignees to follow their facility's safety protocols and if possible, immediately disconnect the STERRAD 50 System from its electrical source. Customers were instructed to contact ASP to schedule an onsite visit to investigate the issue. If the issue is caused by a malfunctioning capacitor in the vacuum pump, the firm will replace the malfunctioning component right away. To correct this situation, an ASP Field Service Engineer (FSE) will visit all customers facilities within the next sixty days to replace the capacitor in the vacuum pump on the affected STERRAD 50 System.
Z-0909-2009 | Feb-24-2009 | 2009 | Terumo Advanced Perfusion System 1 - Terumo Advanced Perfusion System 1; 100V -120V; 15 A (circuit breaker); 50/60 Hz (20A power source required) base. Terumo Cardiovascular Systems; Ann Arbor; MI. Model 801763.The device is indicated for use in the extracorporeal circulation of blood for arterial perfusion and cardiopulmonary bypass procedures by a perfusionist. | Terumo Cardiovascular Systems Corporation | 2 | The power manager board may cause a short circuit; resulting in loss of system power; loss of central control monitor or electrical arcing. | An "Urgent Medical Device Correction" letter was sent to consignees on November 7, 2008. The letter described the reason for correction and provided instructions to customers. Terumo Cardiovascular Systems Corporation will replace the power manager boards on the affected devices. Customers were instructed to complete and fax the Customer Response Form to CVS Customer Service at 1-800-292-6551 (US only) or respond by telephone at 1-800-521-2818. Please direct questions to CVS Customers Service at 1-800-521-2818.
Z-0176-2010 | Nov-11-2009 | 2009 | RESPIRONICS V60 VENTILATOR - RESPIRONICS V60 VENTILATOR | Respironics California Inc | 2 | The recall was initiated to correct an issue with a component on the V60 Ventilator Power Management (PM) Printed Circuit Board (PCB). A component on the PCB has the potential to break off or become damaged if the PCB is removed from the ventilator. If this occurs; the ventilator may not properly switch between back up battery power and AC power. This can result in power loss during ventilation fo | The recall notification was initiated October 26, 2009 with the firm mailing a Urgent Medical Device Field Correction letter to the affected customers. The letter informed the customers of the reason for the notice and how Respironics Field Service Technicians were going to correct this issue. The letter also gave the customers instructions to follow pending the completion of the field inspection/replacement. For questions or additional information contact Respironics US Customer Service at 1.800.345.6443, Option #5 then Option #1.
Z-1600-2009 | Sep-04-2009 | 2009 | Central Station Monitoring System - GE Healthcare CIC Pro Clinical Information Center Central Station (CIC Pro version 5.1; MP100 Hardware).The CIC Pro" clinical Information Center central station is intended to collect information from a network and display this date. This data includes physiological; patient dermographic and/or other non-medical information. | GE Medical Systems Information Tecnology | 2 | GE Healthcare has become aware of a potential issue associated with the failure to restart CARESCAPE CIC Pro v5.1 on the MP100 hardware. This issue may impact patient safety. | A GE Healthcare "Urgent Medical Device Correction" letter dated February 24, 2009 was sent to consignees. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction, and Contact Information. Affected units will be replaced with new units by GE Field Service team members. For questions regarding this notice or experience a failure to restart the device, please contact GE Healthcare Technologies Support Center at 1-800-558-7044, enter Option 2 and ask for Technical Support or your local GE Field Service Representative.
Z-1399-2009 | May-26-2009 | 2009 | Digital Fluidics PCB - Immulite 2000 and Immulite 2500 digital Fluidics PC Board;Part Number 450116-0002.Performs chemiluminescent immunoassays using serum; plasma; and/or urine for in vitro diagnostic testing. | Siemens Medical Solutions Diagnostics | 2 | Design Flaw: The redesign of the Digital Fluidics PCB used in IMMULITE 2000 and IMMULITE 2500 is missing a jumper. | Siemens Healthcare Diagnostics Inc. issued an "Urgent: Device Recall" notice dated January 2009 mailed to all direct customers of the affected product. The letter described the problem, provided "work-around" instructions and also provided a Recall Response Form to be returned to the firm via fax. For further information, please contact your local Siemens Healthcare support representative or Siemens Healthcare Diagnostics Inc. directly at 1-973-927-2828.
Z-0583-2010 | Dec-29-2009 | 2009 | Symbia S Series SPECT System - Symbia S Series SPECT System; Single-Photon Emission Computed Tomography; Siemens Medical Solutions USA; Inc.; Hoffman Estates; IL 60195-2061.Indications for use: SPECT + CT: Perform scans and nuclear imaging studies with the same instrument. To obtain attenuation corrected images and to provide registration of anatomical and physiological images within the patient's anatomy. | Siemens Medical Solutions USA; Inc. | 2 | Some of the hand controllers of the Symbia S & T systems are missing a resistor switch; and the absence of the switch may lead to unintended system motion during camera set up activities. | Siemens issued an Urgent Field Correction Recall letters dated November 18, 2009 and December 11, 2009 to the affected users of the Symbia S and Symbia T Systems, respectivley. The customers were advised of precautionary measures and that Siemens will replace the affected device. They also were advised that a Siemens Customer Service Representative would contact them within a few days to schedule a replacement of the device. For further information, contact Siemens Medical Solutions USA, Inc. at 1-800-888-7436 (USA).
Z-1933-2009 | Aug-18-2009 | 2009 | Boston Scientific Teligen - Boston Scientific TELIGEN; Implantable Cardiac Defibrillator; Models E102 and E110. Sterilized using ethylene oxide. Boston Scientific; 4100 Hamline Avenue North; St. Paul; MN 66112-5798; USA.The product is indicated for patients who are at risk for sudden cardiac death due to ventricular arrhythmias and may require pacing support. COGNIS devices also provide bi-ventricular pacing. | Boston Scientific CRM Corp | 2 | Boston Scientific Cardiac Rhythm Management retrieved devices that could be subject to a potential for acute non-secure lead connections when implanted. Nonsecure lead connections can manifest as high impedance; or lead electrogram artifacts resulting in oversensing and/or inappropriate delivery of therapy. | Boston Scientific Corporation field personnel began retrieving devices containing the original set screw system on Friday, July 17, 2009. If needed a letter, dated 07/17/2009 was left with the hospital administrator letting them know that product has been collected for return to Boston Scientific Corporation. Replacement devices will be provided at no charge. If the field personnel was unable to contact the hospital in the time provided, a notification letter dated July 22, 2009 was sent via Federal Express to the hospital. This letter was directed to the EP Catheter Lab Manager with a cc to the Hospital Risk Manager, Hospital Administrator and Purchasing and Central Supply. The letter asked that the affected devices not be implanted and to contact their local representative for pick-up and return. The firm will replace the affected devices. Please direct questions to your local Sales representative or Boston Scientific Corporation CRM Technical Services by calling 1-800-227-3422.
Z-1118-2009 | Sep-25-2009 | 2009 | Definium 5000 X-ray System - Definium 5000 X-ray System; model number 5220493 and 5220393-2 | GE Healthcare | 2 | A defect in the AEC board in the system's control causes failure to terminate additional exposures at the exposure time selected for the initial exposure. | A notification was sent to all affected customer sites alerting them to this issue, with instructions for safe use of the affected systems.
Z-0561-2010 | Dec-23-2009 | 2009 | Xoran MiniCAT System for ENT - Xoran MiniCAT System for ENT (server and work station); Xoran Technologies; Ann Arbor; MI. Cone beam CT for sinus; temporal bones; and maxillofacial imaging. | Xoran Technologies; Inc. | 2 | The system may overheat; causing poor image quality; interruption in image acquisition and/or damage to the system. Also; if the system is not allowed to cycle down prior to being restarted after an emergency stop; the x-ray and gantry motions will both resume. | A letter dated 9/17/08 was sent to each account telling them that the PLC would be upgraded. The firm visited each account and made corrections.
Z-0083-2010 | Oct-28-2009 | 2009 | Foot Pedal - Foot Pedal; Foot Control for SC2000 and SC2100 Powered; simple cranial drills; burrs; trephines; and their accessories. Catalog Numbers: EMAX2-FP and FPLUS-FP.Indicated to cut and shape bone including bones of the spine and cranium. | The Anspach Effort; Inc. | 2 | The Anspach Effort; Inc. is recalling the Foot Control/ Foot Pedal for SC2000 and SC2100 Powered; simple cranial drills; burrs; trephines; and accessories. The Foot Control Cables lacks the shield connection from the cable to the connector body. The lack of shielding may result in radiated emissions which could possibly cause interference on other electrical equipment adjacent to the foot pedal | Anspach issued an "Urgent Product Removal" Letter dated August 17, 2009 informing customers of the affected product. Consignees are asked to screen their inventory and return any affected product including a completed Product Replacement Form to the firm. For further questions, contact Anspach at 1-800-327-6887 or 1-561-627-1080.
Z-1953-2009 | Aug-18-2009 | 2009 | Medtronic Virtuoso - Medtronic Virtuoso DR; Model D154AWG; Dual Chamber implantable cardioverter defibrillator with atrial and ventricular therapies. Medtronic; Inc.; 710 Medtronic Parkway; Minneapolis; MN 55432-5604 USA.This device is indicated to provide ventricular antitachycardia pacing and ventricular defibrillation for automated treatment of life-threatening ventricular arrhythmias in patients with NYHA functional Class II/III heart failure. In addition; the device is indicated for use in the above patients with atrial tachyarrhythmias; or those patients who are at significant risk of developing atrial tachyarrhythmias. | Medtronic Inc. Cardiac Rhythm Managment | 2 | Medtronic has detected a specific pattern of MOSFET IC malfunctions in its Concerto; Virtuoso and EnRhythm family of devices. The probability of occurrence decreases with time and; to date over 90% of the malfunctions related to the pattern have occurred within the first twelve months after implant. | Medtronic, Inc. personnel retrieved 2 non-implanted devices remaining at 2 hospitals. A Performance Notes letter was left at each facility. The Performance Notes described the issue, actions to address issue and recommendations to physicians. The Performance Notes will also be posted on the firm's website. Direct questions about this recall to Medtronic, Inc. by calling 1-763-514-4000.
Z-1792-2009 | Aug-03-2009 | 2009 | ALARIS PC UNIT - The Alaris PC Unit; Model 8015 (manufactured or serviced between April 2008 and January 2009).The Alaris PC Unit is the central programming and power supply component for the Alaris System modules. | Cardinal Health 303 dba Cardinal Health | 1 | There is a potential in the specified ESD (electrostatic discharge) protection circuitry in the keypad of Alaris PC units manufactured or serviced between April 2008 through January 2009. In these units there is a potential for the keypad not being responsive; key entries without key presses; or key entries that register incorrectly which may result in an over-infusion or under-infusion. An over | The recall was initiated on June 12, 2009 with the firm forwarding via UPS an Urgent: Medical Device Recall Notification with attached Customer Response Card to the Chief Administrative Officer, Director of Pharmacy, Director of Biomedical Engineering, Director of Nursing & Director of Risk Management of all of the affected consignees. Customers were also asked to please promptly complete and return the enclosed Customer Response Card to expedite the correction process. The Recall Notification informed the consignees of: (A) PROBLEM #1: Occlusion Warning Message for Alaris Pump Module; (B) PROBLEM #2 Syringe volume Warning Message for Alaris PCA module; (C) PROBLEM #3 Electrostatic Discharge (ESD) protection circuitry; (D) PROBLEM #4 Fluid Ingress and (E) PROBLEM #5 Alaris System Inter-Unit Interface (IUI) Connectors Labeling. The recall notice included "Required Action for Users & Follow-up actions by Cardinal Health for all five problems. Customers were also asked to report any adverse reactions experienced with the use of this product, and/or quality problems to the FDA's MedWatch Program. In addition, customer were provided the following contact information for the recall: (1) Cardinal Health Recall Center for general questions: By phone at 1-888-562-6018 7am to 5pm (Pacific) ; (2) Customer Advocacy for adverse reports: By phone at 1-800-854-7120, Option 1, Option 1, Option 3 or by email at customerfeedback@cardinalhealth.com 24 hours a day, Sunday-Saturday; (3) Technical Support for Technical Questions regarding the Alaris System: By phone at 1-888-562-6018, 6am-5pm (Pacific). The firm followed-up with the release of a nationwide press statement regarding this issue on July 29, 2009. The press release indicated serial numbers of affected devices, as well as CareFusions short term instructions to customers, and the firms strategy to fix the affected devices can be found at: www.cardinalhealth.com/alaris/medical-device-recall/ (This informatio.
Z-2030-2009 | Sep-14-2009 | 2009 | LIFEPAK CR Plus Defibrillator/Monitor - Medtronic Physio-Control LIFEPAK CR Plus Defibrillator/Monitor.Automated external defibrillator deployed primarily in public access defibrillator environments where the expected use for any one device is once every 20 months. For use on persons requiring defibrillation therapy. | Physio Control; Inc. | 1 | The K1 relay on the Analog printed circuit board assembly may short due to moisture ingress in a highly humid environment. K1 shorting may affect the ECG amplitude causing the Shock Advisory System(SAS) to not analyze the ECG rhythm correctly thus preventing the device from delivering therapy. | Physio-Control, Inc. notified Customers by phone beginning August 18, 2009. In addition, the recalling firm Issued an "Urgent Medical Device Recall" letter dated August 2009 informing Customers of the affected product and informing them that a Physio-Control representative will call and arrange for immediate device replacement. For further information, contact Physio-Control Customer Care at 1-800-442-1142.
Z-1646-2009 | May-26-2009 | 2009 | Roche ACCU-CHEK Spirit Plus insulin pump - Roche ACCU-CHEK Spirit insulin pump. Part numbers 4540492001; 4759702001 and 4759729001. | Disetronic Medical Systems; Inc. | 2 | The "up" and "down" buttons; which are used for changing the program in the menu or to administer additional insulin through a bolus delivery; may experience intermittent failure or completely cease to operate. | The firm issued a press release on 4/30/09 and sent letters with attached Urgent Medical Device Recall (Field Corrective Action) notification dated 5/4/09 to distributors, health care providers and users. Disetronic will replace pumps within the affected serial numbers (SN02119552 to SN10006093) that have intermittent or complete loss of function of the up and/or down button promptly upon contact from the customer with a new pump that is not affected by the potential button failure. For pumps within the affected serial numbers (SN02119552 to SN10006093) which are functioning properly, no replacement is available/possible at this time. Disetronic will gradually phase in replacement pumps for the remaining pumps within the affected range, as they become available. Customers are to complete the enclosed Customer Reply Card to acknowledge receipt of the notification. On 5/17/11, after being notified of the recall's ineffectiveness (in that only 40% of the pumps had been replaced), the firm re-issued recall notification letters including the aforementioned information to their consignees.
Z-1426-2009 | Jun-23-2009 | 2009 | Advance - GE Datex-Ohmeda Advance Anesthesia Gas-Machine. | GE Healthcare | 2 | Unintended shut-down: The anesthesia machine may start shutting down without human intervention due to a faulty On/Standby Switch. Prior to shutdown; the system will issue an alarm and play a message that shutdown will occur in 8 seconds. The failure only affects the electrical portion of the switch. If the clinician does not intervene; the patient could experience hypoventilation leading to h | GE Healthcare contacted consignees via two "Urgent Medical device Correction" letters dated February 5, 2009 and March 10, 2009. The letters are addressed to 3 titles within the affected accounts; Healthcare Administrator/Risk Manger, Chief of Intensive Care and Director of Biomedical Engineering. The letters describes the Safety Issue, Affected Product details, Safety Instructions, Product Correction and Contact Information. The second letter dated March 10, 2009 includes 700 serial numbers that were missing from the first letter. The second letter was only mailed to the consignees of the additional 700 units.
Z-1019-2009 | Mar-09-2009 | 2009 | Medfusion Syringe Infusion Pump - Smiths Medical Medfusion Model 3500 Syringe Infusion Pump; Software version V4.1.5. Manufactured by: Smiths Medical MD; Inc.; St. Paul; MN 55112 USA. The Medfusion 3500 pump is a small; lightweight and portable syringe infusion pup. The pump offers a variety of delivery modes to program to meet specific patient care needs.The Medfusion 3500 Syringe infusion pump is designed for the administration of fluids requiring precisely controlled infusion rates. The Medfusion 3500 syringe pump is indicated for use with the following delivery routes: Arterial; Epidural; Enteral; Intravenous; Spinal; Subcutaneous for delivering the following fluids: Intravenous solutions; drug Solutions; blood or blood products and enteral feedings. | Smiths Medical MD; Inc. | 2 | Smiths Medical has become aware of an increased trend in reports of keypad issues with Medfusion Syringe Infusion Pump; model numbers 3500-306 and 3500-415 ("Pump") which have occurred during implementation or shortly after first being placed into use. These issues may render the Pump inoperable or the Pump may not respond to certain keypad button presses. These issues can be intermittent. Smit | Consignees were initially notified of the recall via a telephone conversation beginning December 18, 2008 and continued through January 2009. After the initial telephone contact, an "Urgent Medical Device Recall" letter dated December 22, 2008 was sent via electronic mail. The letter was addressed to "Affiliates, Distributors, Clinicians, and other Customers". The letter gave a Description of the Problem and Advise on Action to be Taken by the User. Direct questions to Smiths Medical Customer Service department at 1-800-426-2448 (option 1, then option 3). A second Smith Medical "Urgent Medical Device Recall Additional Affected Units"recall letter was sent after the recall was expanded. The second letter dated March 6, 2009 was addressed to Safety/Risk Mangers, Healthcare Providers, Biomedical Engineers, Distributors, Clinicians and other Customers. The letter listed the Details of Affected Devices, Description of the Problem, Advise on Action to be Taken by the User, Results of Investigation and Contact Details.
Z-0946-2009 | Jan-30-2009 | 2009 | Epicor LP Connecting Cable; CC-6-LP-CE - Epicor LP Connecting Cable; CC-6-LP-CE; Catalog Number 12263; an accessory to the Epicor Cardiac Ablation Control System; manufactured by St.Jude Medical; Sunnyvale; CADisposable connecting cable; accessory to the Epicor Cardiac Ablation Control System. The cable is used to connect the Epicor disposable tissue device (UltraCinch or UltraWand) to the Epicor Cardiac Ablation control System. | St. Jude Medical Inc | 2 | Firm has received multiple reports of difficulties connecting the product and/or the presence of error code messages occurring with the use of the connecting cable and the Ablation Control System unit. | The firm issued notification by "Urgent Product Recall" letter dated 12/4/2008, to all consignees. The letter describes the issue, potential risk and states that the product should be returned. For additional information, contact St Jude Medical at 800-722-3423.
Z-0180-2010 | Nov-12-2009 | 2009 | LATITUDE¿ Communicator - Guidant LATITUDE¿ Communicator; model 6481; Boston Scientific 4100 Hamline Avenue North; St. Paul; MN 55112-5798 USA. The LATITUDE¿ Patient Management system enables physicians to periodically monitor both patient and device status remotely. The LATITUDE¿ Patient Management system is intended for use to remotely communicate with a compatible Guidant or Boston Scientific pulse generator and transfer data to a central database. | Boston Scientific CRM Corp | 2 | The L ATITUDE? communicator (Model 6481) for a patient may not have been able to receive measurements from their weight scale or blood pressure monitor afer initial setup. The issue has been resolved. If any patient weight or blood pressure readings are stored in the patient's sensor; they will be uploaded to LATITUDE after the patient's communicator calls in to Boston Scientific's servers. It | Boston Scientific used phone calls as the delivery method for these communications. The communication began August 6, 2009 for the original 6 patients affected. Notifications via phone calls to additional patients affected began on August 17, 2009. Direct questions about the recall to the LATITUDE Customer Service by calling 1-651-582-4000.
Z-0585-2010 | Dec-29-2009 | 2009 | Oxygen Sensor - Oxygen Sensor #5804; Distributed by: Teleflex Medical; Research Triangle Park; NC 27709 | Teleflex Medical | 2 | Sensor causes the oxygen monitor to display prematurely a Low Sensor warning which indicates an early maintenance requirement for the instrument. This may result in interrupted treatment. | Consignees were notified via next day delivery on or about 11/25/2009. Each customer received a letter requesting acknowledgement of the recall and a request to return the affected product. Each consignee that is a distributor was instructed to forward the letter to their consignees to retrieve relevant product from those locations. A second and third mailing to non-responding customers on the consignee list will be made, or as an alternative, the third contact may be by phone, e-mail, or fax. Affected products received by Teleflex Medical will be destroyed or reworked and replacement product or credit will be provided to the consignee. Records of contacts and of destruction or rework will be maintained. Questions are directed to the company at 1-866-246-6990.
Z-2325-2009 | Sep-30-2009 | 2009 | ARCHITECT i System Assay CD ROM - ARCHITECT i System Assay CD ROM - US; list 6E58-22; Version 22.01; and ARCHITECT i System Assay CD-ROM - WW (excluding US); list 6E59-23; Version 23.0; containing the ARCHITECT LH assay file; Abbott Laboratories; Diagnostic Division; Abbott Park; IL 60064 USA; Product List Number: 6E58 (US) and 6E59 (WW) Versions 23 and lower.The Architect LH assay is a Chemiluminescent Microparticle Immunoassay for the quantitative determination of human luteinizing hormone (LH) in human serum or plasma. | Abbott Laboratories | 3 | There continues to be an elevated level of complaint activity for calibration error code 1227; when using ARCHITECT LH reagent lots (6C25-22 or 6C25-27) due to the use of assay disk versions that do not contain the updated ARCHITECT LH assay file contained on ARCHITECT i System Assay CD ROM versions 24 and higher (US LN 6E58; WW LN 6E59). | Abbott issued a "Product Correction-Immediate Action Required" letter dated April 24, 2009. The letter provided users with the steps to take to determine what version of the ARCHITECT LH Assay is on their instrument and instructions for installing the assay file provided with the letter if needed. Users were requested to return a customer reply letter acknowledging receipt of the letter. For additional information, contact 1-877-4ABBOTT.
Z-1815-2009 | Aug-11-2009 | 2009 | Hospira Phoenix Infusion System with MedNet Software; One Channel Infusion System - Hospira Phoenix Infusion System with MedNet Software; Symbiq One-Channel Infusion System; Software versions v2.1 & v3.0; Manufactured by Hospira Inc.; Morgan Hill; CA. | Hospira Inc | 2 | Potential delay/underinfusion of critical therapy-- Devices experience increased frequency of alarms that result in temporary pump stoppage. | Hospira, Inc. issued an "Urgent: Device Field Correction" dated March 31, 2009 via Federal Express to Healthcare Professionals and Valued Hospira Customers. Consignees were informed of the affected devices and will be contacted by a Hospira representative to arrange necessary device upgrades. For further questions, contact Hospira Global Product Safety and Complaints at 1-800-441-4100.
Z-0372-2010 | Nov-23-2009 | 2009 | Philips SureSigns Patient Monitors - SureSigns Patient Monitors (VM3; VM4. VM6. VM8):VM3 3/5 lead ECG/SPo2; Respiration Bedside Patient Monitor with 8.4"display VM4 3 Lead ECG; NBP; SPo2; optional Predictive Temp Bedside Patient Monitor with 8.4" display VM6 3/5 Lead ECG; NBP; Temp; optional SPo2 Bedside Patient Monitor with 8.4" display or 3/5 Lead ECG; NBP; Temp; optional SPo2 option IBP Bedside Patient Monitor with 8.4" display VM8 3/5 Lead ECG; NBP; Temp; IPB; ET CO2; SPo2 Bedside Patient Monitor with 10" display or 3/5 Lead ECG; NBP; Temp; IPB; SPo2 Bedside Patient Monitor with 10" display Part number: VM3:863077;453564041661;453564050161;VM4:863063;453564024371;453564021961;453564024401;453564021971VM6:863064;863065;453564024421;453564022001;453564024441VM8:453564022041;863066;863068;453564022081; 453564024461; 453564024481The SureSigns VM3; VM4; VM6; VM8 and VS3 monitors are used for monitoring; recording and alarming of multiple physiological parameters of adults; pediatrics; and neonates in healthcare environments. Additionally; the monitor is intended for use in transport situations within a healthcare facility. | Philips Healthcare Inc. | 2 | Use of these devices at their maximum volume setting may result in the premature failure of the internal speaker | Customer notifications dated October 2009, in the United States will be delivered via UPS. Customer notifications outside the US will be managed by Philips representatives in each affected geography. Customers will be instructed to lower the volume setting while using existing units during the first notification. As part of the second notification to the customers, the field action correction will be a customer installable software upgrade. Customers may contact Philips Response Center with questions or concerns at 1.800.722.9377. On December 14, the firm issued the follow-up notification letter implementing the action plan to provide the software upgrade kit for installation and instructions.
Z-1390-2009 | Jun-02-2009 | 2009 | HeartStart MRx Defibrillator/Monitor - HeartStart MRx Defibrillator/Monitor; M3535A/M3536A Is used for the termination of ventricular tachycardia and ventricular fibrillation. The device is for use by qualified medical personnel trained in the operation of the device and qualified by training basic life support; advanced cardiac support; or defibrillation. It must be used by or on the order of a physician. | Philips Healthcare Inc. | 2 | Spontaneous turn-on which could deplete the battery; rendering the device unusable until power is restored. There is also the possibility of a failure mode in which the device fails to respond to user initiated turn-on; rendering it unusable for monitoring and therapy. | Philips notified customers via Urgent Medical Device Recall Letter & Notice dated March 2009. Customers are asked to follow the "Actions to be taken by customer/user section of the Urgent Medical Device Recall Notice. In this section, customers are given the following instructions: "During the interim period, as you await the upgrade for your device, you may continue to use the Heartstart MRx. However, to reduce the risk that a spontaneous turn-on could deplete the battery, it is advised that you keep: a charged spare battery with the Heartstart MRx AC power applied, if possible, when the device is not in use You should also consider identifying a readily available backup defibrillator to use in the unlikely event that the device fails to turn on. Devices that exhibit a spontaneous turn-on or failure to respond to user initiated turn on should be removed from service as soon as possible." Philips will replace the therapy switch in affected units free of charge per Mandatory Field Change Order FC086100077. All customers will be contacted by a Philips Healthcare representative to arrange for replacement of the therapy switch in the affected devices.
Z-0926-2009 | Jan-27-2009 | 2009 | Stryker PainPump2 BlockAid - Stryker PainPump 2 Blockaid; 400 mL PainPump for Continuous Nerve Block; sterile; 5 per box. Stryker; Kalamazoo; MI; REF 575-100; Product Number 0575-100-000.The device is used to deliver controlled amounts of medication and narcotics directly to the intraoperative site for pain management; the pump infuses the medication at an hourly flow rate and provides the option for patient controlled bolus doses. Medications and narcotics are infused through intramuscular; subcutaneous and epidural routs. It is also intended for controlled delivery of local anesthetics and narcotics in close proximity to nerves for postoperative regional anesthesia and pain management; in which case routes of administration may be intraoperative; perineural or percutaneous. | Stryker Instruments Div. of Stryker Corporation | 2 | The pump may display an E-4 error message; which causes the pump to stop dispensing the pain medication. | Consignees were notified via letter dated 11/12/08 and instructed to locate and destroy the affected lots. Customers were also instructed to complete and return the enclosed Business Reply form via fax (336-499-8603) to Stryker Instruments Regulatory Department. For questions regarding the recalled product, please contact Stryker Instruments at 1-800-800-4236, extension 3808 or extension 4056.
Z-1729-2009 | Aug-27-2009 | 2009 | Clinac with Version 7.x Software - Clinac with Version 7.x Software; manufactured by Varian Medical Systems; Palo Alto; CA.Part of the Trilogy Radiotherapy Delivery System for radiation therapy intended to delivery megavoltage x-ray treatments for conventional radiotherapy and stereotactic radiosurgery and radiotherapy. | Varian Medical Systems Oncology Systems | 2 | Unexpected Movement: if the stereotactic motion disable function is turned on; the couch can be moved via float mode unexpectedly. | Varian Medical Systems issued an "Urgent Medical Device Correction" dated May 29, 2009 distributed to all affected users, as well as Varian Sales, Marketing and Service organizations, informing them of the issue. Consignees were informed that a Varian Service representative will be in contact to schedule the proper correction. For further questions, contact your local Varian Customer Support District Manager.
Z-0003-2010 | Nov-09-2009 | 2009 | HEARTSTART MRX; MODEL M3535A - HeartStart MRx (M3535A/M3536A) Defibrillator/Monitor.The HeartStart MRx/MRxE is for use for the termination of ventricular tachycardia and ventricular fibrillation. The device is for use by qualified medical personnel trained in the operation of the device and qualified by training basic life support; advanced cardiac support; or defibrillation. It must be used by or on the order of a physician. | Philips Healthcare Inc. | 2 | incorrect default configuration settings; insufficientbattery power as the user may not be alerted that no battery is installed or that a low battery condition exists prior to disconnecting from AC or DC power | Philips Healthcare issued a Urgent Voluntary Medical Device Correction Notification on August 21, 2009 advising users of the problem and Customers are asked to follow the "Action to be taken by customer/user" section of the Urgent Voluntary Medical Device Correction Notification. In this section, customers are given the instructions to correct their devices. For further information or support concerning this issue, please contact Philips representative at 1-800-722-9377.
Z-1720-2009 | Aug-13-2009 | 2009 | CT Perfusion 4 on Advantage Workstation; CT and PET/CT consoles - GE Healthcare; CT Perfusion 4 on Advantage Workstation; CT and PET/CT consoles.Image analysis software package that allows the user to produce image data and to generate information with regard to changes in image intensity over time supporting CT Perfusion images. | GE Medical Systems; LLC | 2 | GE Healthcare has recently become aware of the need to adjust default settings associated with the use of CT Perfusion 4 on the Advantage Workstation; ST and PET/CT consoles that may impact patient safety.The default parameter settings in CT Perfusion 4 may not be optimal in the case of data acquired with longer time sampling intervals such as axial scans with inter scan delay (ISD) greater th | GE Healthcare issued an "Urgent Medical Device Correction" letter dated April 23, 2009. The letter was addressed to Hospital Administrators / Risk managers and Radiology department mangers. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information instructing the user to follow the enclosed User Guide Addendum to mitigate further problems. For further questions, contact GE Healthcare at 1-800-437-1171 (US), 0120-055-919 (Japan); for other countries, contact your local GE Healthcare field service representative.
Z-1094-2009 | Apr-10-2009 | 2009 | GE Healthcare Infinia Hawkeye - GE Healthcare Infinia Hawkeye(R); Model numbers:H2508PY H3000WN H3000WR H3000WW H3000WY H3000WZ H3000YM H3000YS Intended use: To perform general Nuclear Medicine imaging procedures for detection of radioisotope tracer uptake in the patient body; using a variety of scanning modes supported by various acquisition types and optional imaging features designed to enhance image quality in Oncology; Cardiology; Neurology and other clinical diagnostic imaging applications. | GE Medical Systems; LLC | 2 | There is a possibility for misregistration when brain scans are performed with Infinia with the Hawkeye option. This misregistration could present the potential for misdiagnosis when the patient headrest sags under the weight of a head. | Consignees were sent from September 2008 - Feb 2009 a GE "Product Safety Notification" letter dated 4/2008. The letter was addressed to Hospital Administrators, and Managers of Radiology/Cardiology, Radiologists/Cardiologists. The letter described the Safety Issue, Affected Product, Safety Instructions, Product Correction and Contact Information. It instructed consignees to continue to use the Infinia Hawkeye system normally for all types of Hybrid scans and stated that as a general practice tomograms should be reviewed carefully using the Xeleris Volumetrix processing protocol to ensure proper registration. Local sales or service reps should be contacted for further information regarding the notification.
Z-2232-2009 | Sep-21-2009 | 2009 | Capnostream20 - Oridion Medical Capnostream20 capnograph/pulse oximeter monitor and monitor with printer. Part numbers CS 08651; CS 08653; CS08657; and CS08659. | Oridion Medical 1987 Ltd. | 2 | 1) The monitor's ability to withstand electrostatic discharge (ESD) to exposed metal connectors on the back of the monitor which may cause the monitor's display to freeze.2) The unit generates electromagnetic radiation above the limits set by applicable standards for medical devices | Oridian sent a letter to all its distributors June 25, 2009, instructing them to gather all the units to be exchanged for factory reconditioned units that have the hardware and software enhancements and meet all performance requirements in full.
Z-0365-2010 | Nov-18-2009 | 2009 | PICIS - PICIS; CareSuite Family- OR Manager (4235); Preop Manager (4250); Anesthesia Manager (4220); PACU Manager (4245) & Critical Care Manager (1110); patient information systems. | Picis Inc. | 2 | Demographic data; most notably allergy and precaution data; can be overwritten with incomplete data or blanks by the interface between non-Picis clinical systems (your "HIS" or "CIS") and certain Picis applications; if that interface is not configured properly. | On April 8th 2009, Picis initiated a customer notification letter to the installed base of potentially affected users, explaining the potential source of patient data conflict and to provide support in choosing the appropriate configuration option and implementation. The notification letters included with this report package contain additional details regarding configuration options.
Z-2315-2009 | Sep-29-2009 | 2009 | ZOLL R Series Defibrillator/Pacemaker/Monitor - ZOLL R Series Defibrillator/Pacemaker/MonitorUsed by qualified medical personnel for the purpose of converting ventricular fibrillation to sinus rhythm or other cardiac rhythms capable of producing hemodynamically significant heart beats. | ZOLL Medical Corporation; World Wide Headquarters | 2 | Failure of Device to display ECG could delay delivery of defibrillation or pacing therapy. | Zoll notified domestic consignees on August 31, 2009 and international consignees on September 1, 2009 by an Urgent Medical Device Correction letter. Customers were instructed to continue to perform daily tests per the recommendations provided in the R Series Operator's Guide. In the event of a failure, customers are to remove the device from service and contact biomedical or clinical engineering department per their facility's protocol. Units that operate as expected can remain in service and back-up monitor should be available in the event of a failure. Zoll's technical support department should be contacted at 1-800-348-9011 or +1-978-421-9460 to schedule the return of affected devices with the defect for replacement of the suspect component.
Z-1648-2009 | Jul-07-2009 | 2009 | Symbia S Series SPECT System - Siemens Medical Solutions USA; Inc.; Symbia S Series SPECT System; Single-Photon Emission Computed Tomography; Part number 8717741. | Siemens Medical Solutions USA; Inc. | 2 | Patient injury can occur when the patient is lying on the patient bed and the automatic collimator exchanger is undergoing a collimator change operation at the launch of an acquisition workflow. Detector motions required for this change can cause collision with the patient or pallet if it is in brain scan position; i.e. the patient or pallet is into the field of view. | Siemens sent Urgent Field Correction Recall letters dated 1/20/09 to all users of the Symbia S and Symbia T Systems. The accounts were informed of the potential injury to patients while lying on the patient bed and the automatic collimator exchanger is undergoing a collimator change operation at the launch of an acquisition workflow. The customers were advised to perform all necessary collimator changes before the patient is positioned on the patient bed, and to examine each of their workflows that include a Tomo Acquisition activity and make sure the Auto Collimator Change check box on the Camera Parameters Tab is not checked. They also were advised that a Siemens Customer Service Representative would contact them within 60 days to schedule a corrective software upgrade free of charge. Any questions were directed to contact Siemens Medical Solutions USA, Inc. at 1-800-767-2313 (USA) for assistance. The accounts were requested to complete the Mandatory Fax-Back Form acknowledging receipt and understanding of the letter.
Z-1195-2009 | May-19-2009 | 2009 | Digital Fluoroscopic Imaging System - Innova 2121-IQ/3131-IQ Biplane Cardiovascular Imaging System. The Innova Biplane systems are the first and only digital flat panel biplane systems with a full sized lateral plane to cover lateral anatomy without requiring multiple contrast injections and radiation exposures. The product is indicated for use in generating fluoroscopic images of human anatomy for diagnostic and interventional cardiac angiography procedures. These systems can be operated in a mobile or fixed site environment. | GE Medical Systems; LLC | 2 | GE Healthcare has become aware of a potential image storage failure on Innova 2121 IQ/ 3131 IQ cardiovascular imaging systems that may impact patient safety. This intermittent communication failure on the image storage chain may result in the following situations: Record abort; Record not available; Fluorostore function not available; Review not available; Boot/Reset failure. Fluoroscopy imag | A GE Healthcare "Urgent Medical Device Correction" letter dated February 2, 2009 was sent to consignees via mail. The letter was directed to Hospital Administrator/Risk Manger, Managers of Radiology/Cardiology and Radiologists/Cardiologists addressing the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information. Further questions should be addressed to GE Healthcare at 1-800-437-1171.
Z-1931-2009 | Sep-16-2009 | 2009 | iPulse Circulatory Support System - iPulse Circulatory Support SystemCatalog number: 0036-0010The iPulse System Console is a multi-purpose mechanical circulatory support system used to support patients suffering from acute cardiovascular failure. The iPulse console can be used to drive exclusively one of two classes of therapeutic devices: either pneumatically actuated Ventricular Assist Devices (VADs) for total circulatory support; or (i.e.; non-simultaneously) an intra-aortic balloon (IAB) catheter for counter-pulsation therapy. | Abiomed; Inc. | 2 | iPulse Console SC1035 displayed a "Low Pressure / Low Flow" alarm and stopped pumping during patient transport in battery operation. | An "Urgent Voluntary Medical Device Field Correction Notice" was issued April 8, 2009 to affected customers. The customers were notified that the iPulse Console SC1032 displayed a "Low Pressure/ Low Flow" alarm and stopped pumping during patient transport in battery operation. Instructions were provided to users in case they encounter this failure mode. ABIOMED Field Service Engineers will contact customers to schedule a time to visit site and correct affected consoles. Please contact ABIOMED's Field Service at 1.800.554.8666, if you have any questions.
Z-0915-2009 | Feb-11-2009 | 2009 | MAKO SURGICAL TACTILE GUIDANCE SYSTEM (TGS) - Tibial Array and Femoral Array. Model Number(s): 0150ARR00001 and 0150ARR00002.Manufactured by: MAKO Surgical; Corp.The devices are accessories used with MAKO's Tactile Guidance System (TGS). The Tactile Guidance System is intended to assist the surgeon in providing software defined spatial boundaries for orientation and reference information to anatomical structures during orthopedic procedures. | Mako Surgical Corporation | 2 | The product in its current state may fail during use; which would cause the instrument to be in-accurate. | A MAKO employee will visually verify product at each account for any product affected by this recall. In addition, a recall letter dated July 1, 2008 was mailed to users (surgeon and owner of system, e.g. hospital). Customers were instructed to complete the enclosed Acknowledgement form and return it to your Clinical Technical Specialist (CTS) or in the enclosed pre-paid addressed envelope to the Mako Surgical Corporation. Please direct questions to your CTS or call 954-927-2044, extension 604.
Z-0927-2009 | Jan-29-2009 | 2009 | Artiste MV SA Linear Accelerator with syngo RT Therapist Version 4.1 build 67 and build 122 - Artiste MV SA Linear Accelerator with syngo RT Therapist Version 4.1 build 67 and build 122. Manufactured by Siemens Medical Solutions Oncology Care Systems; Concord; CA. | Siemens Medical Solutions USA; Inc | 2 | If a plan contains a sequence of beams; including at least 1 ARC treatment beam and if during the setup stage of a beam following an ARC beam in the delivery sequence a termination occurs; and the user selects the immediate resumption feature; syngo RT Therapist Express systems will display and record a second delivery of a prior beam and not the delivery of the current beam in certain circumstanc | An Urgent Medical Device Customer Safety Letter dated October 23, 2008 was issued to consignees via certified mail. The letter identified the issue and preventive measures for consignees. Monitoring will be done via return cards. Follow-up to be conducted via telephone.
Z-0653-2009 | Jan-27-2009 | 2009 | TheraMap - CAD SCIENCES LLC 3TP TheraMapTM Software; for oncology treatment planning; assessment and monitoring of the patient response. | iCAD; Inc. | 2 | Software modules not approved for this indication. | iCAD initiated a voluntary field correction by telephone and letter on 12/5/08 by disabling these features at all affected sites. Since Precision Point and TheraMap are controlled by a software license key, iCAD will disable these features by updating the license at the affected sites.
Z-0623-2009 | Sep-25-2009 | 2009 | MultiDiagnost Eleva with Flat Detector (MDE FD) - Philips Medical Systems MultiDiagnost Eleva Flat Detector System with software version 4.3.1 or higherUsed for general R/F; fluoroscopy; radiography; and angiography examinations. | Philips Medical Systems North America Co. Phillips | 2 | System fails to meet requirements when it is used at kV levels greater than 115 kV at certain frame rates and in certain acquisition modes; as grid switch leakage may occur. | A Customer Notification letter describing the failure, the means, by which a customer can prevent the failure mode from occurring, and actions planned by Philips to correct the problem were sent to all users of the affected systems. Philips will modify the system software and will be installed at no cost to the customer.
Z-1007-2009 | Mar-10-2009 | 2009 | AED 10 - Welch Allyn AED 10 automated external defibrillator and MRL Jumpstart (collectively "AED 10") | Welch Allyn Protocol; Inc | 1 | Potential for device to deliver less than the programmed energy. | Welch Allyn notified distributors by letter on 12/15/2008 requesting that they notify their customers. Welch Allyn sent an Urgent: Medical Device Recall letter to distributors and customers on February 26, 2009.
Z-1172-2009 | Apr-08-2009 | 2009 | Lead Integrity Alert - Medtronic RV Lead Integrity Alert ; Cat. # SW012; 1.0 software in EnTrust(R) (D153ATG; D153VRC; D154ATG; D154VRC) and EnTrust(R) Escudo (D144DRG; D144VRC) defibrillators. | Medtronic Inc. Cardiac Rhythm Managment | 3 | Medtronic has identified an issue related to installation or removal of the Lead Integrity Alert software in EnTrust(R) and EnTrust(R) Escudo" defibrillators. In those devices only; installation or removal of LIA will inadvertently turn off two audible patient alerts. | Between November 24 and December 9, 2008 US CRDM are responsible to hand deliver the consignee letter, a Medtronic "Urgent Patient Management Information," dated November 2008, to impacted physicians to make them aware of the issue and continue to promote and install LIA (by using the new LIA tipcard). For any physician listed that a confirmation sheet was not received by end of day, December 9, 2008, the letter was mailed to the physician on Dec 10, 2008. The letter described the issue and the product involved. The letter also gave recommendations and informed consignees in regard to a future update to the LIA software for EnTrust devices.
Z-0994-2009 | Apr-22-2009 | 2009 | VNS Therapy System Generator - Cyberonics VNS Therapy Demipulse Generator; Model 103; an implantable; multiprogrammable; vagus nerve stimulator. | Cyberonics; Inc | 3 | Reset/disabling of the VNS Therapy Demipulse Generator and Demipulse Duo Generator due to magnet interference resulting in the loss of stimulation. | Cyberonics notified consignees of recall via letter on 02/07/08. Consignees were notified of reset issue in which during a strong magnetic field, the software may report an error even though no error occurred. This error causes the generator to reset, which disables the generator (i.e., no stimulation delivered). Physicians should inform any potential and currently implanted patients of the possibility of this event. Patients should continue with their regularly scheduled visits as determined by their physician. If this event is suspected, physicians should interrogate the device as described in the Physicians Manual to confirm the generator has been reset and Cyberonics should be contacted for any additional instruction.
Z-0995-2009 | Apr-22-2009 | 2009 | VNS Therapy System Generator - Cyberonics VNS Therapy Demipulse Duo Generator; Model 104; an implantable; multiprogrammable; vagus nerve stimulator. | Cyberonics; Inc | 3 | Reset/disabling of the VNS Therapy Demipulse Generator and Demipulse Duo Generator due to magnet interference; resulting in the loss of stimulation. | Cyberonics notified consignees of recall via letter on 02/07/08. Consignees were notified of reset issue in which during a strong magnetic field, the software may report an error even though no error occurred. This error causes the generator to reset, which disables the generator (i.e., no stimulation delivered). Physicians should inform any potential and currently implanted patients of the possibility of this event. Patients should continue with their regularly scheduled visits as determined by their physician. If this event is suspected, physicians should interrogate the device as described in the Physicians Manual to confirm the generator has been reset and Cyberonics should be contacted for any additional instruction.
Z-0366-2010 | Nov-18-2009 | 2009 | Agfa IMPAX HeartStation ECG Management System - AGFA IMPAX HeartStation ECG Management System. | AGFA Corp. | 3 | HeartStation does not refresh content changes to the PDF reports that are printed emailed or faxed from the Index page. This may lead to a discrepancy in patient demographics between what is displayed in the HeartStation Client interface and the report delivered to the consumer. Questions are being directed to the local Agfa representative; the contact information was included in the letter. | An Urgent Safety Notice was sent via FED-EX to all the sites with the affected units (a total of 6) on May 7, 2009. The three foreign accounts were notified by letter on May 14, 2009. The letter describes the potential issue and mitigation. Acknowledgment via FAX-Back that the information was received and understood was requested from the sites. Customers who use the print, email, fax functionality of this product were verbally notified of the defect and they have been advised of a workaround. In the notice letter, Agfa HealthCare is also providing customers with the recommended workaround. The workaround is to only print from the Viewer screen or to print the ECG once confirmed. The Viewer screen, however, does not allow the user to print batches of reports as does the Index screen. Questions about this matter are being directed to the local Agfa HealthCare representative.
Z-1027-2011 | Aug-21-2009 | 2009 | MOSAIQ - SEQUENCER component of MOSAIQ versions 1.30 (all builds); 1.40 (all builds); 1.50 (all builds); 1.60 (all builds); and 2.0 (all builds). Manufactured by IMPAC Medical Systems; Inc. Sunnyvale; CA.The intended use of SEQUENCER (a verification system) is to assist in the process of patient treatment on a radiotherapy treatment machine. SEQUENCER provides the capability to notify the users of actions that need to take place prior to treatment; to display reference images for set up purposes; and to auto setup the machine to predefined settings. Actual settings are read from the treatment machine; via the machine's communication interface; and compared to the predefined values. If a mismatch occurs between the planned values and the actual machine settings; treatment is prohibited. At appropriate times during the treatment; the actual delivered values are recorded to provide treatment tracking. Treatments are initiated by the operator; via the machine control system; only after the user has verified that the treatment setup is correct. SEQUENCER does not bypass any checks put in place by the machine's control system. | Impac Medical Systems Inc | 2 | Potential for improper treatment. A software error may result in potential major underdosing to targeted areas from jaw blockage of the MLC aperture (jaw intrusion); and dosing to non-targeted areas from interstitial leaf leakage (jaw extrusion). | Impac Medical Systems Inc sent an Important Safety Notice letter to all affected consginees. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to contact Impact Software support as soon as possible if they have the affected configurations identified in the letter. Customers were asked to complete the form included with the letter and and fax to: European customers: +44 1293 654401 American & rest of world customers 702-992-5002 or e-mail to support@impac.com Customers were asked to distribute the Safety Notice to any and all users of IMPAC software at their organization who are potentially affected by this issue. For questions e-mail support@impac.com or call European customers: 44 1293 654320 American & rest of world customers: 800-488-4672.
Z-1237-2009 | May-01-2009 | 2009 | Alaris PCA Module and Alaris PC Unit with software versions 8 thourh 9.1 - Alaris Patient-Controlled Analgesia (PCA) Module (formerly Medley PCA Module) Model 8120 used with the Alaris PC Unit with software versions 8 through 9.1. | Cardinal Health 303 dba Cardinal Health | 2 | The Alaris PCA module can potentially infuse above or below the intended infusion dose. | The field action was initiated on March 11, 2009 with Cardinal Health sending by registered return receipt mail a Safety Alert to all customers using the Alaris PC Unit with software versions 8 through 9.1, which encompasses the subset of those customers that use the Alaris PCA Module. The Safety Alert was addressed to the Directors of Nursing, Pharmacy, Risk Management, and Biomedical Engineering at each hospital facility. If a customer did not respond to the Safety Alert, then Cardinal Heath notified that customer by telephone to verify receipt of the Safety Alert. This Safety Alert Informed the customers about a potential risk, and required them to: (1) Determine whether you use version 8 through 9.1 of the Alaris PC Unit. (2) Forward this Safety Alert to all Data Set owners (3) To immediately identify the potential risks review enclosed FAQs and Instructions for Data Set Review. (4)Continue to follow the Institute for Safe Medication Practices (ISMP) and the Joint Commission recommendation for safe medication administration - two nurses checking to ensure that the correct prescription is being entered into the PCA pump when first ordered and when there are any changes in the prescription (paying close attention to dosing units and infusion modes). The customers were also instructed to call our Support Center at 1-888-562-6018 with any questions. An updated Customer Recall Notification was forwarded by registered return receipt mail by Cardinal Health on April 15,2009. This notice reinforced the information in the first notice in addition to informing the customers of the hazard of an over-infusion or under-infusion of medication. Also, customers were given instruction for reporting any adverse reactions experienced with the use of the affected product and/or quality problems via the FDA's MedWatch Program.
Z-0481-2010 | Dec-03-2009 | 2009 | ezManager Max Diabetes Management System version 2.0.9 - Animas ezManager Max Diabetes Management System version 2.0.9 | Animas Corporation | 2 | Printed hard copy of results is not the same as the system screens- carbohydrate to insulin ratio incorrect. | Animas issued an Urgent Medical Device Correction letter dated 11/2/09 to the pump user and another letter to the health care professional that prescribed the pump. The letters explain the issue. The patient letter is accompanied with a CD that contains the upgraded software and instructions to follow.
Z-0988-2009 | Apr-16-2009 | 2009 | DELTEC COZMO INSULIN INFUSION PUMP WITH COZMONITOR BLOOD GLUCOSE METER - Deltec Cozmo¿ Insulin Pump; Model 1800.Deltec Cozmo¿ Insulin Pumps are ambulatory electromechanical pumps used for delivering insulin to patients with diabetes. The pump is used mainly by homecare patients but; can also be used in a healthcare facility. | Smiths Medical MD; Inc. | 2 | Smiths Medical has become aware of a display irregularity with the Deltec Cozmo? Insulin Pump; Model 1800 and Model 21-1700 containing Model 1800 software. There have been adverse events reported that the amount of Extended Bolus delivered was not accurately displayed by the Pump. The amount of extended Bolus displayed on the Pump Home Screen 2 and in the Bolus Summary Report is less than what w | Consignees were sent a Smiths Medical "Urgent Device Recall" letter dated November 20, 2008. The letter was addressed to Distributors, Clinicians, Patients and other Customers. The letter included a Description of the Problem, Advice on Action to be Taken by the User and requested the return of the Confirmation Form. In particular, users were given the option to check the "Replacement Pump" option on the Confirmation Form initiating the process for pump replacement with a recertified pump followed by a return of the defective product. The Confirmation Form can be faxed to 1-800-628-6322 in the U.S. or 1-651-628-7485 outside the U.S., completed at www.cozmore.com/notice or calling 1-800-501-5748 in the U.S.
Z-1232-2009 | May-04-2009 | 2009 | CIC PRO CLINICAL INFORMATION CENTER CENTRAL STATION - GE Healthcare CIC Pro" Software Versions 5.0.3; 5.0.6 ;5.0.7; 5.0.8 and 4.1.The CIC Pro" Clinical Information Center Central Station is intended for use under the direct supervision of a licensed healthcare practitioner. The intended use is to provide clinicians with adult; pediatric and neonatal patient data in a centralized location within a hospital or clinical environment. The CIC Pro Central Station is intended to collect information from a network and display this data. This data included physiological; patient demographic and/or other non-medical information. Physiological parameters and waveforms from GE Healthcare Information Technologies monitors and telemetry systems can be displayed and printed from the CIC Pro Clinical Information central station. Beat to beat patient information for all parameters and waveforms from he bedsides and telemetry systems can be displayed. The CIC Pro Clinical Information Center central station supports the ability to access information from GE Medical Systems Information Technologies products in web browser format. Additionally; the CICC Pro clinical Information central station supports the ability to access patient information collected from the unity network and stored on a network server. | GE Medical Systems Information Tecnology | 2 | GE Healthcare is aware of potential network time issues associated with the time change aspects of certain CARESCAPE CIC Pro" software versions. These issues may impact patient safety. No reported patient injuries or illnesses for the use of this device in conjunction with this device correction. | Consignees were sent a GE Healthcare "Urgent Medical Device Correction" letter dated February 27, 2009. The letter was addressed to the Chief of Nursing/Chief of Biomedical Engineering and Risk Manager/Healthcare Administrator. The letter describes the Safety Issue, Affected Product Details, Safety instructions, Product Correction, and Contact Information. Direct questions to your local GE Healthcare Field Service Representative or the GE Healthcare Technologies Support Center by calling 1-800-558-7044 (Option 2, Technical Support).
Z-0918-2009 | Feb-25-2009 | 2009 | Powerheart G3 - Cardiac Science Powerheart Automated External Defibrillator G3 9300E. | Cardiac Science Corporation | 2 | Potential for AED to not deliver defibrillation therapy. | On October 13, 2008 Cardiac Science began sending via Fed-Ex letters and software upgrade kits. Mailing of letters and software upgrade kits is anticipated to be completed by November 21, 2008. The letter instructs consignees to install the software update, assure that the AED information test file generated by the update software is returned to the firm, and complete the postage paid response card and return it to the firm. If the affected devices are no longer in the possession of the consignee, the consignee is to contact the firm and forward the update kit and letter to the person who has the AED in their possession. A separate letter is being sent to distributors. The letter advised to install the software upgrades for any units in stock or to forward the software upgrade along with the end user letter enclosed and that the AED information text file generated by the update software must be returned to the firm. The firm will visit the locations of individual national account customers and install the update. OEM partners will be notified and provided the information and update kits to provide to their customers.
Z-0264-2010 | Nov-16-2009 | 2009 | Cyberonics VNS Therapy Programming M250 System - Cyberonics VNS Therapy Programming M250 Programming System; model number 250; version 4.6; 6.1; 7.0 and 7.1. The Model 250 Handheld software is provided on flashcards packaged in software envelopes or pre-installed on handheld computers. | Cyberonics; Inc | 2 | Some VNS Therapy System replacement Demipulse generators reporting low lead impedance readings. In rare instances; a system diagnostic test using Model 250 Programming Software (versions 7.1 and earlier) may report "Lead Impedance: OK" when a short-circuit condition exists. | Cyberonics sent an Important Safety Notice on July 24, 2009, to all physicians explaining the event, risks presented and recommended actions to take. Labeling updates are being developed concurrently with the notification. The firm will mail or hand-deliver the notifications via firm representatives.
Z-0179-2010 | Nov-12-2009 | 2009 | System 1000; TINA; AURORA and AltraTouch - System 1000 family of Hemodialysis Instruments; including the System 1000; TINA; AURORA and AltraTouch; Baxter Healthcare Corporation; Deerfield; IL 60015 U.S.A.The Hemodialysis and Continuous Renal Replacement Therapy devices are used for acute and chronic hemodialysis; including high flux hemodialysis. The device is intended for use by trained operators when prescribed by a physician. The device is intended to be used with hollow fiber or parallel plate dialyzers. | Baxter Healthcare Renal Div | 2 | When the main power is lost temporarily; a software anomaly prevents the instrument from automatically changing the ultrafiltration rate (UFR) to the minimum value as described in the Operator's manual. | Baxter sent an Urgent Device Correction letter dated August 31, 2009 to all System 1000 Series of Hemodialysis Instruments customers via first class mail on the same date, to the attention of the Hemodialysis Administrator. The letters informed the accounts that a software anomaly in software versions 4.B.0.9, 4.B.1.1 and 4.B.1.3 was discovered when the main power was lost temporarily during a simulated treatment. The anomaly prevents the instrument from automatically changing the ultrafiltration rate (UFR) to the minimum value as described in the Operator's Manual. The accounts were informed that instruments with software versions 4.B.0.9, 4.B.1.1 and 4.B.1.3 will be upgraded to version 4.B.1.4 software, provided the instruments have a 486 Motherboard or Single Board Computer, and that instruments containing versions of 4.A software are not affected. The accounts were requested to complete the enclosed Customer Reply form, indicating the instrument serial number and software version for each machine at the site, and if they would be requesting Baxter Technical Services to upgrade their instruments or have the software sent to them to make the upgrades, and fax the sheet to 1-847-270-5457. Any questions were directed to Baxter Global Technical Services at 1-800-553-6898, prompt 4, option 1.
Z-1478-2009 | Jun-26-2009 | 2009 | Philips Avalon Fetal Monitor - Philips Avalon Fetal Monitor FM20; Model Number: M2702A. | Philips Healthcare Inc. | 2 | Interruption of power-up/boot sequence may cause the paper scale and recorder speed settings of the monitor to revert to factory default may lead to misinterpretation of the fetal trace by the user and to potential incorrect clinical decisions | Philips Healthcare issued an "Urgent-Medical Device Correction" notification dated May 15, 2009 to Customers in the United States via UPS and notifications outside the US will be managed by Philips representatives in each affected geography. Customers were asked to follow the guidelines in the Action to be Taken by Customer/User section of the Correction Notice until they receive their software upgrade. Questions can be made to Philips response center at 1-800-722-9377.
Z-0582-2010 | Dec-29-2009 | 2009 | Delta/Delta XL Monitors with Omega S Package - Delta/Delta XL Monitors with Omega S Package | Draeger Medical; Inc. | 2 | monitor is rebooting/resetting | Draeger Medical Systems, Inc. issued and Urgent Medical Device Recall letter dated November 2009 to their affected customers informing them of the problem and a work around until software can be validated. A Technical Service Bulletin has been generated and distributed globally that instructs the affected International Drager subsidiaries to inform customers in their countries. For further information, contact Draeger Medical Systems, Inc. at 1-800-543-5047, extension 1-2349.
Z-2029-2009 | Sep-04-2009 | 2009 | Newport e360 Ventilator - Newport e360 Ventilator; Rx only. The e360 Ventilator System is intended to provide continuous(endotracheal or tracheostomy [ET] tube) or non-continuous(mask) ventilatory support and monitoring for infant; pediatric; and adult patients requiring tidal volumes equal to or greater than 20 milliliters (mL). | Newport Medical Instruments Inc | 2 | In December 2008; Newport initiated a field correction to upgrade e360 ventilators with software versions 3.0 or higher to software version 3.3. As a result of this upgrade; some customers experienced unexpected device alert alarms. Newport's investigation into these complaints confirmed that the 3.3 software responds too quickly to certain situations; giving a false device alert. Newport condu | Newport Medical initally sent out a Field Correction Notice to a limited number consignees dated December 17, 2008 and then the notice was expanded to include all e360 consignees dated June 15, 2009. The firm notified consignees that they have revised the e360 software to version 3.8. A customer verification form was attached to verify upgrade completion. For further questions, contact Newport Customer Technical Support Department at 1-800-451-3111 or +1-714-427-5811 Extension 500.
Z-0285-2009 | Jan-05-2009 | 2009 | GE Healthcare Solar 8000M System - GE Healthcare Solar 8000M System; a multiparameter physiological patient monitoring system intended for use on adult; pediatric and neonatal patients; within a hospital or facility environment | GE Healthcare | 2 | A software error has been discovered when using Combination Monitoring with Solar 8000M/i - PDM - Telemetry in combo mode. This error has only been noted with this particular configuration and mode of monitoring. It has not been observed in other monitoring applications. The software error will result in a reset of the Solar 8000M/i; causing a loss of monitoring for approximately 1-2 minutes a | A GE Healthcare consignee letter was sent out 9/9/08. The letter dated August 26, 2008 was addressed to Healthcare Administrator/Risk Manger. The letter described the problem, product and provided recommendation.
Z-1039-2009 | Apr-06-2009 | 2009 | GE Healthcare S/5 iCentral L-NET05; L-NETC05 - GE Healthcare; S/5 iCentral and iCentral Client;Related to software licenses L-NET05 and L-NETC05; Datex-Ohmeda Inc.; PO Box 7550 Madison; WI 53707-7550; USA; GE Healthcare Finland Oy; Helsinki; Finland PO Box 900. FI-00031 GE.The iCentral maintains the network connections between the Datex-Ohmeda bedside monitors and other network devices in Monitor Network. Network connection consists of hardwire network and/or wireless LAN (WLAN) connection. The iCentral can be used for remote monitor management; storage; printing; viewing; reviewing or otherwise processing of information from several bedside monitors or other network devices. The Dates-Ohmeda Network will be used for patients in the hospital and it is meant for continuous use. The device is to be used by qualified personnel only. | GE Healthcare | 2 | GE Healthcare is providing information about two issues associated with iCentral L-NETO5 and L-NETC05 v4.3; 4.4; 4.5 and 5.0 when used in combination with Dash 2000; 3000; 4000 and 5000; Solar 7000; 8000; 8000M and 8000i; Eagle 3000 and 4000 and ApexPro. The iCentral does not provide audio or visual alarming for the following: A FIB; VT> and PVC(1) alarms when used with ApexPro and VT>2 and PVC | Consignees were sent a GE Healthcare "Urgent Medical Device Correction" letter dated January 26, 2009 addressed to Healthcare Administrator/Risk Manager, Chief of Nursing and Director of Biomedical Engineering. The letter described the safety issue, affected product details, safety instructions, product correction and contact information. Direct questions to GE Healthcare at 1-800-558-7044, Option 2 and ask for Technical Support or contact your local GE Healthcare Distributor.
Z-1939-2009 | Sep-17-2009 | 2009 | Verathon BladderScan BVI 9400 - The BladderScan BVI 9400 is a portable ultrasound instrument that provides a noninvasive measurement of urinary bladder volume. The device consists of an ultrasound probe that scans the patient's bladder; and a battery operated console that provides measurement related information. The BVI 9400 measures ultrasonic reflections on multiple planes inside the body and produces a three dimensional image. Based on the image; the BVI 9400 calculates and displays bladder volume. | Verathon; Inc. | 2 | Three specific issues are being addressed:1. Calculation and use of an incorrect year in dates after 12/31/2009.2. Double scans; continuous scans; and double printing.3. Calibration sensitivity. | U.S. customers (including Government) were sent a post card on or about 5/18/2009 and were called about the upgrade. Units are to be upgraded by the customers using an online tool called ScanPoint. Customers may elect to return their units to the firm for upgrading. Canadian customers were phoned by Verathon. For foreign units sold and distributed through one of the recalling firm's related firms, Service Bulletin SB-0014 was sent to the distributors located in The Netherlands, United Kingdom, and Australia. Contact Verathon Customer Care at 1-800-331-2313 if you have any questions.
Z-1725-2009 | Aug-25-2009 | 2009 | Hologic Discovery - Hologic Discovery series Bone Densitometer System Software versions 12.6; 2.3.; or 2.4.Indicated for the estimation of bone mineral density (BMD); comparison of measurable variables obtained from a given QDR scan to a database of reference values; the estimation of fracture risk; vertebral deformity assessment; body composition analysis; and discrimination of bone from prosthetics. | Hologic; Inc. | 2 | System may not turn off at the end of the x-ray and cause additional x-ray exposure | Hologic issued an Urgent: Medical Device Recall customer notification letter, dated April 17, 2009, by certified mail to inform customers of the problem. With the notification letter, customers will receive the following: -Software version Apex 12.6.2, 2.3.2, or 2.4.2 -Instructions to upgrade their systems with the provided software and to verify the new version on the system as a confirmation that the software was successfully loaded. -New Operator's Manual for installations where the operating characteristics of the system have changed due to the software upgrade. For Customers with systems with software versions other than 12.6, 2.3, or 2.4: Customers were asked to upgrade their QDR software, return their old QDR software disk, and complete and return to Hologic the enclosed Confirmation of Notification.
Z-1809-2009 | Aug-27-2009 | 2009 | Varis; Aria Radiation Oncology - Version 8.1.15; 8.5.11; and 8.6.07 - Varis; Aria Radiation Oncology - Version 8.1.15; 8.5.11; and 8.6.07; Model number HIT-INFSYS; Manufactured by Varian Medical Systems Inc.; Palo Alto; CAThe Varis RV function is designed to assist the operator of a radiation therapy device in providing accurate treatment setups for each patient by monitoring set up parameters and preventing the radiation therapy device from commencing irradiation while any parameter is out of conformance with the treatment plan. Additional applications provide various data management and library functions. | Varian Medical Systems Oncology Systems | 2 | Potential mistreatment-- due to software issues; retired plans may be used and treated repeatedly. | An Urgent Medical Device Correction Letter/ Field Safety Letter, dated May 8, 2009, were distributed to all affected users, with a description of the problem and user corrective action steps. The Letter will also be distributed to the Varian Sales, Marketing, and Service organizations, informing them of the issue. The firm intends to produce a software fix for the problem. If further clarification is needed, users should contact their Varian Customer Support District or Regional Manager.
Z-0942-2009 | Feb-25-2009 | 2009 | Stryker InTouch Critical Care Bed - Stryker InTouch 1.0 Critical Care Bed (Model 2130) with the Integrated PositionPro mattress (Model 2920) option.The device is intended to be used as an Intensive Care Unit bed. | Stryker Medical Div. of Stryker Corporation | 2 | The touch screen on the footboard may display a "tilt over range" error; resulting in a precautionary stop of the bed motions which result in the bed appearing to have locked out some of its bed positioning features. Also; the "hold" button; used to interrupt an intended patient turn on beds with an integrated PositionPro mattress; may only function intermittently; resulting in an inability to ho | Consignees were notified of the problems via letter dated January 5, 2009, advised to check their beds and to take them out of service if the foot board touchscreen displayed a "tilt over range" error, and that a Stryker Medical field service representative would contact them to perform software upgrades on the affected beds. Please direct questions to Stryker Medical at 1-269-324-6644.
Z-0940-2009 | Feb-25-2009 | 2009 | Stryker InTouch Critical Care Bed - Stryker InTouch 1.0 Critical Care Bed; Model 2130.The device is intended to be used as an Intensive Care Unit bed. | Stryker Medical Div. of Stryker Corporation | 2 | The touch screen on the footboard may display a "tilt over range" error; resulting in a precautionary stop of the bed motions which result in the bed appearing to have locked out some of its bed positioning features. | Consignees were notified of the problems via letter dated January 5, 2009, advised to check their beds and to take them out of service if the foot board touchscreen displayed a "tilt over range" error, and that a Stryker Medical field service representative would contact them to perform software upgrades on the affected beds. Please direct questions to Stryker Medical at 1-269-324-6644.
Z-1548-2009 | Jul-17-2009 | 2009 | Gendex VixWin Platinum Imaging Software - Gendex VixWin Platinum Imaging Software; Versions 1.0; 1.1 and 1.2; Gendex Dental Systems; 901 West Oakton Street; Des Plaines; IL 60018.Intended for use of the software is to control capture; display; treatments; analysis and saving of x-ray digital images. | Gendex Dental Systems | 2 | There is a compatibility issue when the VixWin Platinum versions 1.0/1.1/1.2 software is used in conjunction with Easy Dental; PracticeWorks; EagleSoft and SoftDent practice management software; resulting in patient files that may contain images that are mixed between patients. | Gendex issued an "Urgent Medical Device Report" dated March 6, 2009 to the direct account dealers of the affected device via first class mail. The noticed described the issue and provided a software update directed to fixing the problem. In addition, customers were asked to complete and return an Acknowledgement Form by fax to 1-847-718-0601. Further questions should be directed to the Gendex Technical Support team at 1-888-275-5286.
Z-1549-2009 | Jul-01-2009 | 2009 | GE Centricity PACS software - GE Centricity PACS (Picture Archiving and Communication System) software; GE Healthcare Integrated IT Solutions; Barrington; IL 60010.The Centricity PACS is used to receive; store; distribute; display; manipulate; edit and annotate images throughout a clinical environment. | GE Healthcare Integrated IT Solutions | 2 | There is a potential safety issue associated with the use of GE Centricity PACS software related to configuration of DICOM query retrieval with rejected images. | GE Healthcare Integrated IT Solutions sent an "Urgent Medical Device Correction" letter dated April 2, 2009 to their Centricity PACS software versions 2.1.x and 3.0.x customers with attention to the Director of Radiology, the Hospital Administrator, the Head of the Radiology Department and the PACS Administrator. The letters advised users of the patient safety issue associated with the affected device and provided safety instructions to follow until the software is updated. User were asked to complete the enclosed Customer Reply form and fax the completed form to 1-847-939-1479. Centricity PACS Software Versions 1.0.x and 2.0.x customers were sent separate letters explaining the issue and the safety instructions to follow, but there will be no software updates since those software versions are no longer being supported. Any questions should be directed to the Remote Online Center, Field Engineer or the GE Customer Care Center at 1-800-437-1171.
Z-0456-2010 | Nov-24-2009 | 2009 | GE Centricity PACS RA1000 Workstation Software - GE Centricity PACS (Picture Archiving and Communication System) RA1000 Workstation software; GE Healthcare Integrated IT Solutions; Barrington; IL 60010. The Centricity PACS is used to receive; store; distribute; display; manipulate; edit and annotate images throughout a clinical environment. | GE Healthcare Integrated IT Solutions | 2 | There are potential safety issues; including misdiagnosis or delay in treatment; associated with the use of GE Centricity PACS RA1000 Workstation software that may impact patient safety. | GE Healthcare Integrated IT Solutions sent Urgent Medical Device Correction letters dated September 23, 2009 to their Centricity PACS software versions 2.1.x and 3.0.x customers, to the attention of the Hospital Risk Manager, the Hospital Administrator, the Head of the Radiology Department and the PACS Administrator, on 9/25/09. The letters advised the users of the patient safety issues associated with the use of GE Centricity PACS related to unintentionally adding new images to an approved exam; deletion of images from PACS if Cancel Checked or Cancel Selected is chosen in the CA Tool - Archive Queue; inadvertently getting two or more exams out of sync with the corresponding reports resulting in the RIS-IC reports being associated with the wrong patients exam; and when rejecting a duplicate image(s) from an unspecified exam that is still in the Arrived or Verified status, the duplicate image(s) is also rejected from the Dictated, Transcribed, or Completed status exam. The letter provided the users with safety instructions to follow until the software is updated. Any questions were directed to the Remote Online Center, Field Engineer or the GE Customer Care Center at 800-437-1171. The accounts were requested to complete the enclosed customer reply form, supplying the workstation software or build number, and fax the completed form to 847-939-1479. Centricity PACS Software Versions 1.0.x and 2.0.x customers were sent separate letters explaining the issues and the safety instructions to follow, but there will be no software updates since those software versions are no longer being supported. Any questions should be directed to the Remote Online Center, Field Engineer or the GE Customer Care Center at 1-800-437-1171.
Z-1543-2009 | Jul-13-2009 | 2009 | GE Centricity PACS RA1000 Workstation - GE Centricity PACS RA1000 Workstation software (for diagnostic image analysis); GE Healthcare Integrated IT Solutions; Barrington; IL 60010. The Centricity PACS (Picture Archiving and Communication System) Workstation is intended for use as a primary diagnostic and analysis tool for diagnostic images by trained healthcare professionals; including radiologists; physicians; technologists; clinicians and nurses. It is also intended for use as a clinical review workstation throughout the healthcare facility. The workstation interface provides the user with a means to display; manipulate; archive; print; and export images when connected with the Centricity PACS infrastructure. | GE Healthcare Integrated IT Solutions | 2 | Software anomalies may allow the image orientation to be misinterpreted; posing a potential patient safety issue. | GE Healthcare Integrated IT Solutions notified their Centricity PACS RA1000 Workstation Software Versions 2.1.x and 3.0.x customers via an Urgent Medical Device Correction notice dated March 2, 2009 of the affected device resulting in patient safety issues. The letter provided users with safety instructions until the software is updated. Consignees are requested to complete the enclosed customer reply form, supplying the workstation software or build number, and fax the completed form to 1-847-939-1479. Centricity PACS RA1000 Workstation Software Versions PathSpeed 7.x, 8.x, Centricity 1.X and 2.0.X customers were sent separate letters explaining the recall including instructions. There will be no software updates since those software versions are no longer being supported. For questions, call the Remote Online Center, Field Engineer or the GE Customer Care Center at 1-800-437-1171.
Z-1224-2009 | May-29-2009 | 2009 | Terumo APS-1 - Terumo Advanced Perfusion System 1 Central Control Monitor. Terumo Cardiovascular Systems; Ann Arbor; MI; Catalog number 802100.The Central Control Monitor (CCM) provides a centralized display of system information and on-screen control of all system devices. The CCM combines an embedded computer; a touchscreen; and a graphical user interface. | Terumo Cardiovascular Systems Corporation | 2 | The monitor may exhibit start-up failure; video/display issues; calibration issues; and partial or complete loss of functionality during use. | Consignees were notified via an Urgent Medical Device Correction: Safety Advisory letter dated 1/12/09 advised of the reason for the correction, potential problems, and provided clinical information should the problem(s) arise. Customers are to review the notice, assure that all users are aware of the notice, and confirm receipt of the communication by completing the attached Customer Response Form and faxing it to the number indicated on the form. Questions or concerns should be directed to Terumo CVS Customer Service at 1-800-521-2818.
Z-1064-2009 | Apr-10-2009 | 2009 | CELL-Dyn Ruby Hematology Analyzer - CELL-Dyn Ruby Hematology Analyzer; List Number 08H67-01The CELL-DYN Ruby is a multi-parameter; automated hematology analyzer designed for in vitro diagnostic use in clinical laboratories. The instrument has two modes of operation: the Autoloader mode that aspirates samples from closed collection tubes and the Open Tube mode that aspirates samples from open collection tubes. | Abbott Laboratories | 2 | Carryover failures in software revisions 1.0ML and 2.0ML result in elevated PLT background count. | Recall initiated on 11/17/2008. A Product Correction letter and customer reply form is sent to all currently active CELL-DYN Ruby customers. The customers are asked to keep the communication with their CELL-DYN Ruby System Operator's Manual and to note that the information listed in the table replaces the carryover specification provided in all applications of the manual. Questions from customers in the US should be directed to Customer Support at 1-877-4ABBOTT. Customers outside the US should contact their local hematology customer support representative.
Z-1994-2009 | Sep-01-2009 | 2009 | Datex-Ohmeda S/5 Critical Care Monitor - Datex-Ohmeda S/5 Critical Care Monitor with L-CICU01; L-CICU02A; L-ICU05 & L-ICUO5A software. CCMIndicated for monitoring of hemodynamic (including arrhythmia and ST-segment analysis); respiratory; ventilatory; gastrointestinavregional perfusion; B ispectral index (BIS); and neurophysiological status of all hospital patients. | GE Medical Systems Information Tecnology | 2 | GE Healthcare (Datex-Ohmeda) S/5 Critical Care Monitor and S/5 Compact Critical Care Monitor with L-CICU01; L-CICU02A; L-ICU05 & L-ICUO5A software had an issue where the Heart Rate (HR) was displaying as ---(3 dashes) when the monitor was used continuously for 49 days without a cold start. A missed alarm or missed HR value could cause a delay in recognition of a cardiac event that may result i | GE Healthcare Technologies issued a Marketing Bulletin "Ref: MB MS HKI 06007" dated February 6, 2006 to explain a software update.
Z-0969-2009 | Feb-03-2009 | 2009 | GE Centricity PACS RA1000 Workstation - GE Centricity PACS RA1000 Workstation (for diagnostic image analysis); GE Healthcare Integrated IT Solutions; Barrington; IL 60010.The Centricity PACS (Picture Archiving and Communication System) Workstation is intended for use as a primary diagnostic and analysis tool for diagnostic images by trained healthcare professionals; including radiologists; physicians; technologists; clinicians and nurses. It is also intended for use as a clinical review workstation throughout the healthcare facility. The workstation interface provides the user with a means to display; manipulate; archive; print; and export images when connected with the Centricity PACS infrastructure. | GE Healthcare Integrated IT Solutions | 2 | If a user site created a custom default display protocol (DDP) to display the current exam versus the historical exam and the user-defined DDP fails; the system displays the exams using a GE default hanging protocol. This may create a potential for misdiagnosis if the user does not review the date and time stamps in the title bars. | GE Healthcare Integrated IT Solutions notified their Centricity PACS RA1000 Workstation software customers via letter (Urgent Medical Device Correction) dated 1/12/09 of the software anomaly resulting in a patient safety issue involving Default Display Protocols. The issue may create a potential for misdiagnosis if the user does not review the date and time stamps in the title bars. The letter provided the users with safety instructions to follow until the software is updated. Any questions were directed to the Remote Online Center, Field Engineer or the GE Customer Care Center at 800-437-1171. The accounts were requested to complete the enclosed customer reply form, supplying the workstation software or build number, and fax the completed form to 847-939-1479.
Z-1001-2009 | Apr-21-2009 | 2009 | GE QUASAR NUCLEAR MEDICINE SYSTEM - GE Healthcare Infinia; Model Numbers: ASM000886; ASM001391; ASM001333; ASM001465; ASM001473; ASM001620; ASM500016; H3000WM; H3000WT; H3000WY and SYS000041.Product is used to perform general Nuclear Medicine imaging procedures for detection of radioisotope tracer uptake in the patient body; using a variety of scanning modes supported by various acquisition types and optional imaging features designed to enhance image quality in Oncology; Cardiology; Neurology and other clinical diagnostic imaging applications. | GE Medical Systems; LLC | 2 | GE Healthcare has recently became aware of excessive pressure applied by the collimator pressure sensitive cover of Infinia (Infinia; Infinia Hawkeye?; Infinia Hawkeye? 4) that may impact patient safety. It was reported at a customer site that the pressure used to activate the Pressure Sensitive Device (PSD) should be lowered and that the body contouring software requires correction to prevent po | Consignees were sent from June 2008 to February 2009 a GE "Product Safety Notification" letter dated April 2008. The letter was addressed to Hospital Administrators and Managers of Radiology/Cardiology, Radiologists/Cardiologists. The letter described the Safety Issue, Affected Product, Safety Instructions, Product Correction and Contact Information. A GE Healthcare service representative will contact each consignee to schedule the installation of the Pressure Sensitive Device (PSD). Please contact your local sales or service representative for further information regarding this recall. Further information can be found by calling 1-800-437-1171 in the US or Canada or 1-262-896-2890 if outside the US.
Z-0870-2009 | Jan-12-2009 | 2009 | Centricity Web - Centricity Enterprise Web 3.0 software; the software allows physicians convenient; quick access to medical images and related data they need via a web browser; password and a standard PC; GE Healthcare Integrated IT Solutions; Barrington; IL 60010. | GE Healthcare Integrated IT Solutions | 2 | Software error: There are two potential safety situations with Centricity Enterprise Web software - a forced log off may occur while using the Centricity Enterprise Web during an open session; and Centricity Enterprise Web does not utilize calibration performed on image measurements in Centricity RA1000 Workstation. | GE Healthcare Integrated IT Solutions notified their Centricity Enterprise Web software customers via letter (Urgent Medical Device Correction) dated 12/12/08 of the software anomalies resulting in patient safety issues involving a forced log off during an open session while using the Centricity Enterprise Web and measurements displayed in Centricity Enterprise Web will use the calibration provided by the original scanning modality. Centricity Enterprise Web does not utilize calibration performed on image measurements in Centricity RA1000 Workstation. The letter provided the users with safety instructions to follow until the software is updated. Any questions were directed to the Remote Online Center, Field Engineer or the GE Customer Care Center at 800-437-1171. The accounts were requested to complete the enclosed customer reply form, supplying the workstation software or build number, and fax the completed form to 414-918-4653. For further information, please contact GE Customer Care Center by telephone at 800-437-1171.
Z-0541-2010 | Dec-16-2009 | 2009 | GE Healthcare Voluson E6; GE Voluson E8; and GE Voluson E8 Expert Diagnostic Ultrasound Systems - GE Healthcare Voluson E6; GE Voluson E8; and GE Voluson E8 Expert Diagnostic Ultrasound Systems | GE Medical Systems; LLC | 2 | GE Healthcare has recently become aware of an incorrect operating procedure associated with software versions 9.0.0; 9.0.1; and 9.0.2 of the Voluson E6 or Voluson E8 Diagnostic Ultrasound Systems that may impact patient safety.When using the combination of 2D-CRI and PW with a steered Doppler angle; the displayed velocity scale is incorrect. The consequence can be an underestimation of flow v | Consignees were sent on 10/23/09 a GE healthcare "Urgent Medical Device Correction" letter dated 10/22/09. The letter was addressed to Hospital Administration/Risk manger, Primary Physician and Biomedical Dept. Manger. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information. Questions and concerns regarding this issue are directed to the local GE Healthcare service representative or to 1-800-437-1171. GE also plans to update all of the affected GE Voluson E6, GE Voluson E8, and GE Voluson E8 Expert Diagnostic Ultrasound Systems that were distributed.
Z-1090-2009 | Mar-19-2009 | 2009 | BD PrepStain System - BD PrepStain System; Model Number: 799-14000-00; Catalog Number: 05CR000021; 05CR00021R; 799-13000-00; 799-13000-00R; 799-14000-00 and 799-14000-00R; Catalog Number: 05CR000021; 05CR000021R; 799-13000-00; 799-13000-00R; 799-14000-00 and 7914000-00R.PrepStain is a liquid-based thin layer cell preparation device for preparation and staining of Prep smear slides from cervical specimens in a liquid based format (gyn use) and of cytology examination slides from non-cervical specimens such as urine and lung lavage. | Bd Diagnostic Systems Tripath | 2 | Software Error in the non-gyn program of some devices may cause non-gyn samples to be incorrectly transferred from centrifuge tubes to slides potentially result in mismatched patient results. | BD Diagnostics began contacting known non-gyn customers by phone on 09/26/2008. Customers were notified by Urgent - Important Product Information letter on 09/29/2008. They were advised to stop using the system until they complete a list of actions provided in the letter. A telephone contact number for Technical Support was provided in the letter. A permanent software based solution will be provided to customers when completed. They anticipate the software will be available in October 2008.
Z-0129-2009 | Jan-13-2009 | 2009 | Cytomics FC 500 Flow Cytometry System with CXP Software - Cytomics FC 500 Flow Cytometry System with CXP Software Version 2.2; Part Number: 626553. This product is used for the qualitative and quantitative measurement of biological and physical properties of cells and other particles. These properties are measured when the cells pass through one or two laser beams in single-file. | Beckman Coulter Inc | 2 | Software error: This field action was initiated after Beckman Coulter determined that the CXP software does not refresh the dataset automatically under certain circumstances. The User documentation does not clearly state that the user should refresh the dataset after changes are made. | Software error: A Product corrective Action (PCA) letter was mailed via US Postal Service on October 17, 2007 to all Cytomics FC 500 Flow Cytometry System with CXP Software accounts. This letter was entitled "URGENT: PRODUCT CORRECTIVE ACTION." Customers were informed that the user documentation does not clearly state that the user should refresh the dataset after changes are made. The letters provide the customers with an explanation and action to be taken whenever a protocol containing data is modified by making changes to cytosettings, regions or gates. Customer were instructed to review all data plots and results before reporting results out. any uncharacteristic patterns in the data or unexpected results must be verified by running Quality Control and if necessary, preparing new samples. In addition, customer were instructed to please share this information with their laboratory staff and retain this notification as part of their laboratory Quality System documentation. Also they were asked to complete and return the an attached Customer Response Form within ten days. Customers with questions or concerns regarding the notification, could call Beckman Coulter Customer Service at 1-800-526-7694 in the United States or Canada. Outside the US and Canada, they were to contact their local Beckman Coulter Representative.
Z-0455-2009 | Jan-07-2009 | 2009 | FC 500 with CXP Acquisition Software - Cytomics FC 500 Flow Cytometry System with CXP Software; Part Numbers: 626553; Version 2.2 | Beckman Coulter Inc | 2 | Software error: The recall was initiated after Beckman Coulter identified that the CXP User documentation does not adequately characterize the use of the Live Gate feature.The primary use of a Live Gate is for the international removal of unwanted events; such as debris. The data outside of the Live Gate will not be stored to the final listmode file; this reducing the data set size.The softwa | The customer notifications were initiated by Beckman Coulter on July 23, 2007 when a Product Corrective Action (PCA) was mailed via US Postal Service to all affected customers. The PCA informed the customers that the user documentation does not adequately characterize the use of the Live Gate feature. The letter provides the customer with explanation of Live Gate and actions to be taken. For further information, please contact Beckman Coulter, Inc. by telephone at 714-993-5321.
Z-1155-2009 | Sep-24-2009 | 2009 | Fluorostar 7900 Mobile Fluoroscopy System - GE Healthcare Fluorostar 7900 Mobile Fluoroscopy System | Ge Healthcare | 2 | The alarm timer is reset after each exposure and therefore never reaches the five minute alarm state unless one exposure exceeds 5 minutes in duration. | GE Healthcare sent a customer notification out to all sites that have potentially affected systems. GEHC Surgery will also install a software correction into the affected systems in order to resolve this issue at no charge to the customer. Until the GEHC OEC Fluorostar 7900 has received the corrective software installed, they are encouraging facilities to exercise caution in monitoring the amount of cumulative fluoroscopy time when the affected system is in use.
Z-1475-2009 | Jun-19-2009 | 2009 | Digital Fluoroscopic Imaging System - GE Innova 3100 / 3100 IQ Cardiovascular Imaging System; GE Healthcare; Waukesha; WI 53188.The product is indicated for use in generating fluoroscopic images of human anatomy for vascular angiography diagnostic and interventional procedures and optionally; rotational imaging procedures. It is also intended for generating fluoroscopic images of human anatomy for cardiology; diagnostic and interventional procedures. | GE Medical Systems; LLC | 2 | GE Healthcare has become aware that a few customers have experienced an unusual occurrence of system lockups on Innova single plane cardiovascular X-ray imaging systems while using the image auto-archiving function. when occurring during X-Ray imaging; the acquisition fails and the user has to rest the system. To date; one patient death has been reported related to this issue. | GE Healthcare sent a "GE Urgent Medical Device Correction" letter dated May 2008. The letter was addressed to Health Care Professionals describing the problem with the affected devices including instructions to avoid further problems until a GE Healthcare service representative is able to make contact and upgrade the system. Further questions may be addressed by calling 1-800-437-1171.
Z-0861-2009 | Feb-20-2009 | 2009 | Eclipse treatment planning system - Varian Medical Systems Eclipse treatment planning system; Model Number H48; versions 6.5/7.x and Inspiration 8.0. | Varian Medical Systems Oncology Systems | 2 | Possible incorrect wedge orientation in treatment: Software labeling was issued without warning message relating to preservation of user selected values for Siemens wedge orientation; wedge limits and SAD values for Elekta machines; and collimator operating limits for Saturne Machines when the software version is upgraded. | An Urgent Device Correction product notification letter/field safetly letter dated November 21, 2008, was distributed to all affected users, with a description of the problem and user corrective action steps. The letter was also distributed to Varian Sales, Marketing and Service organizations, informing them of the issue.
Z-1233-2009 | May-07-2009 | 2009 | Siemens Mammomat Novation DR - Siemens Mammomat Novation DR; full field digital mammography system. Model Number 06646900.The device is used for full field digital mammography system. | Siemens Medical Solutions USA; Inc | 2 | Image may appear dark; requiring repeat image acquisition. | A Customer Safety Advisory Notice dated December 16, 2008 was issued to affected customers via Update Instructions SP040/08/S. The letter described the potential issue and provided instructions to avoid its occurrence. A software update to correct the issue is currently being developed and will be installed on affected systems when it becomes available.
Z-0259-2010 | Nov-13-2009 | 2009 | Integrated Power Console System - Medtronic Integrated Power Console System; IP X1REF 1898001 or REF EC300Medtronic Xomed; Inc.; 6743 Southpoint Drive North; Jacksonville; FL 32216-0980 USA.Distributed by: Medtronic Powered Surgical Solutions; 4620 North Beach Street; Fort Worth; TX 76137 USA.No difference between two referenced Consoles (REF numbers differ for internal sales tracking purposes). | Medtronic Xomed; Inc. | 2 | An error code (Code 13) is displayed on some of the consoles when a Medtronic Stylus or EHS handpiece is connected to the IPC indicating that the handpiece has not been recognized by the console. | A Product Correction Notification was issued to each consignee on June 15, 2009. Medtronic sales representatives will visit each facility to advise consignees of the issue and to install a software upgrade that will correct the handpiece recognition issue on the console(s).
Z-0928-2009 | Jan-27-2009 | 2009 | GE Centricity PACS RA1000 Workstation - GE Centricity PACS RA1000 Workstation (for diagnostic image analysis). GE Healthcare Integrated IT Solutions; Barrington; IL 60010.The Centricity PACS Workstation is intended for use as a primary diagnostic and analysis tool for diagnostic images by trained healthcare professionals; including radiologists; physicians; technologists; clinicians and nurses. It is also intended for use as a clinical review workstation throughout the healthcare facility. The workstation interface provides the user with a means to display; manipulate; archive; print; and export images when connected with the Centricity PACS infrastructure. | GE Healthcare Integrated IT Solutions | 2 | There is a potential patient safety issue involving patient jacket content; which may result in unintended historical image(s) being displayed. | GE Healthcare Integrated IT Solutions notified their Centricity PACS RA1000 Workstation software customers via letter (Urgent Medical Device Correction) dated 11/21/08 of the software anomaly resulting in a patient safety issue involving patient jacket content, which may result in unintended historical image(s) being displayed. The letter provided the users with safety instructions to follow until the software is updated. Please direct any questions to the Remote Online Center, Field Engineer or the GE Customer Care Center at 800-437-1171. The accounts were requested to complete the enclosed customer reply form, supplying the workstation software or build number, and fax the completed form to 414-918-4653.
Z-0473-2009 | Jan-21-2009 | 2009 | Innova 2100IQ - GE Innova 2100IQ Versatile Cardiac and Vascular Imaging System; GE Healthcare; Waukesha; WI 53188 ( The system consists of an a monoplane positioner; a vascular or cardiac table; an X-ray system and a digital detector) The product is indicated for use in generating fluoroscopic images of human anatomy for vascular angiography diagnostic and interventional procedures and optionally; rotational imaging procedures. It is also intended for generating fluoroscopic images of human anatomy for cardiology; diagnostic and interventional procedures. | GE Medical Systems; LLC | 2 | Software error; computer: GE Healthcare has became aware of an issue with Error Management affecting the Innova 2100; 3100/3100 IQ & 4100/4100IQ. It was reported that when the foot switch was pressed; the live monitor screen was blank and it looked like there were no X-rays being emitted. But the message "Acquisition in progress" was displayed on the DL monitor; even after the foot switch was r | Consignees sent a "GE Urgent Medical Device Correction" letter dated February 28, 2008. The letter was addressed to Hospital Administrators/Hospital Risk Management Department/Managers of Radiology-Cardiology/ Radiologists-Cardiologists. The letter described the affected products, safety Issue, short term recommendations, long term solution and contact Information. The first batch of letters was received by customers on 5/28/08. For further information, please contact GE Healthcare at 262-544-3894.
Z-0665-2009 | Jan-28-2009 | 2009 | Medtronic Maximo II DR - Medtronic Maximo II DR D28DRG; for ventricular pacing and defibrillation. | Medtronic Inc. Cardiac Rhythm Managment | 2 | Maximo II VR ICD; DR ICD and CRT-D inadvertently include the Capture Management algorithm (Atrial/Right Ventricular/Left Ventricular Capture management). The Capture management algorithm monitors daily pacing thresholds and may increase or decrease pacing amplitudes in response to patients' needs. Although the algorithm is operating correctly; it is not accessible to the clinician for reprogramm | A Medtronic Inc. letter dated July 30, 2008 was sent to Physicians. The letter described the product and problem. The letter advised to contact Technical Services for specific questions regarding the care of patients or how to program the Maximo II devices. The letter instructed to collect and hold all non-implanted Maximo II devices and that a Medtronic Representative will contact the doctor to make arrangements to update the programmer 2090.
Z-1568-2009 | Jul-06-2009 | 2009 | MX8000 IDT CT SYSTEM - Dental Application on the Extended Brilliance Workspace v4.0 and v4.0.1. Model #455011202691 for Computed Tomography (CT); Model #455011002031 for Nuclear Medicine (NM).The "Extended Brilliance Workspace (EBW)" Dental Application v4.0 and v4.0.1 is a component of the Computed Tomography X-Ray System that is intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. | Philips Medical Systems (Cleveland) Inc | 2 | Philips Healthcare has decided to recall the Dental Application on the Extended Brilliance Workspace v4.0 and v4.0.1 due to the fact that when images are saved or filmed that contain the Curved Multi-planar Reformation Section (cMPRS) views that were created with a short curve the Right/Left annotations will be reversed. | An "URGENT - Field Safety Notice" dated April 27, 2009 was sent to all consignees via certified mail. The letter described the affected products, issue, and actions for the customer/user. The notification letter provides a work around until the new software is available. For further information or support concerning this issue, please contact your local Philips representative: Customer Care Center (1-800-722-9377, option 5: Diagnostic Imaging, option 5: Radiation Therapy, option 3: Oncology) or to your local Philips Healthcare office.
Z-0065-2010 | Oct-19-2009 | 2009 | CELL-DYN Emerald Hematology Analyzer - CELL-DYN Emerald Hematology Analyzer; model number 09H39-01. Distributed by Abbott Diagnostics Division; Santa Clara; CA.Automated analyzer intended for in-vitro diagnostic use in the clinical laboratory. | Abbott Laboratories | 2 | Software bug. The CELL-DYN Emerald WBC L1 flag is configured incorrectly. The flag is not generated according to system requirements. | Abbott Diagnostics issued a "Product Correction" notice dated July 16, 2009 describing the affected product. The notice was sent to all currently active users and included a customer reply form, a software update and installation instructions. For further information, contact Abbott Customer Support at 1-877-4ABBOTT. Customers outside the U.S. should contact your local hematology customer support representative.
Z-0647-2009 | Jan-08-2009 | 2009 | syngo Dual Energy Viewer - CT syngo Dual Energy Viewer on the syngo Multimodality Work Place (MMWP) and on the CT workplace; Model number 10094833.Picture Archiving and Communications System. | Siemens Medical Solutions USA; Inc | 2 | Indicated orientation does not match actual orientation of the patient. | The recalling firm issued a Customer Safety Advisory dated 10/8/08 to affected customers via Update Instructions CT052/08/S. The letter informs customers of the potential issue and provides instructions to avoid its occurrence. Customers should consider "no saving of calculated Dual Energy data sets from non-supine and mirror-option scans via "Save Sets CT"-button." The firm also highly recommends for customers "to check the correct display of "left", "anterior" and "posterior"markings, e.g. by checking the position of the spine with respect to the posterior marking in the image." The error should be fixed within certain upgrade and installation activities on the affected software.
Z-1542-2009 | Jul-13-2009 | 2009 | Tumor Localization (Tumor LOC) software application - TumorLoc software for the Brilliance Big Bore (BB) and the CT Extended Brilliance Workstation (EBW); Model #4550 116 02431 (TumorLoc).The BB is a CT X-ray system intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. This device may include signal analysis and display equipment; patient and equipment supports; components and accessories. | Philips Medical Systems (Cleveland) Inc | 2 | A software error in the TumorLoc software was discovered upon receipt of a customer complaint. There is the potential for data to be lost during reconstruction of images in which slices can either be discarded; or interleaved based upon their series ID. | Philips Medical issued an 'URGENT -Medical Device Correction' document dated January 28, 2009. This notification informs consignees of the TumorLoc imaging issue, the inherent hazards involved should this problem occur, the circumstances under which this problem may occur and corrective actions which the user/customer should perform in the event that this problem is encountered. It also informs that user/customer that Philips will have their Service Engineers contact the customers within 6 months to install a TumorLoc upgrade. For further questions, contact the Phillips Customer Care Center at 1-800-722-9377.
Z-1719-2009 | Aug-14-2009 | 2009 | GE Healthcare MUSE Cardiology Information Systems - GE Healthcare MUSE Cardiology Information Systems vs. 7.0; 7.0.1; 7.0.2; 7.1; 7.1.1.The MUSE Cardiology Information System is intended to store; access and manage cardiovascular information on adult and pediatric patients. | GE Healthcare | 2 | GE Healthcare has become aware of an issue resulting from a workflow situation related to the automated labeling of patient data in the MUSE Cardiology Information System. When a prior record has been confirmed with an invalid (default) patient identifier (PID); MUSE will accept future records with the same default PID as correct. If subsequent records don not contain a name; MUSE will pre-popu | GE Healthcare issued an "Urgent Medical Device Correction" letter dated October 13, 2008, addressed to Healthcare Administrator/Risk Manager, Cardiology Administrators and Head of Cardiology. The letter listed the Safety Issue, Affected Product Details, Safety Instructions, Product Correction, Background Information and Contact Information. For further questions, contact GE Healthcare at 1-800-558-7044. Outside the US, please contact your local GE Healthcare Service Representative.
Z-1398-2009 | May-22-2009 | 2009 | Cell-Dyn Ruby Hematology Analyzer - Cell-Dyn Ruby Hematology Analyzer with System Software Version 2.0ML; List Number 08H67-01; Manufactured by Abbott Diagnostics Division; Santa Clara; CA.Product is a multi-parameter automated hematology analyzer designed for in vitro diagnostic use in clinical laboratories. The instrument has two modes of operation: autoloader that aspirates samples from closed collection tubes; and open tube mode that aspirates samples from open collection tubes | Abbott Laboratories | 2 | Burn Risk-- when power is cycled to the instrument; there is a possibility that the Mixer Motor Assembly may reach 100 degrees Celsius (212 degrees Fahrenheit ). If the customer touches the motor; burns may occur. | A "Product Correction" letter dated February 12, 2009 and "Customer Reply" form were sent to all currently active CELL-DYN Ruby customers using Software Version 2.0 ML. The letter described the affected product, necessary actions and contact information for customers. Customers were instructed to acknowledge receipt of the "Product Correction" letter by completing the enclosed Customer Reply form and return via fax to 1-800-777-0051. A product information letter is being added to the accessory kits shipped with new instruments and version 2.0 ML System Software upgrade kits. For questions regarding this information, contact Abbott Laboratories Customer Support by calling 1-877-422-2688. Customers outside the US, please contact your local hematology customer support representative.
Z-0150-2010 | Nov-10-2009 | 2009 | ABL 80 CO-OX SYSTEM - ABL 80 CO-OX SYSTEM; software versions 1.30 and 1.31Additional Product Codes: CEM; CGA; CGZ; GHS; GKR; GLY; JFP; JGS; JIX; JJY; JPIThe ABL 80 CO-OX SYSTEM is a portable; automated analyzer that measures pH; blood gases; electrolytes; glucose and oximetry in whole blood. The ABL 80 CO-OX SYSTEM is intended for use by trained technologists; nurses; physicians and therapists. It is intended for use in a laboratory environment; near patient or point-of-care setting. | Sendx Medical Inc | 2 | Inverness have recently become aware of a potential error contained in V1.30 and V1.31 software for the ABL 80 CO-OX platform. The error can occur if the analyzer is set to automatically upload patient records to Radiance or an HIS/LIS system. This can result in the reporting of incorrect results for one or more CO-OX derived parameters. Incorrect results can be displayed on the screen and the | Radiometer sent out a letter to consignees informing them of public reason for recall, and telling them that they have provided a copy of the upgrade software and instructions to correct this potential problem immediately. Software upgrade instructions were included with the recall letter. Consignees were informed that if they preferred that their local service representative perform the upgrade, to follow the Instructions for disabling auto upload of patients records. SOFTWARE UPGRADE INSTRUCTIONS: The upgrade software must be installed via the Software Upgrade option under Menu>Utilities>Software Upgrade. It is not a "boo table" CD. ~The Upgrade process will take approximately 10 minutes. The analyzer should be allowed to complete the entire process without user interaction. Do not attempt to turn the analyzer on or off manually at any time during the process. ~The upgrade process will include the following general steps: 1. Automatic power down and reboot of the analyzer 2. File transfer process which includes a variety of progress bars to communicate the status 3. Automatic power down and reboot of the analyzer a second time 4. A dialog box on the main menu screen that states, "ABL80 FLEX software has been upgraded to version 1.32." This signifies a successful upgrade process. ~If a solution pack and a sensor cassette are already installed on the analyzer, their installation status will not be affected by the software upgrade. ~All stored data will be retained in all logs They were also asked to complete and return the attached fax form, with their signature, to 1-800-736-0601. If their were any questions regarding the letter, or upgrade, to contact Techline at 1-800-736-0600 opt.
Z-1401-2009 | May-27-2009 | 2009 | Panorama Patient Monitoring Network - Panorama Patient Monitoring Network with ambulatory telepacks; and software version 8.5.5; 8.5.6; 8.6.1 or 8.7.Datascope Patient MonitoringIntended for use in a fixed location; in the healthcare facility setting; as a central viewing station. | Mindray DS USA; Inc.; dba Datascope Patient Monitoring | 2 | Software Anomaly: Software anomaly can occur when the Panorama is attempting to play an audio file; either a physiological or technical alarm sound; at which time the system software may cause the system to freeze due to an interaction in the alarm system. | Urgent Product Field Correction letters, dated March 27, 2009, were sent by certified mail, return requested, and signature required. The letters identified the affected product and the problem. It also stated the adverse effect on patients and that customers' Panorama software will be upgraded to a current software revision. A Datascope Service Representative will contact customers to arrange for the software upgrade of the Panorama(s) in their facility. This will be done at no cost to the customer.
Z-1928-2009 | Aug-27-2009 | 2009 | AMPLILINK Software; version 3.1.0 to 3.2.2 - AMPLILINK Software; version 3.1.0 to 3.2.2 with COBAS AmpliPrep. Catalog number 03563383001; catalog number 04855094001; catalog number 04807197001; catalog number 04862392001; catalog number 05093236001; and catalog number 05201357001.Automated sample preparation for nucleic acid analysis on either the COBAS TaqMan or COBAS TaqMan 48. COBAS AmpliPrep/COBAS TaqMan is capable of tunning HBV; HCV; and HIV tests. | Roche Molecular Systems; Inc. | 2 | Software bug: A unique series of events involving the COBAS AmpliPrep instrument running AMPLILINK software version 3.2.2 that results in specimens being amplified and detected with the wrong parameters. The error can occur on either COBAS AmpliPrep/COBAS TaqMan 48; COBAS AmpliPrep/COBAS TaqMan (docked) or cobas s 201 system (docked) running Amplilink Software versions 3.1 or 3.2 series up to and | Urgent Medical Device Correction Letters were sent on June 2, 2009 by first class mail. The letter described the issue; impact on results; clinical implications; the workaround; and the actions required, which include following the workaround until a new version of AMPLILINK software is available and filing the letter for future reference. Questions or concerns should be directed to Roche Diagnostics Technical Support at 1-800-526-1247.
Z-2243-2009 | Sep-21-2009 | 2009 | HI-ART SYSTEM - TomoTherapy Hi-Art System¿; Version 3.XThomoTherapy Incorporated 1240 Deming Way; Madison; WI 53717.The TomoTherapy HI-ART System¿¿ is intended to be used as an integrated system for the planning and delivery of intensity modulated radiation therapy (IMRT). The HI-ART System provides precise delivery of radiation to tumors or other targeted tissues while minimizing the delivery of radiation to vital healthy tissue. | TomoTherapy Incorporated | 2 | TomoTherapy received one customer report indicating that after starting a patient procedure from the StatusConsole; the radiation on light and the audible indicator on the Status Console came on as expected;however the Operator Station (OS) did not display the dose rate and the procedure time; or the couch position updates. Within a few minutes; the customer pressed the Stop button on the Sta | Consignees were sent a TomoTherapy "Urgent Medical Device Correction Field Safety Notice" letter dated August 10, 2009. The letter addressed the Issue, Product Affected, Investigation, Recommended Actions and Resolution. Customers were recommended to perform the following actions (in the event that the dose rate, procedure timer, and couch position updates are not present on the OS, after pressing Start on the status console and the radiation on light and audible indicators are active): 1) Immediately press the Stop button on the Status Console and 2) contact your authorized TomoTherapy Service Representative for further direction. The affected issue will be corrected in the next software version, which is scheduled for release in early 2010. The TomoTherapy Customer Contact Center can be reached 24 hours a day, 7 days a week, by e-mail at support@tomotherapy.com or by phone at 1-866-368-4807 (North America).
Z-1883-2009 | Sep-06-2009 | 2009 | Pyxis Anesthesia System 3500 - Pyxis Anesthesia System 3500; Cardinal Health; Clinical Technologies and Services; San Diego; CA.The Pyxis Anesthesia Station 3500 is a medication dispensing product that provides medication and inventory management for the anesthesiologist at the point of care within the operating room. The system tracks and stores controlled substances; medications and supplies for anesthesia providers and automates verification of patient usage records; documents waste; manages inventory levels and tracks utilization. | Cardinal Health Manufacturing LLC | 2 | Cardinal Health has sent out a Safety Alert Notification to inform its customers that the SR2b Software for the Pyxis Anesthesia System 3500 (PAS 3500) has received several customer reports of system "non-responsiveness" (lock ups) during the use of the product since its release in September 2008. | The "Safety Alert Notification" letter was mailed to all Pyxis Anesthesia System 3500 customers on December 17, 2008 by registered return receipt mail. The letter was addressed to the Director of Pharmacy at each facility. The mailing packages contain the "Safety Alert Notification" letter and the revised pages of the Pyxis Anesthesia System 3500 User Guide and Safety Manual. The "Safety Alert Notification" letter reinforced to customers how to maintain back-up processes to allow quick access if a system lock-up condition is experienced. The letter informed customers the situation will be corrected and a solution will be implemented. If you experience any system unresponsiveness during the use of the product, please report the issue to the Cardinal Health Technical Support Center at 1-800-727-4102. If you have any questions regarding this issue, please contact your local field support representative or the Customer Advocacy Department at Customer.Advocacy@cardinalhealth.com.
Z-2048-2009 | Sep-03-2009 | 2009 | Carl Zeisss Meditec; Inc. Stratus OCT - Stratus OCT Model 3000 instruments distributed with version 6.0.0 software; product number 2660021119710 (new instrument) or 2660022119710 (refurbished instrument); and Stratus 6.0.0 Software Kit; product number 2660021133164.The Stratus OCT is a high resolution tomographic device for the viewing and axial cross sectional imaging of posterior ocular structures. It is used for in vivo imaging and measurement of the retina; retinal nerve fiber layer; macula; and optic disk. The Stratus OCT with Retinal Nerve Fiber Layer and Macula Normative Database is a quantitative tool for the comparison of retinal nerve fiber layer in the human retina to a database of known normal subjects. It is intended for use as a diagnostic device to aid in the detection and management of ocular diseases; including; but not limited to; macular edema; central serous retinopathy; and glaucoma. | Carl Zeiss Meditec Inc | 2 | Mis-labeled image. A bug in version 6.0.0 software and 6.0.0 software kit used with Stratus OCT Model 3000 can produce a report that incorrectly labels OD or OS. | Firm initiated notification to customers on July 2, 2009, by sending a letter via FedEx to US customers, and via e-mail and FedEx to outside US consignees. Customers were informed of the issue and steps needed to correct it. A revised software version will be provided to the customers. Any questions should be directed to Carl Zeiss Meditec, Inc. Customer Care at 1-800-341-6968.
Z-2251-2009 | Sep-20-2009 | 2009 | BrainSCAN and iPlan RT Dose - BrainLab Radiotherapy Treatment Planning Software; Catalog number 20610 - Radiosurgery 3.0Catalog number 20620 - Radiosurgery 3.5 andCatalog number 20630 - Circular ARC SRS/SRT Planning.The software is intended for use in stereotactic; conformal; computer planned; LINAC based radiation treatment of cranial; head and neck; and extracranial lesions. It is intended to be used by experienced and trained health professionals. | Brainlab AG | 2 | Failure to conduct important safety checks when using the BrainLab radiotherapy treatment planning software in combination with BrainLab conical collimators could result in unintended radiation outside the conical shaped field; which may lead to serious injury of the patient. | BrainLAB AG issued an "Urgent Field Safety Notice" letter dated August 12, 2009 informing all customers of the BrainLab radiotherapy planning software using conical collimators of the important safety checks to avoid the failure to apply and verify an adequate jaw-shaped field size and field position at the LINAC for the used conical collimator(s) or the failure to mount and verify the correct, intended conical collimator. Consignees were instructed on proper steps to take to avoid further issues. For further questions, contact BrainLAB AG by phone at 1-800-597-5911. *** On February 1, 2010, BrainLab sent supplemental letters to 50 of the accounts who received the August 12, 2009 letter and have Novalis Dose Delivery Systems (excluding all Novalis Tx systems). The accounts were informed that in addition to the safety checks listed in the August letter, BrainLab will be installing a mandatory software modification on their systems that limits the field size to a maximum of 4x4 cm2 when the conical collimators are inserted. The accounts would be contacted by a service representative within the next six weeks to schedule the update. Any questions were directed to BrainLab at 1-800-597-5911. ***.
Z-2239-2009 | Sep-22-2009 | 2009 | ILUMAVision - ILUMAVision Software; version 2.1.0 Picture Archiving and Communication systemProduct used as a tool for the pre-surgical planning of the placement of dental implants and pre-surgical fabrication of stents to aid in dental implant insertion. In addition; it is used for diagnostics and treatment planning in orthodontic; ENT and similar applications. | IMTEC IMAGING L.L.C. | 2 | Under certain circumstances; software Version 2.1.0 has potential for incorrect data imaging and patient identification. | Firm sent notifications to all ILUMAVision customers on 8/3/2009 and offered upgraded software. The notification letter indicated the the reason for the recall and identified the potential issues associated with the device. Contact 3M IMTEC Technical Support at 1.800.879.9799, if there are any questions.
Z-2316-2009 | Sep-24-2009 | 2009 | Centricity Enterprise Archive - Centricity Enterprise Archive (EA) versions 3.0 to 3.0.7 software; GE Healthcare Integrated IT Solutions; Barrington; IL 60010.The software delivers a scalable and flexible DICOM storage solution for Cardiology and Radiology images. | GE Healthcare Integrated IT Solutions | 2 | Software error: There is potential safety issue with Centricity Enterprise Archive (EA) 3.0.x software where study split operations are not correctly replicated to a secondary 'shadow' archive. | GE Healthcare Integrated IT Solutions sent Urgent Medical Device Correction letters dated August 7, 2009 to the attention to the Director of Radiology, the Hospital Administrator, the Head of the Radiology Department and the PACS Administrator. The letters advised users of the potential safety issue and provided safety instructions to follow until the software is updated. Users were asked to complete the enclosed Customer Reply form and fax the completed form to 1-847-939-1479. Any questions should be directed to the Remote Online Center or the GE Customer Care Center at 1-800-437-1171.
Z-0320-2009 | Jan-05-2009 | 2009 | Somatom Sensation 40 X-Ray System - Somatom Sensation 40 X-Ray System; Model number 8874427. This is a computed tomography x-ray system which is intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data from the same axial plane taken at different angles. | Siemens Medical Solutions USA; Inc | 2 | software error - does not allow protocol to properly synchronize with ECG triggers | The recalling firm issued a Customer Safety Advisory Notice dated 9/3/08 to affected customers via Update Instructions CT045/08/S. This letter informs customers of the potential issued and provides instruction to avoid its occurrence. For additional information, contact Siemens Medical Solutions USA, Inc. at 1-610-448-4500.
Z-0836-2009 | Jan-23-2009 | 2009 | GE Centricity PACS RA1000 Workstation - GE Centricity PACS RA1000 Workstation (for diagnostic image analysis); GE Healthcare Integrated IT Solutions; Barrington; IL 60010.The Centricity PACS Workstation is intended for use as a primary diagnostic and analysis tool for diagnostic images by trained healthcare professionals. It is also intended for use as a clinical review workstation throughout the healthcare facility. The workstation interface provides the user with a means to display; manipulate; archive; print; and export images when connected with the Centricity PACS infrastructure. | GE Healthcare Integrated IT Solutions | 2 | There is a potential patient safety issue involving the use of the F12 keyboard Approve & Complete function. The issue may cause erroneous information to be added to a report when trying to approve and complete it. | GE Healthcare Integrated IT Solutions notified their Centricity PACS RA1000 Workstation software customers via letter (Urgent Medical Device Correction) dated 12/12/08 of the software anomaly resulting in a patient safety issue involving the use of the F12 keyboard Approve & Complete function. The issue may cause erroneous information to be added to a report when trying to approve and complete it. The letter provided the users with safety instructions to follow until the software is updated. Any questions were directed to the Remote Online Center, Field Engineer or the GE Customer Care Center at 800-437-1171. The accounts were requested to complete the enclosed customer reply form, supplying the workstation software or build number, and fax the completed form to 414-918-4653.
Z-0002-2009 | Sep-24-2009 | 2009 | Revolution XR/d - Revolution XR/d system control room PC model 5117866-2 with Software collector 5115434 | GE Medical Systems; LLC | 2 | GE Healthcare has recently become aware of a potential risk of observing a blank image after exposure onRevolution XR/d systems equipped with HP 8200 PC user consoles that may impact patient safety.It was reported that images acquired with off-center collimation for extremities with theposition selected as Digits to front or Digits to back are displayed as blank. Investigation intoth | Consignees were sent a GE Healthcare "Product Safety Notification" letter dated March 7, 2008. The letter was addressed to Hospital Administrators, Mangers of Radiology and Radiologists. The letter listed the Safety Issue, Affected Product, Safety Instructions, Corrective Action and Contact Information.
Z-1654-2009 | Jul-07-2009 | 2009 | IntelliVue - Philips IntelliVue Clinical Information PortfolioSoftware Version D.0; for patient data collection; storage; and management. | Philips Healthcare Inc. | 2 | Error in therapy scheduling: If scheduled orders are documented and edited prior to the orders' original schedule time; ICIP can create a second instance of the same order marked as pending. A second instance of an order may result in a patient receiving duplicate therapy or patients may not receive their scheduled therapies. | Philips notified customers by Urgent Medical Device Correction Notification letter on May 6, 2009. The Customer notifications in the United States will be delivered via UPS. Customer notifications outside the US will be managed by Philips representatives in each affected geography. Customers will be asked to follow the guidelines in the Action to be taken by Customer/User section of the Urgent Medical Device Correction Notice until they receive their software corrections. The corrections will be in the form of software upgrades and will be performed by Philips service personnel. On May 27, 2009, Philips issued the second Urgent Medical Device Correction notification. While customers await the correction of their devices, customers are given the following instructions: Clinicians should not rely solely on the worklist within lCIP to determine actionable orders for the patient until their software is upgraded. Clinical users should always use the medication administration record (MAR) and the administration record in determining the care their patients receive in addition to the critical care worklist.
Z-1436-2009 | May-27-2009 | 2009 | CareVue Chart - CareVue Chart Release C.OPart Number: 862246Intended for use in data collection; storage; and management with independent bedside devices; and ancillary systems that are connected either directly or through networks. It is indicated for use by healthcare providers whenever there is a need for generation of a patient record and computation of drug dosage. | Philips Healthcare Inc. | 2 | Software: A software defect under certain circumstances could lead to the unintended administration of a duplicate dose of a scheduled medication within a single 24 hour period | Philips Medical notified accounts by letter on 5/6/09 the Urgent - Medical Device Correction notification. While customers await the correction of their devices, customers are given the following instructions: Clinicians should not rely solely on the worklist within CareVue Chart to determine actionable orders for the patient until their software is upgraded. Clinical users should always use the medication administration record (MAR) and the administration record in determining the care their patients receive in addition to the critical care worklist. Philips is currently developing a correction that will be available at no additional charge. A Philips representative will notify customers when the software correction becomes available. If further assistance is needed, customers are to contact the Philips Customer Care Center at 1-800-722-9377.
Z-0133-2010 | Nov-09-2009 | 2009 | VARiS 6.2; 6.1 and 6.0 - VARiS 6.2; 6.1 and 6.0 when used with LINK; Model numbers H88 and H85. Manufactured by Varian Medical Systems; Palo Alto; CA. | Varian Medical Systems Oncology Systems | 2 | Software issue: Under certain circumstances; when a therapy treatment field is transferred to VARiS Vision 6.2; 6.1 and 6.0 via LINK procedures; the rotation direction will be reversed. | Recall initiated September 15, 2009. An "Urgent - Field Safety Notice" letter was distributed to all affected users, with a description of the problem and corrective action steps. The letter was also distributed to Varian sales, marketing, and service organizations, informing them of the issue. For further information in the USA and Canada, contact Varian at 1-650-424-5731 or 1-888-VARIAN5 (888- 827-4265). In Europe, contact +41 41 749 8844.
Z-0121-2010 | Nov-05-2009 | 2009 | ARTISTE syngo RT Therapist - Siemens ARTISTE syngo RT Therapist; Therapist Assist; Therapist Express Basic; and Therapist Express Assist Systems; manufactured by Siemens Medical Solutions; Concord; CA | Siemens Medical Solutions USA; Inc | 2 | A failed data transfer is possible that may result in data not being recorded in the interfaced system. Siemens Medical Solutions; Inc. USA became aware of this potential issue with its ARTISTE Linear Accelerator's Syngo RT Therapist v4.1 Workspace (containing the Siemens brand of Verify and Record application called PRIMEVIEW) and the LANTIS Oncology Information System product. | The firm issued a TH015/09/S Safety Modification syngo RT Therapist 4.1 Customer Advisory letter, dated September 10. 2009. the letter stated the reason for modification, the affected systems, working time, deadline, reimbursement, and reporting. The firm is preparing a software upgrade for the affected products to correct the problem. The letter was sent via certified mail.
Z-1889-2009 | Sep-02-2009 | 2009 | Agfa Healthcare Corp. NX 2008 Central Monitoring System - NX 2008 Central Monitoring System (CMS). Product code: E2FB6Agfa's Computed Radiography Systems with NX Workstations are intended for use inproviding diagnostic quality images to aid the physician with diagnosis. | AGFA Corp. | 2 | After opening and closing a study on a Central Monitoring System while a study is still open on an in-room NX system; a problem of image mix-up can occur on the in-room NX system. | An "Urgent Safety Notice" was sent to all potentially affected sites on/about June 22, 2009 to make them aware of the issue. A FAX Back form is requested with each letter to ensure someone at the account acknowledges the information was received and understood. If a FAX Back form is not received, the account will be contacted by telephone. A software service update to solve this issue is expected to be released by August 15, 2009. A Mandatory Service Bulletin will be deployed to correct the problem in the software. If you have questions, please contact Agfa Healthcare at 1-800-581-2432 prompts 1, 3 and then 2 for CR Support.
Z-0160-2010 | Nov-10-2009 | 2009 | Guardian Services - Guardian Services; Guardian Software; Part Number 0240-009-430 (Guardian Services) with Part Numbers 0250-003-016; 0250-003-017 and 0250-003-018 (OrthoPad)Intended use of Guardian Services is to act as a medical data backup and disaster recovery plan. | Stryker Imaging | 2 | Guardian Services data backup does not contain 100% data for all OrthoPad customers on a daily basis. | Firm notified consignees by an Urgent: Device Correction letter, dated 7/31/2009. The letter states the problem and that Stryker is upgrading its internal infrastructure to remediate the problem. The letter also states the firm's recommendations to reduce risk to patients. Stryker will not bill customers for the Guardian service until the problem is remediated and the firm notifies customers. If there are any questions, customers are to contact the firm at 972-410-7100.
Z-0003-2009 | Sep-22-2009 | 2009 | Revolution XR/d - Revolution XR/d systems with control room PC model 5117866-2 with Software collector 5194381. GE Healthcare; 3000 North Grandview; Waukesha; WI 53188.Intended for use in generating Tomographic images of human anatomy. | GE Medical Systems; LLC | 2 | GE Healthcare has recently become aware of an error in the generator software on the Revolution XR/dconfigured with HP 8200 PC consoles that reports a mAs reading that is higher than is actually exposed to patients. | Consignees were sent a GE Healthcare "Product Safety Notification" letter dated March 7, 2008. The letter was addressed to Hospital Administrators, Mangers of Radiology, and Radiologists. The letter listed the Safety Issue, Affected Product, Safety Instructions, and Contact Information.
Z-0985-2009 | Mar-02-2009 | 2009 | Orthosoft Navitrack; OS Unicondylar Knee 1.0 - Orthosoft Navitrack; OS Unicondylar Knee 1.0 - Universal; software application for orthopedic surgical stereotaxic instrument. | Zimmer Inc. | 2 | Incorrect parameters: If the user switches from one surgeon profile to another surgeon profile after completing anatomical landmark acquisition tasks; the application will continue to use the initial surgeon profile parameter values in the calculation of the predicted hip-knee-ankle alignment estimates; which may have different resection or component thickness parameters than are appropriate for | Zimmer notified consignees via letter dated 8/19/08, which advised them not to change the selected surgeon profile after the acquisition of landmarks has begun, and the issued will be fixed with the next software release.
Z-2253-2009 | Sep-30-2009 | 2009 | VEST 6.6 or earlier software - VEST 6.6 or earlier software for use with Neuro Kinetics products.The device is used for vestibular testing. | Neuro Kinetics | 2 | Marketed without a 510k or PMA submission to include the normative data display. | A notification letter was emailed to customers starting on August 10, 2009. The letter stated that the normative display needed to be removed pending FDA clearance. Direct questions to Neuro Kinetics, Inc. by calling 1-412-963-6649.
Z-0999-2009 | Apr-07-2009 | 2009 | Viasys Healthcare; Sonora Transcranial Doppler (TCD) System - Viasys Healthcare; Sonora Transcranial Doppler (TCD) System. Model Sonara. Part Number of Units: 982A0470 Sonara Bilateral.The Sonara Transcranial Doppler system is a medical ultrasound device for noninvasively measuring the blood flow velocities in arteries and veins. | Cardinal Health; Inc | 2 | Internal testing has identified several several scenarios where unexpected System lockups; shutdowns or system errors would render the Sonara unavailable for use until the Sonara can be restarted. These issues can be corrected by installing the Version 6.0 b045 software updating the Audio Driver and updating to XP Service Pack 3. Additionally; improper wiring could cause the cooling fan to not op | Consignees were sent a Cardinal Health "Urgent: Medical Device Correction" letter dated January 12, 2009 beginning the week of February 2, 2009 and continuing into week of February 9, 2009. The letter was addressed to customers describing the problems and systems affected and solutions. Cardinal Health requested customers to complete the actions listed in the letter to install the Field Correction and return the completed "Verification Form" via mail, e-mail or fax to Cardinal Health-NeuroCare. Further information can be obtained by calling 1-800-356-0007 ext 2191 or fax to 1-877-244-3805.
Z-1810-2009 | Sep-14-2009 | 2009 | CyberKnife Robotic Radiosurgery System - CyberKnife Treatment Delivery System; a subsystem of the CyberKnife Robotic Radiosurgery System; manufactured by Accuray Inc.; Sunnyvale; CA.Indicated for treatment for planning and image-guided stereotactic radiosurgery and precision radiotherapy of lesions; tumors and conditions anywhere in the body when radiation treatment is indicated. | Accuray Inc | 2 | Targeting accuracy out of specification; Error alert does not render the system down; which may result in mistreatment in the wrong area. | Accuray, Inc. issued an "Urgent Device Correction" notice dated June 22, 2009. The consignees were informed of the defect including methods for detecting the error. The firm is in the process of developing a software patch to address the issue. For further information, contact Accuray, Inc. Customer Support at +1-877-668-8667 (US) or +1-408-716-4700 (outside US) or customersupport@accuray.com.
Z-0168-2010 | Nov-10-2009 | 2009 | Guardian Services - Guardian Services; Part Number 0240-009-430 with Part Numbers: 0240-030-004 and/or 0240-030-002; 0240-030-008 (Net Practice).The intended us of Guardian is to act as a medical data backup and disaster recovery plan. | Stryker Imaging | 2 | Product serving as data backup protection does not contain 100% data backup for all NetPractice customers on a daily basis. | Firm notified consignees by an Urgent: Device Correction letter, dated July 31, 2009. The letter states the problem and that Stryker is upgrading its internal infrastructure to remediate the problem. The letter also provides Stryker's recommendations for reducing risk to patients. Technical instructions on the backup storage for the NetPractice are included with the letter. Stryker will not bill customers for the Guardian Service until the problem is remediated and Stryker notifies customers. If there are questions, customers should contact Stryker at 972-410-7100.
Z-1235-2009 | Apr-28-2009 | 2009 | syngo Imaging - syngo Imaging; picture archiving and communications system (PACS); Model number 10 014 063 | Siemens Medical Solutions USA; Inc | 2 | Incorrect and/or incomplete data: Several potential safety issues while working with the syngo Image product including; but not limited to; data loss; errors in patient data due to different patient IDs; delays in image availability; inaccurate results during image measurements; incorrect review or diagnosis of displayed or printed images; measurement inaccuracies from the host modality and image | syngo Imaging has issued an Addendum to the User Manual, which was distributed to affected customers via Update Instructions IM118/08/S in November, 2008. Two safety related hotfixes (VB20G_HF4 and VB30A_HF26) will be released to those affected only after Update Instructions IM118/08/S has been performed. The addendum informs the user about safety relevant notes which pertain to syngo Imaging versions VB20G and VB30A. Update is a prerequisite for the following safety related hotfixes: VB20G_HF4 and VB30A_HF26.
Z-0624-2009 | Feb-04-2009 | 2009 | O-ARM(R) - Medtronic O-Arm Intraoperative Imaging System Catalog Number: BI-700-00027 | Medtronic Navigation; Inc | 2 | Navigation accuracy problem: Accuracy may fall outside of specification when used in combination with the Medtronic Synergy Experience StealthStation System Spine Software with the O-Arm Intraoperative Imaging System and tracker. | Medtronic Navigation notified Field representatives on 11/27/07 to perform site visits to implement FCA 2007-005 - O-Arm Intraoperative Imaging System Accuracy, and deliver the Customer Letter notification. Confirmation of the visit was documented.
Z-0418-2010 | Nov-24-2009 | 2009 | Philips medical Systems - IntelliVue Clinical Information Portfolio (ICIP) Critical Care Release D; Version D.02.Intended for use in data collection; storage; and management with independent bedside devices; and ancillary systems that are connected either directly or through networks. It is indicated for use by healthcare providers whenever there is a need for generation of a patient record and computation of drug dosage. | Philips Healthcare Inc. | 2 | A software defect exists in the ICIP Release 0.02 scheduling function for order management such that if a user edits either the Schedule Time or the Stop Time for an existing order; the corresponding intervention(s) for the order may not be created. | Philips Medical Systems notified accounts by letter on October 21, 2009, of the release ICIP version D.03, which contains a software correction . A Philips representative will schedule the implementation of this upgrade. Please contact the Philips Customer Care Center with questions or concerns about this correction at 1-800-722-9377.
Z-1888-2009 | Sep-24-2009 | 2009 | ARTISTE MV Systems - ARTISTE MV Systems (Part Number 8139789); equipped with a syngo RT Therapist (Part Number 8162815) or RT Therapist Assist system (Part Number 8162807); with Software version NET VA10A for RTT4.1.122 or RTT 4.1.125; manufactured by Siemens Medical Solutions; Concord; CA.Medical charged-particle radiation therapy system. | Siemens Medical Solutions USA; Inc | 2 | Unexpected treatment plan alteration-- Software issue may result in unintended modifications to treatment | Siemens Medical Solutions USA, Inc. notified consignees of the affected product by advisory letter via Certified Mail on April 09, 2009. For further information, contact Siemens Medical Solutions at 1-925-602-8157.
Z-1651-2009 | Jul-07-2009 | 2009 | CXR4 Computed Tomography X-Ray System - Hitachi CXR4 Computed Tomography (CT) Scanners; Software Version 1.51 or less. | Hitachi Medical Systems America Inc | 2 | Incorrect scale on image: A software error can occur if two (2) different Field-of-View (FOV) settings are registered in a protocol using Multi-Reconstruction. This could cause the image generated to be displayed with the wrong scale factor rendering anatomical measurements incorrect. | HMSA issued an URGENT - DEVICE CORRECTION notification to their customers on 4/14/2009. The notification describes the potential software error and the condition under which these image distortions can occur. The notification goes on to provide the customer with a work-around procedure which they can use to avoid encountering this issue with their current software configuration(s). The letter states that the firm will be installing updated software on all affected systems beginning in June, 2009, and that a Hitachi Service Rep. will be contacting each customer and making arrangements to come out and install a software upgrade on the affected system(s) at no cost to the customer. Lastly, the letter provides a telephone # to the firm's Applications Helpline: 1-877-228-2777, which the customer can use to obtain additional information, if desired.
Z-1582-2009 | Jun-29-2009 | 2009 | BRILLIANCE VOLUME - Brilliance iCT software v2.5.0; Model 728306. The Brilliance iCT is packaged INA carton and the software is contained within the Brilliance CT scanner.The 'Brilliance iCT' is a Computed Tomography X-ray System intended to produce cross-section images of the body by computer reconstruction of x-ray transmission data taken at different Ganges and planes. This device may include signal analysis and display equipment; patient; and equipment supports; components and accessories. | Philips Medical Systems (Cleveland) Inc | 2 | Software anomaly. An artifact may appear in brain scans and brain perfusion studies; when using 80kV and a UA; UB; or UC filter with this CT scanning device. | An "URGENT - Field Safety Notice" dated December 22, 2009 was issued via certified mail to affected customers. The letter described the affected product, problem, hazards involved, and actions for customers to take. Philips will provide customers with a software update for all affected systems within next 6 months. The firm also sent a second notification in the form of an "URGENT -Medical Device Correction" report dated January 29, 2009 which contained information noted in the "URGENT - Field Safety Notice". Direct questions to your local Philips representative or Philips Customer Care Center (1-800-722-9377), option 5: Other Imaging Systems, option 1: CT or to your local Philips Healthcare office.
Z-0608-2009 | Jan-08-2009 | 2009 | CXR4 Computed Tomography Scanner - Hitachi CXR4 Computed Tomography Scanner. This product is an x-ray imaging device that produces cross-sectional images of he body at different angles. The system reconstructs; processes; displays; and stores the collected images. The device output can provide an aid to diagnosis when used by a qualified physician and is intended for general purpose CT applications. | Hitachi Medical Systems America Inc | 2 | Software error: The device has a software error that can result in a condition where the displayed image orientation does not match the display orientation markers (right-left reversal). | On 11/5/2008 Hitachi Medical Systems sent recall notification letters entitled "URGENT - DEVICE RECALL CORRECTION" to its customers informing them of the problem with the software. They stated that a Hitachi service representative will contact each customer to schedule a software upgrade. For further information, please contact Hitachi Medical Systems America Inc. by telephone at 330-425-1313.
Z-1263-2009 | May-14-2009 | 2009 | Carina Home Ventilator - Carina Home Ventilator.The device is used as a continuous respirator/ventilator for home use. | Draeger Medical; Inc. | 2 | The device may not generate a pilot line disconnect alarm if the pilot line becomes disconnected; which may potentially result in a substantial reduction in tidal volume delivery. | An "Urgent - Medical Device Recall" Letter dated January 2009 was issued to consignees via certified mail. The letter described the issue, specific conditions, background information, and precautions for customers. An additional letter dated January 22, 2009 was included with the "Urgent - Medical Device Recall" letter which instructed consignees to inform an enduser of the contents of the notification letter because some of the consignees are home care device/service providers. When the new hardware/software solution is available, affected devices will be updated free of charge. Direct questions about the notification letter to Drager Medical, Inc. by calling 1-800-543-5047 (press "1" at the prompt and then press "2349#"). Direct questions regarding updating your device when the new software is available, contact Drager Service Technical Support at 1-800-543-5047 (press "4" at the prompt).
Z-0948-2009 | Jan-30-2009 | 2009 | iSite PACS Versions 4.1.x up to and including 4.1.51.3 - iSite PACS (Picture Archiving and Communications System) Versions 4.1.x up to and including 4.1.51.3; manufactured by Philips Healthcare Informatics; Foster City; CA.Device is an image management system intended to be used by trained professionals; and is a software package used with general purpose computing hardware to acquire; store; distribute; process and display images and associated data throughout a clinical environment. The software performs digital image processing; measurement; communication and storage | Philips Healthcare Informatics; Inc. | 2 | Product may display clinical reports with the study date and time in GMT format; which may not be obvious to the end user; who may expect to see the local time in the report. | Product notification was issued on 12/29/2008, instructing users to identify affected products, and instructing users to refer to image overlays for local study date and time information. The firm is planning a field correction expected to be available in "Q1 2009". For additional information, contact Philips at 650-293-2624.
Z-0857-2009 | Jan-14-2009 | 2009 | Artiste MV Digital Linear Accelerator - Artiste MV Digital Linear Accelerator; when used in combination with Syngo RT Therapist Version 4.1 build 67 and automatic Motion of Flat Panel; part number 8139789; manufactured by Siemens Medical Solutions; Oncology Care System.Intended use is to deliver x-ray radiation for therapeutic treatment of cancer. | Siemens Medical Solutions USA; Inc | 2 | Under certain conditions; the flat panel may move unexpectedly and collide with the patient. | The firm issued an Urgent Medical Device Customer Safety Letter, dated 10/23/2008. With the letter included a detailed description of the issue, the condition under which this issue may occur, and recommendation to avoid them. Customers are to read the information carefully and distribute to the responsible persons. The Customer Safety Letter should be included their syngo RT Therapist Express manual and remain there until the next release of software is installed. In the interest of safety, the preventive measures should be performed and all affected personnel should be informed immediately. The firm is tracking responses via return cards and completion protocol.
Z-0271-2010 | Nov-16-2009 | 2009 | Syngo Imaging XS - Syngo Imaging XS VA60 and VA70 in combination with JPEG Compression | Siemens Medical Solutions USA; Inc | 2 | Distorted image. The "Distorted" image cannot be used for diagnostic findings; and; depending on the customer scenario; a repeated exam could be necessary if the image cannot be reported on any other workplace or recovered from another source. | The recalling firm issued a Customer Safety Advisory notice dated 9/14/09 to all affected customers via Update Instructions IM100/09/S. The notice informs the customer of the potential issue and provides instructions to avoid its occurrence. To avoid the potential risk of the issue, the local compression function can be disabled as described in the letter. Depending on the image data, the disabling of compression will result in increased storage usage. The issue will be resolved in the release of version VA70B. For all potentially affected customers with "JPEG Compression" license an update instruction will be released which defines the necessary corrective actions in case of "Distorted" images. The update instruction will be performed on request by the customer service group. Further questions regarding this issue are directed to the local SIEMENS Uptime Service Center or +49 (9131) 84-2056.
Z-1397-2009 | May-15-2009 | 2009 | Centricity TriWin Laboratory - Centricity TriWin Laboratory Information System; GE Healthcare Integrated IT Solutions; Barrington; IL 60010The Centricity TriWin Laboratory product is a laboratory information system. It is designed to facilitate the general clinical; anatomic pathology; and cytology laboratory workflow such as order entry; results entry; instrument interfacing; results reporting and patient record retrieval. | GE Healthcare Integrated IT Solutions | 2 | A software defect in the Centricity Lab (TriWin) product may incorrectly link patient information to a wrong report. Also; when viewing results in Patient Viewer; and a coagulation result is higher or lower than the linearity limit; then all of the results following the higher or lower result will show the same result as the first flagged result. | GE Healthcare notified their Centricity TriWin Laboratory customer via letter dated 11/17/08 of the software anomalies resulting in patient safety issues involving the patient information incorrectly being linked to a wrong report, and when viewing results in Patient Viewer, and a coagulation result is outside of a linearity limit, then all of the results following the out of linearity result will show the same result as the first flagged result. The user was informed that a software patch has been prepared to resolve the issues and will be installed on their system. Any questions were directed to GE Technical Support at 888-778-3375 or 905-479-1249, select option 1. The accounts was requested to complete the enclosed customer reply form, supplying the workstation software or build number, and fax it back to 905-305-0046 to schedule the software upgrade.
Z-0970-2009 | Feb-03-2009 | 2009 | GE Centricity PACS RA1000 Workstation - GE Centricity PACS RA1000 Workstation (for diagnostic image analysis); GE Healthcare Integrated IT Solutions; Barrington; IL 60010.The Centricity PACS (Picture Archiving and Communication System) Workstation is intended for use as a primary diagnostic and analysis tool for diagnostic images by trained healthcare professionals; including radiologists; physicians; technologists; clinicians and nurses. It is also intended for use as a clinical review workstation throughout the healthcare facility. The workstation interface provides the user with a means to display; manipulate; archive; print; and export images when connected with the Centricity PACS infrastructure. | GE Healthcare Integrated IT Solutions | 2 | Unread exams may inadvertently get marked as dictated in PACS workstation when the user signs off the report in Talkstation and opens the next exam. | GE Healthcare Integrated IT Solutions notified their Centricity PACS RA1000 Workstation with Talk Extend software customers via letter (Urgent Medical Device Correction) dated 11/10/08 of the software anomaly resulting in a patient safety issue involving Talk Extend, which could contribute to a delay of care or misdiagnosis due to delay in communicating pertinent exam results or failure to treat due to lack of an attached dictated report by the radiologist for physician review. The letter provided the users with safety instructions to follow until the software is updated. Any questions were directed to the Remote Online Center, Field Engineer or the GE Customer Care Center at 800-437-1171. The accounts were requested to complete the enclosed customer reply form, supplying the workstation software or build number, and fax the completed form to 847-939-1479.
Z-1266-2009 | Apr-28-2009 | 2009 | syngo Imaging - syngo Imaging model number 10014063. Affected systems are V30A or earlier with an NFS connected archive and V30A systems where image data was stored to LTS using software version older than V30A.Picture Archiving and communications system | Siemens Medical Solutions USA; Inc | 2 | An issue concerning archived image data that is sporadically saved with corrupted data may occur. This issue may occur when syngo Imaging is configured to archive image data on a network file system (NFS) connected via a connected archive (Mount Point Solution). | A Customer Safety Advisory Notice, dated 02/09/2009 was sent to affected customers via Update Instruction IM013/09S. The letter informs customers of when the malfunction occurs and the potential issue, provides instructions to avoid its occurrence, and how the issue will be resolved. Customers are to ensure that the Safety Advisory is placed in the system's instructions for use and contact their SIEMENS Uptime Service Senter if they have any questions.
Z-0001-2009 | Jan-26-2009 | 2009 | NEUVIZ DUAL MULTI-SLICE CT SCANNER SYSTEM - NeuViz Dual Multi-Slice Computed Tomography (CT) Scanner System (software).Whole body x-ray computed tomography scanner which features a continuously rotating tube-detector system and functions according to the fan beam principle. The system software is an interactive program used for x-ray scan control; image reconstruction; and image archive/evaluation. It is intended to produce cross-section images of head and body by computer reconstruction of x-ray transmission data taken at different angles. The function of Time Scan is intended to use in contrast scan include single phase and multi-phase scan; that is according to different density of contrast in vessels or organs between different phase to get necessary information to support doctor's diagnosis. | Philips Medical Systems North America Co. Phillips | 2 | An issue associated with the "Timed Scan" application was found. The system does not wait for the programmed time interval and starts the x-ray after selection of the scan button when certain criteria is met. | A letter was issued to customers prior to the release of the field correction during June 2008. The system software will be updated to correct the non-compliances. For information or support concerning this issue, please contact the Neusoft Medical Systems Service Support Department (helpdesk@pnms.neusoft.com) or contact Philips Medical Systems North America Co. Phillips at 425-487-7602.
Z-1885-2009 | Aug-27-2009 | 2009 | IntelePACS Picture Archiving Communications System - IntelePACS Picture Archiving Communications System InteleViewer module; Model #3.X.X. | Intelerad Medical Systems Inc. | 2 | Product software versions 3.5.2.P1 to 3.5.2.p56 inclusive and 3.6.1.P1 to 3.6.1.P82 inclusive were defective. | Intelerad Medical Systems issued an Urgent Medical Device Filed Safety Notice" dated October 24, 2008 describing the affected product and steps required for system upgrade. Users are requested to complete and return an acknowledgement form via email to quality@intelerad.com. For further information, contact Intelerad Medical Systems at 1-514-931-6222.
Z-1650-2009 | Jul-07-2009 | 2009 | Syngo Imaging Version V31 - syngo Imaging Version V31 Picture Archiving and Communication System; model number 10014063. | Siemens Medical Solutions USA; Inc | 2 | Risk of misinterpretation of patient response to treatment: Date displayed may be incorrect. Filtering based on examination date may lead to an incorrect list of examinations. | syngo Imaging has issued a Customer Safety Advisory Notice, dated May 28, 2009, to affected customers via Update Instruction IM019/09/S. The notice informs the customers of the potential issue and provides instructions to avoid its occurrence. A software update to correct this issue is currently being developed and will be installed on affected systems by sygo's service representatives when it becomes available.
Z-2221-2009 | Sep-18-2009 | 2009 | ACCELERATOR Automated Processing System - ACCELERATOR APS System Input/Output Module (IOM); List Number: 7L01-01; manufactured by Inpecco SpA; Segrate (MI); Italy. | Abbott Laboratories; Inc | 2 | When an ARCHITECT c16000 System is connected to an ACCELERATOR APS System; the software on the ACCELERATOR APS System may not generate an error warning when a Sample Presentation/Sample Queue error occurs; which could result in contamination of subsequent samples. | Abbott Laboratories notified Customers of the affected device via a "Product Correction" letter dated October 27, 2008. Consignees provided with a temporary work around until a service representative upgrades the system software. For further information, contact your local Abbott Laboratories, Inc. service representative.
Z-1577-2009 | Jul-09-2009 | 2009 | XORANCONNECT - XoranConnect; software version 4.1.0.6; Xoran Technologies; Ann Arbor; MI. A web-based software which provides online multi-planar image viewing; offsite data storage; administrative reporting tools; and teleradiology workflow. Server designed to store and serve stored imaging content on a RAID5 storage device. | Xoran Technologies; Inc. | 2 | An incorrect patient name label may appear on the screen displaying the selected patient's scan. | The firm released software version 4.1.0.7 by uploading it through a remote software upgrade on February 3, 2009. User accounts were notified by letter dated February 9, 2009. Direct questions to Xoran Technologies at 1-800-709-6726.
Z-1578-2009 | Jul-06-2009 | 2009 | PHILIPS PLUS COMPUTED TOMOGRAPHY SYSTEM - TumorLoc software application; Software Versions: 23.1.1; 20.7.13; and 20.8.15 for use on the GEMINI TF 16 and GEMINI TF 64 Slice PET/CT Systems; Catalog Numbers: 882470 - GEMINI TF 16 Slice; 882471 - GEMINI TF 64 Slice; and 882473 - GEMINI TF 16 Slice Base Configuration. Model Numbers: 4535 679 83931 - GEMINI TF 16 Slice; 4535 679 94741 - GEMINI TF-64 Slice and 4535 674 41711 - GEMINI TF 16 Slice Base Configuration.Tumor Localization is an application used in radiation oncology departments for localizing target volumes for radiation therapy planning. It includes exclusive features for viewing Respiratory Correlated CT data sets and analyzing motion of target and surrounding anatomy. | Philips Medical Systems (Cleveland) Inc | 2 | A software anomaly in the TumorLoc software for the firm's Brilliance CT scanners was discovered. The same software anomaly was subsequently determined to be present in the firm's GEMINI PET/CT units. The software defect results in a the potential for data to be lost during reconstruction of images in which slices can either be discarded; or interleaved based upon their series ID. | An "URGENT - Field Safety Notice" dated February 12, 2009 was sent via FedEx to domestic customers. The notice described the affected products, issue, involved hazard should problem occur and corrective actions for the customer/user. The letter also notified users/customers that a Philips Healthcare Field Service Engineer will contact and visit their site within 6 months to install a TumorLoc upgrade. Notification will also be sent to the international consignees by Philips personnel. For further information or support concerning this issue, please contact your local Philips representative: Customer Care Center (1-800-722-9377, option 5: Diagnostic Imaging, option 5: Radiation Therapy, option 3: Oncology) or to your local Philips Healthcare office.
Z-2223-2009 | Sep-18-2009 | 2009 | BioTek Gen5CL Microplate Data Collection and Analysis Software - BioTek Gen5CL Microplate Data Collection and Analysis Software.Intended for clinical lab use with ELx800 and ELx808 microplate readers; offers preprogrammed clinical diagnostic assays within a PC-based application. | BioTek Instruments; Inc. | 2 | Software design allowed false negatives to be read when wells are dark. | BioTek Instruments, Inc. issued a "Product Recall Notification" dated March 13, 2009 information customers of the affected product. Follow-up letters with software installation instructions were sent March 25, 2009. For further information, contact BioTek Instruments, Inc. at 1-888-451-5171 extension 405.
Z-0134-2010 | Nov-09-2009 | 2009 | Guardian Services - Stryker Imaging; Guardian Services Software; Part Number 0240-009-430 with Part Numbers 0240-020-013 and 0240-020-022 (Office PACS Power).Intended use is to act as a medical data backup and disaster recovery plan. | Stryker Imaging | 2 | Firm has determined Guardian Services does not contain 100% data backup for 12 OfficePACS Power customers on a daily basis. | Stryker Imaging issued an "Urgent: Device Correction" notice dated August 18, 2009. Customers were notified of the defective product and informed that Stryker will be in contact shortly to coordinate the product correction. Actions to be taken during the interim were also detailed. For further information, contact Stryker Imaging at 1-972-410-7100.
Z-1808-2009 | Sep-04-2009 | 2009 | Horizon - American Optisurgical; Horizon Phacoemulsification System; Model: HZN-2030; for ophthalmic surgery. | American Optisurgical Inc | 2 | Device operation interrupted: A software bug causes Phaco to stop when vacuum goes above 300 mmHg while in Multi Burst mode. Use of this product at the parameters specified above could lead to a delay in surgery that might present a health risk. | American Optisurgery sent an E-mail to customers on April 2, 2009 stating the reason for recall and informed them to immediately examine their inventory and quarantine product subject to recall. In addition, if they have further distributed this product, please identify customers and identify them at once of this product recall. Their notification to customers may be enhanced by including a copy of this recall notification letter. They were then told to refer to the enclosed Recall Questionnaire for instructions on what to do with the recalled product: The recall should be carried out to the user level and return the enclosed response form as soon as possible. If there are any questions, they are advised to contact David Salzberg at (800) 576-1266.
Z-2054-2009 | Sep-09-2009 | 2009 | Hi Vision ultrasound diagnostic scanners - EUB-6500/HI VISION 6500 Diagnostic Ultrasound Scanners; DICOM Option EZU-FC5; Software Versions: V01-03; V01-05; V02-04; V03-02; V04-02; V04-03; V04-04; V05-03; V05-05A; V05-07A and V05-30A. | Hitachi Medical Systems America Inc | 2 | The firm's parent company discovered a DICOM UID (Universal Identifier) code mix-up which can occur due to a software error. This error can cause the images from two (2) different patient scans to become intermingled and consequently; labeled incorrectly. | HMSA issued a 'DEVICE CORRECTION' notification to their customers on 5/12/2008. The notification describes the potential data transmission problem (potential to archive studies done on different patients into the same patient folder) and the conditions under which this problem can occur. The notification cautions the user/customer to verify all images stored on the PACS (Picture Archiving and Communications System) in order to ensure that they do not contain images from a different patient. The DEVICE CORRECTION letter also states that the firm will be installing updated software on all affected systems beginning in June, 2008, in order to address this issue. It also states that a Hitachi Service Rep. will be contacting each customer and making arrangements to set up a convenient time to come out and perform a software upgrade. Lastly, the letter provides a toll-free telephone # to the firm's 'Customer Service Center: 1-800-800-4925, which the customer can use to obtain additional information, if desired.
Z-0945-2009 | Feb-03-2009 | 2009 | IMPAX Cardiovascular Polling Procedure version 1.00.09 - IMPAX Cardiovascular Polling Procedure version 1.00.09; used in IMPAX 6.2.1 and 6.3.1.; Catalog numbers: E18R4000; L5NQJ000; L57OP000; LBCDB000; and E2ZCK000.IMPAX Cardiovascular is a picture archiving and communications system intended to be used as a networked cardiovascular information management system. The IMPAX Centric Polling Procedure provides a mechanism to ensure that patient demographic and merge/split operations performed on the IMPAX database utilizing the IMPAX Service Tool are captured and applied to the IMPAX CV database; in order to keep patient updates from the IMPAX database synchronized with the IMPAX CV database. | AGFA Corp. | 2 | Polling procedure may fail when multiple merge/split is carried out. | All sites potentially affected by this product were notified of the issue directly via telephone. Each site was directed to observe a ten minute polling window as an interim mitigation until the normal mitigation by software Service Update is performed at the site by Agfa Support. An Urgent Safety Notice was sent via FED-EX on/about 12/24/2008 to all sites. The letter describes the potential issue and both interim and formal mitigation. Acknowledgment, via FAX-Back, that the information was received and understood was requested from the sites.
Z-1040-2009 | Mar-24-2009 | 2009 | QCA and iQCA Software - QCA and iQCA software; a video microscopy software system for quantitative estrogen receptor immunohistochemistry; Cell Analysis Inc.; Evanston; IL 60201 or Highland Park; IL 60035.Intended to measure and quantitate the percentage and intensity of positively stained nuclei in formalin-fixed; paraffin-embedded tissue specimens immunohistochemically stained for estrogen receptors. Indicated for use as an aid in the management; prognosis and prediction of therapy outcomes of breast cancer when used with reagents validated for those indications. | Leica Microsystems; Inc. | 2 | A change in the software to open up the software architecture to include other analytes other than Estrogen Receptor was made without FDA review and approval. | Leica Microsystems sent Medical Device Correction letters dated 11/6/08 to the direct accounts, informing them that all versions 3.7 and above of QCA and iQCA software are not in compliance with the original 510(k) cleared software for scoring of immunohistochemical stains in clinical applications. The QCA/iQCA software is only cleared for Estrogen Receptor staining under clinical applications. The accounts were provided with red warning stickers that state: "WARNING The clinical application of QCA software has been cleared only for the measurement of images from microscopic slides of breast cancer specimens stained for the presence of estrogen (ER) nuclear receptor protein only. Performance characteristics of the product, when used outside of this intended use, have not been established." They were instructed to apply the stickers to any existing manuals and or at each point of use, and to instruct all users about the notification. The accounts were also requested to complete and return the enclosed reply form confirming that the changes have been made.
Z-1092-2009 | Mar-12-2009 | 2009 | GE Healthcare Innova 3100 / 3100 IQ - GE Innova 3100 / 3100 IQ; Cardiovascular Imaging System; GE Healthcare; Waukesha; WI 53188 (The system consists of an a monoplane positioner; a vascular or cardiac table; an X-ray system and a digital detector)Indicated for use in diagnostic and interventional angiographic procedures of human anatomy. Intended to replace image intensifier fluoroscopic systems in all diagnostic or interventional procedures. Not intended for mammography applications. | GE Medical Systems; LLC | 2 | GE Healthcare reported an issue with the Exposure Switch Redundancy system with Innova Systems 3100 & 4100; in which this error stays on; even after the pedal is released; inhibiting further x-rays; and a system reset is required to recover. | On site visits were performed from 3/6/08 - 12/19/08 to update software on all 5 units.
Z-1236-2009 | May-01-2009 | 2009 | Cerner Bridge Medical Transfusion Administration and Specimen Collections - Cerner Bridge Medical Transfusion Administration and Specimen Collections; Versions 3.4.0; 3.4.1 and 3.4.2. | Cerner Corp | 2 | Specimen labels printing incorrect patient demographics. | Cerner Corporation issued a "Priority Review Flash" dated January 15, 2009 via e-mail to consignees using the software. The letter described the violative nature of the software and explained the proper navigation controls needed to prevent further issues. Further questions should be addressed to Bridge Client Support at 1-800-786-0239.
Z-0629-2009 | Jan-17-2009 | 2009 | Cyberknife Robotic Radiosurgery System - Cyberknife Robotic Radiosurgery System. A radiation therapy device; MultiPlan (MP) Treatment Planning Software and Iris Variable Aperture Collimator; Software version 3.0.Product is indicated for treatment planning and image guided stereotactic radiosurgery and precision radiotherapy of lesions; tumors and conditions anywhere in the body when radiation treatment is indicated. | Accuray Inc | 2 | System may use random incorrect data to calculate dose. Resulting dose calculation error can exceed 100% of correct dose which may lead to serious patient injury. | An Urgent Advisory Notification letter dated November 14, 2008 was issued to all sites via Fed Ex. The letter identifies the issue and required actions. The user was instructed not to modify current extrapolation depth settings and to contact Accuray Incorporated Customer Support prior to making any changes. Please direct any questions or concerns regarding this issue to Accuray Incorporated Customer Support at 1-877-668-8667 (USA) or 1-408-716-4700 (outside of USA) or customersupport@accuray.com.
Z-0174-2010 | Nov-10-2009 | 2009 | Syngo Imaging XS - Syngo Imaging XS VA60 and VA70 in combination with a NAS or HSM license | Siemens Medical Solutions USA; Inc | 2 | Incomplete amount of image data; specifically; The image count in the series table of the online and archive database will display the full amount of images received at the Syngo Imaging XS as archived. However; the zip file on the NAS or HSM will contain fewer images-those which have been archived until the time of the error/interruption. | The recalling firm issued a "Customer Safety Advisory Notice" dated 9/14/09 to all affected customers via Update Instructions IM101/09/S. The notice informs the customer of the potential issue and provides instructions to avoid its occurrence. Questions regarding this issue are directed to the local SIEMENS Uptime Service Center.
Z-2231-2009 | Sep-22-2009 | 2009 | Eclipse Proton Convolution Superposition DCF; version 8.6.15; Model H48 - Varian Medical Systems; Eclipse Proton Convolution Superposition DCF; version 8.6.15; Model H48; radiation therapy treatment planning system. | Varian Medical Systems Oncology Systems | 2 | Software Issue/Mistreatment Potential-- Due to software issues; dose distribution may be incorrect and unintended radiation doses may be delivered. | Varian Medical Systems Recall initiated on July 17, 2009. A product notification letter was distributed to the affected user, with a description of the problem and user corrective action steps. The letter will be distributed to the sales, marketing and service organizations via certified mail. A software correction is planned.
Z-1791-2009 | Aug-03-2009 | 2009 | ALARIS PC UNIT - The Alaris PC Unit; Model Numbers 8000 and 8015 (formerly Medley PC Unit) software version 8 and above when used with the Alaris PCA Module (Model 8120).The Alaris PC Unit is the central programming and power supply component for the Alaris System modules. | Cardinal Health 303 dba Cardinal Health | 1 | Affected Unit: Alaris PC Unit (Models 8000 and 8015) with software a versions 8 and above when used with the Alaris PCA module (Model 8120) There is a potential risk of over infusion related to a warning message on the user interface of the Alaris PC unit when used in conjunction with the Alaris PCA module. When the Alaris PC unit detects a programmed infusion for the Alaris PCA module that exceed | The recall was initiated on June 12, 2009 with the firm forwarding via UPS an Urgent: Medical Device Recall Notification with attached Customer Response Card to the Chief Administrative Officer, Director of Pharmacy, Director of Biomedical Engineering, Director of Nursing & Director of Risk Management of all of the affected consignees. Customers were also asked to please promptly complete and return the enclosed Customer Response Card to expedite the correction process. The Recall Notification informed the consignees of: (A) PROBLEM #1: Occlusion Warning Message for Alaris Pump Module; (B) PROBLEM #2 Syringe volume Warning Message for Alaris PCA module; (C) PROBLEM #3 Electrostatic Discharge (ESD) protection circuitry; (D) PROBLEM #4 Fluid Ingress and (E) PROBLEM #5 Alaris System Inter-Unit Interface (IUI) Connectors Labeling. The recall notice included "Required Action for Users & Follow-up actions by Cardinal Health for all five problems. Customers were also asked to report any adverse reactions experienced with the use of this product, and/or quality problems to the FDA's MedWatch Program. In addition, customer were provided the following contact information for the recall: (1) Cardinal Health Recall Center for general questions: By phone at 1-888-562-6018 7am to 5pm (Pacific) ; (2) Customer Advocacy for adverse reports: By phone at 1-800-854-7120, Option 1, Option 1, Option 3 or by email at customerfeedback@cardinalhealth.com 24 hours a day, Sunday-Saturday; (3) Technical Support for Technical Questions regarding the Alaris System: By phone at 1-888-562-6018, 6am-5pm (Pacific). The firm followed-up with the release of a nationwide press statement regarding this issue on July 29, 2009. The press release indicated serial numbers of affected devices, as well as CareFusions short term instructions to customers, and the firms strategy to fix the affected devices can be found at: www.cardinalhealth.com/alaris/medical-device-recall/ (This informatio.
Z-1006-2009 | Mar-09-2009 | 2009 | AED 10 - Welch Allyn AED 10 automated external defibrillator and MRL Jumpstart (collectively "AED 10") | Welch Allyn Protocol; Inc | 1 | Reliability issues - potential to shock a non shockable rhythm or not shocking a shockable rhythm; | Welch Allyn sent a Customer Recall Notice dated February 26, 2009, to all customers (customer list supplied by distributors). Welch Allyn will upgrade all units with software version 2.02 and lower to version 2.06.02 by Proc. AED10 Prod Service Min HW/SW Update Rev. E 820-0178-00.
Z-1005-2009 | Mar-10-2009 | 2009 | Welch Allyn AED - Welch Allyn AED 10 Automated External Defibrillator | Welch Allyn Protocol; Inc | 1 | Potential for device to shut down prematurely under certain circumstances. | On 4/13/07 the firm released Service Bulletin part no. 810-2614-00, SB - AED 10 S/W V020600 DOWNGRADE to V020500. That mandatory bulletin, for units with software version 2.06.00 and up indicates that all returning AED 10 units shall have their software version downgraded to 2.05.00.
Z-0122-2010 | Nov-20-2009 | 2009 | Stryker Instruments Navigation System II- CART - Stryker Instruments Navigation System II- CART; Stryker Navigation; Kalamazoo; MI; Stryker REF 7700-100-000.Intended to be used as a component of the Navigation System II. | Stryker Instruments Div. of Stryker Corporation | 1 | Multiple software-related malfunctions may occur. The device may suddenly stop working; the screen may freeze; the screen may only be updated very slowly; the unit may not respond to user input; or the unit may only respond to user very slowly. | The firm issued a press release on 11/4/09. Consignees were notified by an Urgent: Medical Device Recall Notification letter dated 10/23/09, which advised of the affected products, the reason for the recall, the recall plan, and the actions to be taken. Customers were to immediately review this notification and distribute to all departments affected. Also, customers were to immediately quarantine any affected product and label as recalled product; complete and sign the enclosed Business Reply Form and fax a copy to 866-521-2762. Upon receipt of the form, a Stryker Navigation Field Service Rep would contact them to schedule the repair of their equipment. All U.S. units were corrected by 10/30/09.
Z-1663-2010 | May-26-2010 | 2010 | Bausch & Lomb Millennium Microsurgical System - Bausch & Lomb Millennium Microsurgical System. | Bausch & Lomb Inc | 2 | The AC power cords used on these products are subject to an FDA Alert -- Safety Investigation of Certain Medical Device Power Cords: Initial Communication; Date Issued: October 19; 2009. | A "URGENT: POWER CORD RECALL" letter and copies of the FDA Alert dated Oct. 19, 2009 were sent to all customers on or about Jan. 15, 2010, by Stericycle, a recall notification service. The letter describes the product, problem and actions to be taken by the customers as well as FDA's recommendations and actions. The customers were made aware of the FDA Alert concerning the AC power cords and notified that these power cords were used in the devices. They were requested to inspect the power cords for damage and notify the recalling firm if damage is noted or if they have experienced sparking, charring, or firing from the power cords and/or cracking or failure of the prongs. In the absence of those conditions, customers were given the option of returning the power cord for replacement and/or during the regularly scheduled servicing of the device. If you would like a new power cord, please contact Customer Service at 1-800-338-2020 for a free replacement.
Z-1271-2010 | Apr-08-2010 | 2010 | French 100-120V 860 TC C02 & 02 Monitor - French 100-120V 860 TC C02 & 02 Monitor; Respironics/Philips product.Used for trending O2 and CO2 tension at the skin surface for neonates and adults as an adjunct to arterial O2 and CO2 measurements. | Philips Healthcare Inc. | 2 | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock. | Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside plug. The issue affects medical device power cords equipped with a plug that has a prong and ground-pin insert design and a black plastic bridge connecting the terminal prongs on the plug ("taller bridge" design). Medical device power cords that do not have a black bridge connecting the terminal prongs on the plug are not affected. Urgent Medical Device Correction letters were sent to Distributors describing the problem and giving them instructions for notifying customers and actions to be taken. Customers are to visually inspect the affected cords, complete the Customer Reply Power Cord Order Form, and follow the instructions to dispose of affected cords. Further information or support concerning this issue can be obtained by contacting a Philips representative at 1-800-722-9377.
Z-0323-2011 | Nov-08-2010 | 2010 | RS-4i Sequential Stimulator; - The device is labeled in part: " RS Medical RS-4i Stimulator". The device is sold in a kit that includes a carrying case; connector cables; pads (electrodes); battery charger; patient document; instructional DVD; return mail bag; and the operation manual.This device is used with cables and electrodes that are placed on skin in the treatment area; and allow an electrical micro-current to relieve pain. The device requires a prescription from a physician. It is programmed by an RS medical employee (or Account Manager) to the physician's specific treatment plan for each patient. The patient is given instructions on how to utilize the device (proper electrode placement). Powered MUSCLE Stimulation indications for use including: 1) Relax muscle spasm. 2) Prevention or retardation of disuse atrophy. 3) Maintain or increase range of motion. 4) Increase local blood circulation. 5) Re-educate muscle. 6) Immediate post-surgical stimulation of calf muscles to prevent venous thrombosis. INTERFERENTIAL Stimulation indications for use including: 1) Relieve acute pain. 2) Relieve and manage chronic pain. | International Rehabilitative Sciences; Inc. | 2 | The charge current is leaking through two Diodes; which are supposed to block the reverse current; when the device is warm during the main battery charge operation; it then causes the lithium battery to swell around two years of use and it causes the device to be inoperable. | Since June 11, 2008, RS Medical has taken the following actions when the device was sent back through a consumer complaint, or when the device was returned for routine maintenance, or through the rental service. The firm has upgraded the blocking diodes with higher temperature rated diodes, replaced the Lithium coin cell battery, and the software register bit has been changed to denote for subsequent automated testing that the diode upgrade/Lithium coin cell battery replacement has occurred. The firm reported that the device must be returned for the rechargeable NiMH battery replacement around two years of use, and the battery should not be replaced by patients. On 10/29/10, the firm began sending out URGENT NOTICE of VOLUNTARY RECALL letters to their customers. The firm will continue to send out the letters through 2011. The letter states the recall product as RS-4i Sequential Stimulator, Affected Serial Numbers 4i300000 to 4i467119. The customers are informed that "RS-4i Sequential Stimulators, only those within the affected serial number range indicated above, are susceptible to early expiration of the lithium coin cell battery that provides back-up power to the device memory when the stimulator is powered OFF. Customer experience indicates that the 3V lithium coin cell battery can be depleted as early as one third of its rated life. The RS-4i Stimulator will perform as intended until the lithium coin cell battery is fully depleted. When this happens, the Stimulator will turn OFF and fail to turn ON, even if it has been charged by battery charger connection to a standard electrical wall outlet". Customers are advised to make a schedule for a free Return and Repair of a RS-4i Sequential Stimulator and should call the Toll Free Service Line at 877-363-5753 (English) or 877-363-5754 (Spanish) from 6AM to 5PM (Pacific Time) Monday-Friday. If the customers have any other questions or to review additional information before scheduling a device .
Z-1440-2010 | Apr-26-2010 | 2010 | Medtronic EnRhythm pacemaker - Medtronic EnRhythm; P1501DR; Dual Chamber Rate Responsive Pacemaker with RapidRead Telemetry. The EnRhythm pacing system is an implantable medical device that monitors; detects; and treats atrial tachyarrhythmia episodes. It also provides bradycardia pacing and monitoring of ventricular tachycardia (VT) episodes. Manufactured at: Juncos; Puerto Rico; USA. Medtronic; Inc.; Minneapolis; MN 55432 USA. | Medtronic Inc. Cardiac Rhythm Disease Managment | 2 | Medtronic is informing doctors of two specific battery issues with EnRhythm pacemakers that will be addressed by a Medtronic software update available mid-2010. If the software update referenced above is not implemented; there will be a potential risk of loss of device functionality or risk for loss of therapy at or near ERI in a small number of devices. The software will eliminate this issue; fo | Consignees were given a Medtronic "Important : Medical Device Correction" letter addressed to "Dear Doctor" and dated February 2010. In the letter, Medtronic is informing doctors of two specific battery issues with EnRhythm pacemakers that will be addressed by a Medtronic software update available mid-2010. The letter included recommendations including that Medtronic recommends physicians continue to use the ERI notification to determine time for device replacement. At this time, no other action, reprogramming or change in the frequency of follow-up is recommended. For additional information or assistance, please contact your local Medtronic Representative or Medtronic Technical Services at 800-505-4636.
Z-1317-2010 | Apr-09-2010 | 2010 | Philips Medical HeartStart MRx Defibrillator Monitor - HeartStart MRx Defibrillator Monitor (Models M3535A; M3536A and M3536J) used with Philips 14.8-Volt Lithium Ion Battery (ModeI M3538A).The HeartStart MRx is for use for the termination of ventricular tachycardia and ventricular fibrillation. The device is for use by qualified medical personnel trained in the operation of the device and qualified by training in basic life support; advanced cardiac support; or defibrillation. It must be used by or on the order of a physician. | Philips Healthcare Inc. | 2 | Batteries older than 2 years have a higher potential for intermittent electrical contact; may contribute to an interruption in monitoring or therapy if the monitor/defibrillator is operating on battery. | An "Urgent Medical Device Correction" letter with a Field Safety Notice was issued on March 4, 2010 to affected customers. The Instructions for Use Addendum/Application Note is included with the communication describing when to replace the unit's 14.8-Volt Lithium Ion Battery (Product #M3538A) Number M3538A in the the Philips HeartStart MRx Monitor/Defibrillator. Should you have any questions or concerns about the Device Correction, please contact your local Philips representative.
Z-0642-2010 | Jan-14-2010 | 2010 | VNS Therapy Demipulse Generator and VNS Therapy Demipulse Duo Generator - VNS Therapy Demipulse Generator; Model 103 and VNS Therapy Demipulse Duo Generator; Model 104. The Cyberonics VNS Therapy system consists of an implantable VNS Therapy Generator; the VNS Therapy Lead; and the external programming system used to change stimulation settings The Demipulse Generator is an implantable multiprogrammable; device that delivers electrical signals to the vagus nerve. Electrical signals are delivered from the Demipulse Generator to the vagus nerve by the lead. The lead and the Pulse Generator make up the implantable portion of the VNS Therapy System. The external programming system includes the Programming Wand; the programming software; and a compatible computer. The programming software is used with the Programming Wand to allow a physician to read and change the Demipulse Generator settings transcutaneously.Indicated for use as an adjunctive therapy in reducing the frequency of seizures in adults and adolescents over 12 years of age with partial onset seizures. | Cyberonics; Inc | 2 | Under certain conditions; product's battery life can be reduced. | Firm sent "An Important Safety Notice" letter to all physicians explaining events, risks presented and recommended actions. Letter was mailed November 20, 2009. Physicians should complete and return the effectiveness card as soon as possible to acknowledge that you have read and understood the Safety Alert Letter. If further information is needed, customers in the US can contact Clinical Technical Support at 866-882-8804 or via email at clinicaltechnicalservices@cyberonics.com.
Z-0620-2011 | Dec-14-2010 | 2010 | Spacelabs Medical Flat Panel Display - Spacelabs Medical Flat Panel Display is a touchscreen display; 19-inch thin film transistor liquid crystal display (TFT-LCD). The PNs are 010-1619-00 and 010-1655-00; Models 94260-19 and 94256. | Spacelabs Healthcare; Llc | 2 | The power supply of the Spacelabs Medical Flat Panel Display may fail; causing the display to go blank. | Spacelabs Healthcare sent an URGENT: MEDICAL DEVICE RECALL Spacelabs Medical Flat Panel Power Supply PNs: 010-1619-00 & 010-1655-00 letter dated March 20, 2009 to US consignees. On March 27, 2009, Spacelabs sent the same recall letter with translation option to their international consignees.The letter stated the firm would be sending a replacement power supply. While waiting for the replacement of the power supply, Spacelabs Medical recommended their consignees brief the staff on the possibility of the display going blank to help prevent any adverse event from happening. The consignees were advised to dispose the old power supply per their premise protocol or local regulations. The consignees from US can contact Spacelabs Medical at 1-800-522-7025, #2 for Technical Support , and consignees outside the United States can call 1-425-657-7200 x5089 for any questions about the recall.
Z-2478-2010 | Sep-23-2010 | 2010 | Fabius GS - Fabius GS Anesthesia Machine; OR/Anesthesia; Draeger Medical; Inc.; 3135 Quarry Road; Telford; PA 18969inhalation anesthesia machine for use in operating; induction and recovery rooms | Draeger Medical Systems; Inc. | 2 | potential failures to Electri-Cord Manufacturing AC power cord | Draeger Medical, Inc., sent an Urgent - Medical Device Recall letter dated August 2010 to all customers. The letter described the product, the problem, and actions to be taken by the customer. Customers were instructed to: 1) Immediately take the cord out of service if the plug has bent or cracked prongs, burned plastic, or shows signs of excessive wear. 2) Monitor the affected cords regularly and be mindful of excessive wear and tear, misuse, or abuse until the affected cord is replaced. 3) Customers were also asked to complete and return the enclosed Customer Power Cord Form via mail: Draeger Medical, Inc., 3135 Quarry Road, Telford, PA 18969; fax: 1-215-721-5808; or e-mail: kathy.benedetto@draeger.com. 4) Disable the cords by cutting off the plug and dispose of the cord in accordance with their facility's disposal procedure when affected power cords are removed. NOTE: The firm will send replacement power cords free of charge. Draeger will only replace power cords originally provided by Draeger (identified with the white flag/part number 4117266. If you have any questions regarding this letter please call 1-800-543-5047 (press 1 at the first prompt, 2 at the second prompt and then 32349#).
Z-2234-2010 | Aug-17-2010 | 2010 | Hospira Accessory Power Pack - Hospira Accessory Power Pack; a 4 outlet power strip with a custom pole clamp bracket to attach to the IV pole with the infusion pump; Distributed by: Hospira; Inc.; 275 N. Field Drive; Lake Forest; IL 60045; part number 39003-04-01; Modo #HSP1700; Model Tripp-Lite PS-415-HGAccessory power pack used for use with medical devices. | Hospira Inc. | 2 | The Power Pack Accessory is not certified to UL 60601 standard or Subject 1363A for relocatable power taps intended for use with medical equipment. Potential problems could include electrical shock; electrical short or electrical fire. | Hospira sent Urgent Device Recall letters dated 7/1/10 to their direct accounts on 7/2/10 via Federal Express 2 day delivery, informing them of the potential for electrical shock, electrical short, or electrical fire, since the power packs were not certified to UL 60601 standard or Subject 1363A of UL standard 1363 for relocatable power taps intended for use with medical equipment. The accounts were instructed to examine their inventory for the affected product and quarantine it immediately, and Stericycle at 1-800-321-7388 to obtain a return kit for returning the affected product. If they distributed the affected product further, they were instructed to sub-recall the product from their sub-accounts. They were also requested to complete the enclosed reply form and return it to Hospira by fax at 1-262-577-6921, indicating the amount of product placed on hold for return, if they further distributed the product and if they notified their retail customers. For medical inquiries, the accounts were directed to contact Hospira Medical Communications at 1-800-615-0187.
Z-0918-2010 | Mar-02-2010 | 2010 | 9200 Advisor Vital Signs Monitor - 9200 Advisor Vital Signs Monitor uses power cord WW3005 (Electri-Cord Manufacturing Co.; 312 Main St.; Westfield; PA 16950). | Smiths Medical PM; Inc. | 2 | Smiths Medical is conducting a voluntary recall of a limited number of medical device AC power cords; manufactured by Electri-Cord Manufacturing Co.; for use with BCI 3180 Pulse Oximeter(Human); 9200 Advisor Vital Signs Monitor(Human or Veterinary); 8400 Capnocheck II Capnometer(Human or Veterinary) and V6400 Invasive Pressure Monitor (Veterinary).The affected Electri-Cord AC power cords are eq | Consignees were sent a Smiths Medical PM, Inc letter dated January 12, 2010 starting on 1/15/10. The letter was addressed to Risk/Safety Managers, Distributors, Clinicians, Critical Care Physicians, Pediatricians, Pediatric Intensivists, Neonatologists, and Nursing, Clinical/Biomedical Engineering, and Anesthesia Professionals. The letter described the problem and products involved. It provided "Advice on Action to be Taken by the User" and requested the return of the "Urgent Recall Notice Confirmation Form." Customers are advised to monitor affected equipment. To return Power Cords, for questions regarding this Urgent Recall Notice, or to report any issues with these Power Cords, please contact the Smiths Medical Technical Service Department at 1-800-558-2345.
Z-1939-2010 | Jun-30-2010 | 2010 | Power cord used on Bio-Console; autoLog; Sequestra - Hospital Grade Power Cord; UL; Part # 66319; 055476; manufactured by Electri-Cord Manufacturing Co. beginning in February 2003. The cords are a detachable unit; supplied with the instruments listed below: Medtronic Bio-Console¿ 550 centrifugal pump console; Model #95180; Catalogue # 95180; E95180; R95180.Medtronic Bio-Console¿ 560 centrifugal pump consoles Model # 560BC; Catalogue # 560BC; 560BCS; 560BCS1.Medtronic autoLog¿; Model # ATLG110; Catalogue # ATLG110; ATLG110EATLG110R.Medtronic Sequestra¿ 1000; Model #SEQ1000; Catalogue # SEQ1000; SEQ1000E; SEQ1000R. | Medtronic Cardiovascular Revascularization & Surgical Therap | 2 | The potential risks during the use of this power cord include electrical shock; delay in setup and therapy; interruption of therapy; device failure; and fires. Although Medtronic has received no reports of adverse events related to the affected cord; a failure may potentially lead to serious adverse health consequences; depending on the device; and therapy being interrupted. This Field Notificatio | Consignees were sent on 4/21/10 a Medtronic "Urgent Medical Device Recall Notice" letter dated April 19, 2010. The letter was addressed to their customers and described the problem and the product involved. They recommended to discontinue use of the AC power cords and to destroy them. Requested consignees to fill out the attached Medtronic Recall Certificate (FCA 1005), and return it to Medtronics.
Z-2029-2010 | Jul-17-2010 | 2010 | Medtronic StealthStation S7 System - Medtronic StealthStation S7 System | Medtronic Navigation; Inc | 2 | Excessive wear at the power cord retractor or incorrect assembly of the power plug may cause sparking; intermittent loss of power; or electrical shock. | Customers were notified by letter on Jun 14, 2010, advising them of the problem. Instructions were provided to continue using existing retractable power cords until they could be replaced with non-retractable cords by Medtronic, beginning on Aug 17. Further information is available from the company by calling 800-595-9709.
Z-0718-2011 | Dec-17-2010 | 2010 | Mobilett XP Hybrid/Digital/CR - Mobilett XP Hybrid/Digital/CRIntended use: mobile x-ray system | Siemens Medical Solutions USA; Inc | 2 | battery fuse | SIEMENS released a field update dated October 25, 2010, via SP011/10/S to the representatives. The update described issue and corrective action plan. The representatives are to visit each account to upgrade the systems with a new battery fuse and a new holder. If you have any questions regarding this report, please contact Quality Engineer by phone 610-448-3237 or email: meredith.adams@siemens.com and Director, Product complaints & Field Support by phone 610-448-4634 or email: josephine.mcbride@siemens.com.
Z-0359-2011 | Nov-15-2010 | 2010 | GE Healthcare Dash 3000; 4000 and 5000 patient monitors - Power supply boards on Dash 3000; 4000 and 5000 Monitors.GE Healthcare; 8200 West Tower Avenue; Milwaukee; WI 53223. | GE Healthcare; LLC | 2 | The power supply board in GE Healthcare Dash 3000; 4000 and 5000 patient monitors may fail due to insufficient solder of a component. Insufficient solder on the power supply board could result in excess EMC emissions; smoking of components and eventual failure of the power supply; tripping of the fuse and immediate shutdown resulting in a loss of monitoring that could delay recognition of a clinic | GE Healthcare issued an "Urgent Medical Device Correction" letter dated October 8, 2010, which described the loss of monitoring issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information. Customers were instructed to remove the units from use. Local service representatives will contact firms to schedule installation of replacement power supply boards. GE Healthcare can be contacted at 1-800-558-7044, Option 2, Technical Support, about this correction.
Z-1944-2010 | Jul-01-2010 | 2010 | Tabs Professional Monitor; Model 25022 - Tabs Professional Monitor; includes one wall bracket; Model 25022. Stanley Senior Technologies; Lincoln; NE. | Stanley Security Solutions; Inc. | 2 | Product might not indicate low battery alarms or exit alarms when powered by a DC power supply. | The firm made phone calls beginning 3/26/10 to customers who purchased the 12-volt power supply informing them of the recall and instructing them to determine the number of units that will need to be returned for replacement. Important Safety Notice letters dated 3/25/10 were issued via regular mail on 3/26/10 and also providing the reason for the recall, instructions, and customer service contact information to arrange for return of the units.
Z-1082-2010 | Mar-11-2010 | 2010 | CADD¿-Solis ambulatory infusion pumps - Electri-Cord AC Power Cords; IEC 60320 Coupler; Product Reorder Number 21-2145-01 only for use with Smiths Medical CADD¿¿-Solis ambulatory infusion pumps; Model 2100. These cords were sold as an accessory for these pumps under Manufacturing Number A23.164.098.0A; Electri-Cord part Number PC2000403.Electri-Cord AC Power Cords are accessories for use with the CADD¿¿-Solis Ambulatory Infusion Pump. These power cords allow users to run the pump via AC power rather than battery power. The CADD¿¿-Solis pump is indicated for intravenous; intra-arterial; subcutaneous; intraperitoneal; in close proximity to nerves; into an intraoperative site (soft tissue; body cavity/ surgical wound site); epidural space; or subarachnoid space infusion. The pump is intended for therapies that require a continuous rate of infusion; patient-controlled demand doses; or both. | Smiths Medical ASD; Inc. | 2 | Potential for the power cord's prongs to crack and fail at/or inside the plug. Other medical device manufacturers have reported incidents of sparking; charring; and fires from the affected power cords used with their devices.Smiths Medical has received no reports of incidents regarding the use of the affected Electri-Cord powercords with their devices. | Consignees were sent a Smiths Medical, Inc. "Urgent Medical Device Recall" letter dated January 27, 2010. The letter was addressed to Risk/Safety Managers, Clinicians, Critical Care Physicians, Pediatricians, Pediatric Intensivists, Nenatologists, and Nursing, Pharmacy, Clinical/Biomedical Engineering, and Anesthesia Professionals. The letter described the problem, products involved and provided "Advice on Action to be Taken by the User". Consignees were requested to complete and return the "Urgent Recall Notice Confirmation Form". To return Power Cords, for questions regarding the recall notice or to report any issues with the affected product, please contact the Smiths Medical Technical Service Department at 1-800-558-2345.
Z-1409-2010 | Apr-16-2010 | 2010 | Bionic NAVIGATOR Clinician Programmer Model Number SC-7150-4 - Bionic NAVIGATOR Clinician Programmer Model Number SC-7150-4 | Boston Scientific Neuromodulation Corporation | 2 | Toshiba has determined that certain Sony battery packs installed in Toshiba portable computers can overheat; posing fire hazard to customers. Based on Toshiba's investigation with Sony Corporation; they concluded that the battery cells in specific manufacturing lots could be affected. | In a letter dated March 19, 2010 Boston Scientific informed consignees that it determined that certain Sony battery packs installed in Toshiba portable computers can overheat, posing fire hazard to customers, and telling them that they may be in possession of an affected product. They were advised that a Boston Scientific representative will contact them to arrange for an inspection of the Clinician Programmer (SC-7150-4). Until they have determined if the battery pack is affected, consignees were told they can only use their Toshiba portable computer with the original AC adapter power cord supplied with that computer, and were asked to remove the battery pack until its status is determined. Included with the letter was a Reply Verification Tracking Form which consignees were asked to complete the form following the instructions on the form and return to Boston Scientific in the enclosed postage paid return envelope or fax it to (877) 856-3546 by March 31, 2010. If they experience any issues they were told to cease using the Clinician Programmer and contact Boston Scientific's 24-hour Patient Care Hotline at (866) 360-4747. Customers can also contact the Field Action team by phone at (866) 684-7504 or by fax at (877) 856-3546 if they had any questions.
Z-1896-2010 | Jun-24-2010 | 2010 | Trilogy 100 Ventilator - Trilogy 100 Ventilator; Model/Catalog Number 1054260; U1054260 (Domestic). Philips Respironics; Monroeville; PA 15146.Provides continuous or intermittent ventilatory support for the care of individuals who require mechanical ventilation. | Respironics; Inc. | 2 | It is possible that the battery could electrically short; resulting in a potential for fire or a thermal event. | Philips Respironics has contacted all of its consignees by phone and instructed them to remove the potentially affected external battery packs from their inventory and any that may be on their patients/end users. These consignees have been instructed to quarantine batteries they have removed away from flammable materials. Product return/replacement instructions were provided. For further information, contact Philips Respironics at 1-877-387-3311.
Z-1781-2010 | Jun-10-2010 | 2010 | Lifeline AED DDU-100 semiautomatic external defibrillator used with the DBP-2800 Battery Pack - Lifeline AED DDU-100 semiautomatic external defibrillator used with the DBP-2800 Battery Pack; Manufactured by Defibtech; LLC. Seymour; CT 06483.The DBP-2800 battery pack is an accessory to the AED and does not have its own intended use. The AED is intended for use on victims of sudden cardiac arrest when the patient is unconscious; unresponsive and not breathing. | Defibtech LLC | 1 | The AED when used with an affected battery pack; may falsely detect an error condition; cancel charge and not provide therapy | Defibtech, LLC issued an "Urgent Medical Device Safety Information and Correction" notification dated May 20, 2010 by email and hard copy via US Mail. The notification instructed Consignees on proper identification of affected product and steps to take for product replacement. For further information, contact your distributor or Defibtech Technical Support at techsupport@defibtech.com or call 1-877-453-4507 or 1-203-453-4507, 8:30 AM to 5:00 PM Monday - Friday, EST.
Z-1903-2010 | Jul-01-2010 | 2010 | LIFEPAK 20 defibrillator/monitor. - LIFEPAK 20 defibrillator/monitor.Product Part numbers: 3202487-xxx; 3202488-xxx; U3202487-000; 3202488-015; and U3202488-000. | Physio Control; Inc. | 1 | A failure on the power supply assembly can result in either "No DC power" or "No DC or AC power." A failure of DC (battery) power can result in a delay of defibrillation therapy if no AC (line) power is available and the device will not operate. | Firm began mailing Urgent Medical Device Correction notifications to affected customers on May 26, 2010 regarding the power supply to the device. All affected power supplies will be updated; this will be accomplished in two phases. Devices at a higher risk will be updated first in phase 1. The remaining lower risk devices will be updated afterward in Phase 2. Customers are advised to keep the defibrillators in service, and to follow recommended daily Operator Checklist steps, while service updates are scheduled. Customers with questions should contact the firm at 800-442-1142 between 6 am and 4 pm Pacific Time, Monday - Friday.
Z-0593-2011 | Dec-10-2010 | 2010 | i-STAT Analyzer Model 200 - i-STAT Analyzer;Abbott Point of Care; Princeton; NJ 08540Model 200. the 37 degree 37 degree C tests for sodium; potassium; ionized calcium; hematocrit; pH; PCO2; PO2 and the i-stat portable analyzer as part of the i-STAT system; are intended for use in the in vitro quantization of arterial; venous or capillary whole blood. | Abbott Point Of Care Inc. | 2 | i-STAT Analyzers batteries may overheat due to a circuitry failure; resulting in the analyzer becoming uncomfortably hot in the area of the battery component. | Abbott Point of Care issued Urgent Recall Notice letters dated September 20, 2010 to customers. The letter informs users to of the possibility of the unit becoming hot to touch, to remove the unit if this is the case, not to remove batteries unless instructed to do by APOC and to forward notification of the recall to customers if it has been further distributed. Abbott Point of Care can be contacted at 800 366-8020, option 1, about this action.
Z-0470-2011 | Nov-26-2010 | 2010 | HeartStart FR2+ automated external defibrillator (AED) - HeartStart FR2+ automated external defibrillator (AED) Model M3860: Philips brand; configurable ECG on screen; configurable manual charge in advance mode. | Philips Medical Systems | 2 | Failure of the voltage detector; U8 may cause the AED battery to drain more rapidly than normal or make the AED unusable. | Philips Medical Systems issued an Urgent Medical Device Recall letter dated May 28, 2010 identifying the affected device and actions to be taken. Customers were instructed to review their inventory for affected devices. If affected devices are found, the customer should contact Philips to arrange for a replacement device. The customer should also run a battery insertion test. If the device passes the test, the AED may remain in service until a replacement is received. For further information, contact Philips at 1 800 263-3342, option 6.
Z-0430-2011 | Nov-22-2010 | 2010 | Medtronic Dual-Chamber External Temporary Pulse Generators Model 5388 - Medtronic Dual Chamber Temporary Pacemaker; model 5388. Manufacturer: Medtronic; Inc.; 710 Medtronic Parkway; Minneapolis; MN 55432-5604; USA. | Medtronic Inc. Cardiac Rhythm Disease Managment | 2 | A subset of Model 5388 Dual-Chamber External Temporary Pulse Generators worldwide may be unable to power up or may power down unexpectedly. The issue presents itself in one of two ways. First; during startup; and prior to initiating patient therapy; the instrument may power down in 1 - 2 seconds. Second; the instrument may power up correctly; but power down at a later time while in use. | Medtronic issued an Important Medical Device Correction letters, dated October 2010 and November 2010 to hospital's biomedical engineering department. The letter described the issue, how it presents, and stated that Medtronic will contact the hospitals to return affected product for servicing. The second letter reiterated that Medtronic would notify customers when FDA approves their design upgrade for returning devices for service. In the interim, hospitals should continue to use normally functioning devices. Cusotmers may contact Medtronic Instruments Technical Service at 800-638-1991 concerning this action.
Z-0833-2011 | Dec-29-2010 | 2010 | Care E Vac 3 Portable Suction - Care E Vac 3 Portable Suction; Battery/AC Portable Aspirator; Ohio Medical Corporation; 1111 Lakeside Drive; Gurnee; IL 60031; Model 758000The device is a portable aspirator to help evacuate saliva; mucous; vomit or other aspirant from the mouth and or airway to allow adequate respiration or ventilation of the patient. | Ohio Medical Corporation | 2 | The D-14 diode on the circuit board may fail; resulting in a damaged; overheated circuit board. Also a screw used to attach one of the feet of the unit may damage the battery inside the device if the screw is over-tightened. | Ohio Medical sent Urgent Medical Device Correction Notification letters dated November 4, 2010, to all affected customers on November 8, 2010. The letters identified the product, the problems, and the action to be taken by the customer. Customers were instructed to take the affected product out of service, remove diode 14, and replace the foot screw contained therein. Customers were instructed to visit the micro-sites established for these repair procedures for step-by-step instructions on performing the repairs. Accounts with questions after viewing the video demonstration were directed to the Ohio Medical technical assistance line at 224-430-1522. Customers were asked to complete the attached Customer Acknowledgment Form and return the form to the Quality Assurance Department at (fax) 1.847.855.6304, or e-mail to: Complaints@OhioMedical.com, or mail to: Ohio Medical Corproation Quality Assurance 1111 Lakeside Drive Gurnee, Illinois 60031. For questions regarding this recall call 847-855-6270.
Z-0798-2010 | Feb-22-2010 | 2010 | Dragonfly 2-Channel Laryngeal Surface Electrode - 2-Channel Laryngeal Surface Electrode (Model LSE500DCS and LSE500DCL)For use as a Surgical stimulator/locator for recurrent laryngeal nerve. | Neurovision Medical Products Inc | 2 | This recall has been initiated due to a conductive glue used to attach the lead wire to electrode; which has been found in a small number of electrodes; to cause a short circuit in one of the channels. This is due to migration of the adhesive from one pad to the other creating an electrical short. Use of one of these defective electrodes could lead to a failure to monitor. Our investigation has c | URGENT RECALL: Medical Device Letter dated 1/7/10. The letter identified the affected product and provided a list of necessary actions for customers to take. The customers should immediately examine their inventory and quarantine product subject to recall. In addition, if they have further distributed this product, please identify their customers and notify them at once of this product recall. Their notification to their customers may be enhanced by including a copy of this recall notification letter. This recall should be carried out to the user level. Your assistance is appreciated and necessary to prevent any failure to monitor. Please complete and return the enclosed response form as soon as possible. If you have any questions, call Ryan Rea or Kevin Rea, at 866-815-6999.
Z-0288-2011 | Nov-02-2010 | 2010 | Spacelabs Medical Flat Panel Display; Model 94260-19 - The product is a 19" Spacelabs Medical Flat Panel Display; Model 94260-19; manufactured by ELo Touchsystems | Spacelabs Healthcare; Incorporated | 2 | The Spacelabs Medical Flat Panel Display; Model 94260-19 may fail prematurely and causing the screen to go blank. This inability to view data may result in delayed patient treatment. | On 10/23/09, Spacelabs issued URGENT: MEDICAL DEVICE RECALL notification letters dated October 23, 2009 to their consignees. The notification states the recalled product as Spacelabs Medical 19" Flat Panel Display, Model 94260-19. The consignees are informed that there is a possibility that the internal power distribution circuit board will fail prematurely, causing the display to go blank. Audio alarms will continue to function normally. The firm's Field Service Engineer will be contacting the consignees to set up a convenient time to schedule the upgrades of the affected displays. Displays with part number ET1920L-7SKR-1-35S-SPC are not part of this recall and should not be sent in for upgrading. The firm recommends that their consignees brief their staff on the possibility of the display going blank to help prevent an adverse event from happening. The firm will be providing loaners while the affected displays are upgraded at free of charge. If the consignees have any questions about the recall, they could contact Spacelabs Medical at 1-800-522-7025 and select 2 for Technical Support . Consignees outside the United States may call 1-425-657-7200 x 5089.
Z-0811-2010 | Feb-22-2010 | 2010 | LIFEPAK 12 DEFIBRILLATOR/MONITOR - Biphasic LIFEPAK 12 defibrillator / monitor. | Physio Control; Inc. | 2 | Affected devices contain a Biphasic PCBA that may have a solder defect on the H Bridge component that could prevent the device from providing a defibrillation shock. There could be a delay in therapy or prevention of defibrillation therapy if the defibrillator were needed in a cardiac arrest situation. Delay or prevention of defibrillation therapy may contribute to the death of a viable patient. | Consignees were sent a notification letter dated March 2009 that explained firm representatives would follow up with each customer within 30 days to arrange an update to affected devices. Affected Biphasic PCBAs will be replaced. Customers advised to keep the devices in use and to continue to perform daily User Test.
Z-1618-2010 | May-24-2010 | 2010 | Philips Heartstart MRx Defibrillator Monitor - HeartStart MRx Defibrillator Monitor; Model M3535AHeartStart MRx Defibrillator Monitor; Model M3536AThe HeartStart MRx is for use for the termination of ventricular tachycardia and ventricular fibrillation. The device is for use by qualified medical personnel trained in the operation of the device and qualified by training in basic life support; advanced cardiac support; or defibrillation. It must be used by or on the order of a physician | Philips Healthcare Inc. | 2 | ECG signal processing chip may fail causing failure to deliver therapy in AED mode or the failure to indicate poor electrical contact (very high impedance) with the patient in Manual Defib mode. | The firm, Philips Healthcare, sent an "URGENT voluntary Medical Device Correction" Field Safety Notice, letter dated April 7, 2009, to all customers via UPS. The letter describes the product, problem and the actions to be taken by customers, with precautions to follow if the device is used before upgrade is installed. The customers were instructed to check their serial numbers for the affected product, if the device is affected, it may remain in service prior to correction using the following precautions 1) if using AED, have a back up AED; 2) if using external defibrillator paddles, disregard paddles and manually ensure good paddlet-to-patient contact, and 3) if monitoring, be prepared to switch back to the primary ECG that was initially obtained for a given patient if other leads are lost. Customers will be contacted and assisted by Philips field service engineers to install the software upgrade which corrects the problem. Should you have any questions or concerns about this Device Correction, please contact your local Philips representative at 1-800-722-9377.
Z-1237-2010 | Apr-05-2010 | 2010 | Axiom Artis Systems operating with A100 Generator - Axiom Artis Systems operating with A100 Generator | Siemens Medical Solutions USA; Inc | 2 | Due to a contact problem at the low voltage power supply of the generator controller; in rare cases where a user has invoked radiation; it may occur that no radiation is released. Under certain circumstances; radiation is possible by repeated release of the x-ray footswitch or button.To remedy the issue described above; a revised cabling has been installed in all potentially affected syste | The recalling firm issued a Customer Safety Advisory to affected customers via Update Instructions AX006/10/S. The letter informs customers of the potential issues and provides additional information on the implemented resolutions. For additional information, please contact the firm at 888-826-9702.
Z-1194-2010 | Mar-23-2010 | 2010 | Heartstart - Philips Heartstart (HS1) Onsite Automated External Defibrillator (AED) Catalog number M5066A | Philips Medical Systems | 2 | Potential for failure of a capacitor during use which would prevent the AED from delivering effective defibrillation therapy when indicated. | The firm notified their US distributors of the recall and instructed them to contact customers and provide information how to return the product to Philips for a free replacement. Firm contacted distributors via phone and email on 11/2/09 and by letter on 11/9/09. A web site providing information related to the recall was activated on 11/3/09.
Z-2205-2010 | Aug-10-2010 | 2010 | Pulmonetic Systems - Pulmonetic Systems P/N 19189-001 Patient Circuit SPU W/O PEEP 15 FT 5/Pkg.These circuits are intended for use with the MR Conditional LTV 1200 System with the ventilator utilizing an external exhalation valve; as single patient use (SPU) only patient circuits for adult and pediatric patients. The LTV 1200 ventilator is intended to provide continuous or intermittent ventilatory support for the care of individuals who require mechanical ventilation. The ventilator is suitable for use in institutional; home and transport settings; and is applicable for adult and pediatric patients weighing at least 5 kg (11 lbs.); who require the following types of ventilatory support: -Positive Pressure Ventilation; delivered invasively (via ET tube) or non-invasively (via mask). -Assist/Control; SIMV; CPAP or NPPV modes of ventilation. -Breath types including Volume; Pressure Control and Pressure Support. | CareFusion 203; Inc. | 2 | CareFusion is recalling the P/N 19189-001 Patient Circuit (SPU W/O PEEP 15ft) because some of the circuit sense lines may have been reversed during the manufacturing process. If the sense lines have been reversed; immediately upon connection of the patient circuit to the ventilator and turning the ventilator on; the ventilator will begin to autocycle (with increased pressure delivered) with both | Customers were sent a CareFusion "Urgent Medical Device Recall Notification" dated July 2, 1010. The letter was addressed to Respiratory Care Manger. The letter described the product , problem and advised consignees to remove the affected product from inventory and promptly return it to CareFusion for replacement. Please contact CareFusion Technical Support at 800-754-1914, option 2, upon receipt of this notification to arrange for prompt return and replacement of any the patient circuits in your inventory.
Z-1617-2010 | May-19-2010 | 2010 | Newport e360 Ventilator - Newport e360 VentilatorThe e360 Ventilator System is intended to provide continuous (endotracheal or tracheostomy [ET] tube) or non-continuous (mask) ventilatory support and monitoring for infant; pediatric; and adult patients requiring tidal volumes equal to or greater than 20 milliliters (mL). The device is for use by prescription only. The intended environments include hospital; hospital-type; and intra-hospital transport environments. Hospital use typically includes general care floors; operating rooms; special procedure areas; emergency rooms; and intensive and critical care areas within the hospital. Hospital-type use includes facilities such as or similar to surgicenters; sub-acute centers; and special nursing facilities outside of the hospital. Intra-hospital transport includes patient transport within the hospital or hospital-type facility. | Newport Medical Instruments Inc | 2 | The recall was initiated due to a recent increase in complaints related to the early failure of a low pressure transducer which is on Printed Circuit Board (PCB) 2104A in the Newport e360ventilator. Ventilators manufactured between September and December 2009 areaffected by this field correction. All affected ventilators require the replacement of the PCB2104A board to avoid early failure of t | The firm, Newport Medical, sent a "Field Correction Notice" dated April 14, 2010, to all customers. The notice describes the product, problem and actions to be taken by the customers. The customers were instructed to: 1) provide your affected customers with a copy of this letter and advise them to pay close attention to pressure errors which can be an early warning sign of the pressure transducer failure; 2) Schedule their ventilators for replacement of the PCB2104A board as quickly as possible to avoid any unanticipated disruption in their use of their product; 3) Contact our Technical Support Department (800.451.3111 for US customers, 714.427.5811x500 for international customers) to request the replacement kits you will need; and 4) Complete the attached form and return it to Newport Medical as you complete your PCB2104A replacements. This will allow us to ensure that all affected ventilators have been addressed and close this field action. For questions, customers were instructed to contact Newport Medical Technical Support Department (800.451.3111 for US customers, 714.427.5811x500 for international customers).
Z-2421-2010 | Sep-14-2010 | 2010 | CELL-DYN Sapphire Hematology Analyzer - CELL-DYN Sapphire Probe Detector PCB Assembly (also called the Aspiration Bottom Sensor) used on the Cell-Dyn Sapphire Hematology Analyzer; Manufactured by Abbott Laboratories; Santa Clara; CA | Abbott Laboratories | 2 | Under certain conditions in closed mode of operations; the Aspiration Probe Detector PCB assembly may fail to operate properly and provide inaccurate results. | Abbott Diagnostics Division of Santa Clara, CA initiated a field correction of the CELL-DYN Sapphire Probe Detector Assembly used on the CELL-DYN Sapphire Hemotology Analyzer, on April 5, 2010. The correction is being addressed through field service representatives and a mandatory Technical Service Bulletin (TSB). The TSB requires inspection, testing, and potential replacement of the affected assembly. The firm will not be issuing a seperate notification. For questions regarding this correction please contact (408) 982-4800, or fax (408) 982-4863.
Z-2238-2010 | Aug-19-2010 | 2010 | Innova 2000 - Digital Flouroscopic Imaging System - Innova 2000; manufactured by GE Medical Systems; SCS | GE Healthcare; LLC | 2 | A hardware component (potentiometer and encoder)mismatch in Innova 2000 digital fluoroscopic imaging systems may cause a Source to Image Distance (SID) failure; which will cause the imager to fail to detect the distance between the imager and the patient. This will result in loss of the x-ray image if the recovery process is not performed. The error message to use the Emergency Back-Out Button for | GE Healthcare sent a GE "Urgent Medical Device Correction" dated June 9, 2010letter to consignees. The letter was addressed to Hospital Administrators / Risk Mangers; Mangers of Radiology/Cardiology and Radiologists/Cardiologists. The letter described the Safety Issue, Affected Product Details; Safety Instruction; Product Correction and Contact Information. The letter detailed the method for identifying SID failure and the use of the emergency back-out buttoms to recover the X-ray image. A GE service representative will affix a label to the device on-site to remind operators how to recover from loss of X-ray from SID failure. Consignees may contact GE at 800 437-1171, option 4.
Z-1198-2010 | Mar-23-2010 | 2010 | Pnemosure Insufflators - Pnemosure Insufflators; model numbers 620-040-600 and 620-040-610; manufactured by Stryker Endoscopy; San Jose; CA. | Stryker Endoscopy | 2 | Unexpected powering down/failure to boot up: A problem with a transistor in the heater/temperature sensor component in the device may cause the device to power off during surgery or to fail to boot up at all. The risk includes delay and interruption of ongoing surgical procedures. | Recall initiated 12/29/2009. Sales representatives were notified via phone and e-mail, while product advisory notices were sent to consignees via FedEx. Consignees are provided with a response form, and are instructed to run a software correction.
Z-0349-2011 | Nov-12-2010 | 2010 | 19 ViewSonic LCD VG930M Display - ViewSonic 19" flat panel LCD display Model VG930M which is used as part of the Welch Allyn Acuity Central Monitoring System. The affected serial numbers are those less than QC2082880001.The ViewSonic 19" flat panel LCD display Model VG930M is used as part of the Welch Allyn Acuity Central Monitoring System (Acuity System.) The Acuity System consists of a central monitoring station; Acuity System software; and a collection of other commercially available networked products | Welch Allyn Protocol; Inc | 2 | The 19" Model VG930M ViewSonic flat panel LCD display may have faulty capacitors at certain locations in the power supply section. The faulty capacitors may cause a loss of power to the display resulting in the display screen going dark and loss of both visual and secondary patient alarms. | On May 1, 2009, Welch Allyn Protocol, Inc. sent out recall notification to their consignees in via UPS and US Postal Certified Mail. The letter describes the recalled product as 19" model VGM930M ViewSonic flat panel LCD display with serial numbers less than QC2082880001. For the Welch Allyn service contract consignees/customers, the firm's Field Service Engineer will replace the LCD monitor for the consignees. The consignees could contact Technical Support at 1-800-289-2501. For the Warranty/non-Warranty contract consignees/customers-For the Welch Allyn service contract consignees, Stericycle Return Management Services (RMS) will manage all recall and recovery initiatives. Consignees/customers can call Welch Allyn Customer Care at 1-800-535-6663 for any questions or concern about this recall.
Z-0457-2011 | Nov-24-2010 | 2010 | Respironics V60 Ventilator - Respironics V60 Ventilator; manufactured by Respironics California Inc. in Carlsbad; CA. | Respironics California Inc | 2 | Respironics California Inc has received customer complaints of an odor or burning smell coming from the V60 Ventilator; and identified a component on the Power Management PCB that exhibited thermal damage. | Phililps Healthcare issued an Urgent Medical Device Field Correction letter dated July 2, 2010 concerning the RespironicsV60 Ventilator issue. Customers will be contacted by US Field Service Technicians and international distributors to schedule and perform inspections and replacement of the affected component. The unit may continue to be used pending completion of the field inspection. Customers may contact Respironics US Customer Service about this action at 1 800 345 6443, option #5, then option #1.
Z-0625-2011 | Dec-16-2010 | 2010 | LifeVest Wearable Defibrillator - Model 4000 Monitor component of the LifeVest Wearable Defibrillator | Zoll Lifecor Corporation | 2 | The response button function may become intermittent. This can result in the device not progressing beyond the response button test during startup; which may prevent a conscious patient from delaying a treatment shock. | Zoll Life Vest contacted patients that are expected to be in the device more than 60 days by letter and phone to arrange replacement monitors. Instructions for the return of the recalled monitor accompany the replacement monitor The firm will also contact the prescribing physician via telephone for long-term use patients to inform the physician of this voluntary recall and to assist in contacting any patients the firm was unable to reach by telephone. Zoll Life Vest can be contacted at 1-800-543-3267 about this recall.
Z-0034-2011 | Oct-08-2010 | 2010 | Philips IntelliVue Information Center - The device's marketing brochure states: "IntelliVue M3150 Information Center...Philips Medical Systems".The beige-colored PAS-210 external speakers can be identified by their use of DB15-to-3.5 mm mini-phone jack cable (Agispec-081Y-4 or Agispec-081-6) to connect to the rp5700 PC.Indicated for central monitoring of multiple adult; pediatric; and neonatal patients. | Philips Medical Systems North America Co. Phillips | 2 | Potential for Internal Sound Amplifier to malfunction which could delay treatment. | On 11/20/08, Philips began sending the URGENT-Medical Device Correction letter to their consignees via e-mails. On 11/21/08, the firm sent out the recall letter via United State Postal Service. The letter states the affected device as Philips IntelliVue Information Centers, running on the HP rp5700 PC platform equipped with an active speaker solutions using a DB-15 to 3.5 mm cable set. The consignees are advised to ensure that the audible alarm system is clear, free of noise and distortion and adjustable to different sound levels which is explained in p. 5-47 "Adjusting Alarm Tone Volume" of IntelliVue of the Instruction for Use, Release K Manual, supplied with the hardware, Philips Part Number 453564062301. Consignees can call Philips Healthcare Customer Care Service Center at 1-800-722-9377, #3, #1 and reference FCO 99862012.
Z-0166-2011 | Oct-28-2010 | 2010 | Siemens brand 550 TxT Patient Table - Siemens brand 550 TxT Patient Table; Part No. 7346534; Distributed by and/or Manufactured by Siemens AG; Medical Solutions; Germany | Siemens Medical Solutions USA; Inc | 2 | An increase in voltage may cause a short circuit of a power transistor that may cause a smoldering and smoking effect of the electrical wiring and also an unexpected lowering of the Table. Both issues present the potential for injury to a patient or the operator. | Siemens sent a Customer Information Letter on March 31, 2010, to all consignees. The letter identified the product, the problem, and the action that should be taken. Siemens implemented minor modifications to address the issues. No further action was required for the continual operation of the 550TxT Table. For questions regarding this recall call 925-602-8083.
Z-2449-2010 | Sep-22-2010 | 2010 | BIPLANE X-Ray Imaging Systems with Velara Generator - BIPLANE X-ray Imaging Systems included Allura XPER FD10/10; FD20/10; FD20/20; Integris BH5000; and Integris Allura 12-15 Biplane. The BIPLANE system has two C-arms and it allows the operator to image in two planes: Lateral and Frontal. It has a C-arm on the stand and C-arm suspended from the ceiling.The product's brochure is labeled in part: "Philips Healthcare Global Information Center...P.O. Box 1286...5602 BG Eindhoven...The Netherlands". | Philips Medical Systems North America Co. Phillips | 2 | A lock-up might occur inside the Velara Generator due to intermittent shorts inside the system. The problem can only resolved by a warm or cold system restart. In exceptional cases the power converters of the KVmA unit in the generator are damaged and the system can not be restarted. | On 02/04/2010, Philips Healthcare began sending out the URGENT-Device Correction Notices to their consignees. The firm issued two notifications, one for the Allura Integris and Xper FD10C, FD10F, FD10F/C and FD20 Monoplane X-Ray imaging systems with the Velara Generator. The second notification is for the Allura Integris Xper FD10 and FD20 Biplane X-ray imaging systems with the Velara Generator. This notice informs the consignees that a system locks up during a critical interventional case and it might impose a serious injury to patients. This lock up issue could only be resolved by a warm or cold system restart. A warm restart would take about 90 seconds or a cold restart would take about four minutes, and the device could not be used during this time. If the system does not work after the restart, the consignees should contact the Philips Healthcare Customer Service Representative at 1-800-772-9377 #5 (then follow prompts). A service representative will visit each consignee and conduct a software and hardware upgrade.
Z-1238-2010 | Apr-06-2010 | 2010 | Vertier Surgical Table - Vertier Surgical Table; Catalog numbers 0788100000; 0788100000S; 0788100001; and 0788100001S | Stryker Communications Corp | 2 | In certain positions; product is susceptible to fluid instrusion which could possibly result in short circuit or unsolicited movement. | Firm notified all consignees by letter sent 11/13/2009. Contact the firm at 972-410-7100 with questions relating to this recall.
Z-2443-2010 | Sep-21-2010 | 2010 | Oscor PACE 203H; Dual-Chamber External Pulse Generator - Oscor PACE 203H; Dual-Chamber External Pulse GeneratorThe PACE 203H / Model 3085 external pulse generator / temporary cardiac pacemaker is designed to be used with cardiac pacing lead systems for temporary atrial or ventricular or A-V sequential pacing. The PACE 203H / Model 3085 has applications where such pacing modes are indicated for therapeutic; prophylactic; or diagnostic purposes. | Osypka Medical Inc | 2 | The recall was initiated after Osypka Medical received two (2) complaints from their Japanese distributorregarding an incompatibility of the PACE 203H with some new types of 9 V batteries.The investigation revealed that PACE 203H of a certain manufacturing period malfunction(i.e.; not stimulate) when a type of 9 V batteries with higher than usual open circuit voltage (OCV) is used. Failure t | OSYPKA Medical sent "Urgent Medical Device Recall" letters dated June 30, 2010 and July 7, 2010 to the direct customers. The recall notice informed the customers of the Reason for Recall, the Potential Hazard, the Temporary Fix of the Problem, the Permanent Fix of the Problem and Recommended Actions to be taken by the customer. The customers were instructed to contact Osypka Medical via telephone (858) 454-0021 or e-mail to m.osypka@osypkamed.com to verify that the Model 3085 and/or PACE 203H is subject to this product removal, obtain a return-to-manufacturer-authorization (RMA) number ,and then return the affected devices to Osypka Medical. If you have any questions call (858) 454-0021 or e-mail to mail@osypkamed.com.
Z-0589-2011 | Dec-10-2010 | 2010 | GE Responder 2000 Defibrillator/Monitor - GE Responder 2000 Defibrillator/Monitor; models 2025653; 2026109; and 2026114 that were manufactured between July 25; 2008 and October 2; 2009 by GE Healthcare. | Cardiac Science Corporation | 2 | Some units may be non-operational during a rescue; due to a defective PCB interconnect on traces connected to the reed relay. This will result in an inability to analyze the patient's heart rhythm or deliver therapy during a rescue. | Cardiac Science issued a Urgent Medical Device Corrective Action letter dated April 28, 2010 requesting GE Healthcare cease distribution of affected devices and to instruct end users of distributed devices to schedule a service call to have the device corrected. Cardiac will send rework kits no later than June 30, 2010. CSC can be contacted at 44 161 926 0011 to arrange return material authorization and replacement of affected devices.
Z-1732-2010 | May-28-2010 | 2010 | Axiom Artis systems with Motor Controller Unit - Axiom Artis systems with Motor Controller Unit. Siemens Medical Solutions USA; Inc. Malvern; PA 19355.Intended for x-ray; angiographic use. | Siemens Medical Solutions USA; Inc | 2 | sporadic malfunctions of the Artis Systems with Motor Controller caused by tolerances on electronic parts that might increase during the lifetime of the system. | Siemens issued an "Urgent Field Safety Notice" to consignees describing the affected device instructing users to contact the recalling firm's service organization to schedule maintenance. For further information, contact Siemens Medical Solutions USA, Inc. at 1-610-448-3237.
Z-1156-2010 | Mar-17-2010 | 2010 | CircuCool Fluid Circulation Pump - CircuCool Fluid Circulation Pump used in the Chilli Cooled Ablation System; Model Number M00480050; manufactured by Barnant ThermoFisher Scientific; Barrington; IL. Distributed by Boston Scientific Corporation; San Jose; CA.The Chilli Cooled Ablation System is indicated for cardiac electrophysiological mapping; delivering diagnostic pacing stimuli and radiofrequency ablation of mappable ventricular tachycardias attributable to ischemic heart disease or cardiomyopathy in patients who have failed drug therapy. | Boston Scientific Corp | 2 | Product failure-- Product may cease operation as a result of potential speed control board failure when the device is shut off; and may not turn back on. A component on the circuit board fails unexpectedly because of a power surge; causing a diode to burn out. | The firm has issued notification letters via Federal express with reply verification tracking forms. A follow up letter and telephone survey as needed will be sent to follow up with non-responders. The firm is asking its consignees to return any recalled product. If there are any questions, please call 1.408.935.4952.
Z-0548-2011 | Dec-07-2010 | 2010 | ALLURA XPER FD20 - ALLURA XPER FD20The intended uses of the Allura XPER X-ray systems involved with this report are angiographic X-ray systems with a solid state x-ray imaging device for cardiovascular and vascular diagnostic and interventional procedures. | Philips Medical Systems North America Co. Phillips | 2 | The Printed Circuit Board (PCB) in the Main-cabinet. Further investigation revealed that a manufacturing error in the PCB is responsible for the reliability issue. As a result; the complete system may shut down and not re-start. | Philips began mailing out an URGENT-Device Correction letter on May 26, 2009. Philips sent out a second URGENT-Device Correction letter dated June 4, 2009, to all consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were advised to contact their local Philips representative immediately if their system completely shut down and would not restart. A mandatory Field Change Order 72200144 was issued for replacement of the Printed Circuit Board (PCB) in the Allura XPER FD system. A mandatory Field Change Order FCO 70800091 was issued for replacement of the Printed Circuit Board (PCB) in the MultiDiagnost Eleva with FD system. The consignees were to be contacted by their Philips Representative for implementation of the corrective action. The firm's Field Engineer would replace the power tray of the systems. If the consignees needed any further information or support concerning this issue, they were to contact the call Center at 800-722-9377, # 5 and reference FCO 72200144 or FCO 70800091 with any questions.
Z-1757-2010 | Jun-04-2010 | 2010 | Treadmill models TM55; TM65; and CR60 (brand names include Quinton; Burdick; and Cardiac Science). - Treadmill models TM55; TM65; and CR60 (brand names include Quinton; Burdick; and Cardiac Science). | Cardiac Science Corporation | 2 | 28 stress and rehab treadmills manufactured between September 2006 and January 2009 may not have received adequate electrical safety testing prior to being shipped from the factory. If these components break down in the field; a patient or operator could be exposed to excess touch currents on the surface of the product. Because safety testing may not have allowed for complete evaluation of the el | On April 3, 2009 the recalling firms notified consignees by certified letter to domestic customers and via FedEx to international customers with delivery confirmation signature required. Customers were advised to remove treadmill from service until it is tested by a Field Service Engineer. For customers with questions or those that have not already scheduled a service call with Cardiac Science Field Service Engineers, please contact the firm at 1.800.426.0337, option 1.
Z-1899-2010 | Jun-24-2010 | 2010 | Horizon Cardiology Hemo Monitoring System - Horizon Cardiology Hemo Monitoring SystemA programmable diagnostic computer intended for acquiring digitizing; storing; displaying ; and reviewing hemodynamic data for use in hospital cardiac catherization laboratories. | McKesson Provider Technologies - Medical Imaging Group | 2 | McKesson has identified a potential hardware configuration problem in some Horizon Cardiology Hemo systems which may cause delay and/or loss of the patient's physiological parameters on the Horizon Cardiology Hemo screen and on the patient record. | The firm, McKesson Israel Ltd., sent notification letter dated January 27, 2010 to customers. The letter describes the product, problem and action to be taken by customers and the firm. McKesson's will be contacting the customer to coordinate replacement of the sound component of the Horizon Cardiology Hemo system and will provide the additional sound adapter for all affected systems at your facility at not cost. The customers were instructed to pleasel immediately conduct the following instructions to disable sound operations-1) Open McKesson Management Console and locate the Windsurfer>Monitor Settings; 2) Uncheck "Enable QRS beep" in the sounds section; 3) Locate Windsufer>Hardware Settings in McKesson Management Console, and 4) Uncheck "Enable QRS beep" in the Analog Output, QRS Beep Section. If you have any questions, please do not hesitate to contact Ronen Gans directly at (973) 355-9900.
Z-1424-2010 | Apr-27-2010 | 2010 | Powerheart; CardioVive; Nihon Kohden and/or Responder AED - Cardiac Science; Automated External Defibrillator (AED) under the following brand names (with respective model numbers): Powerheart AED G3: 9300A; 9300C; 9300D; 9300E; 9300P; 9390A/E; CardioVive: 92532 and 92533; Nihon Kohden: 9200G; and 9231; and Responder: 2019198 and 2023440. Cardiac Science; Bothell; WA 98021.Indicated for emergency treatment of victims exhibiting symptoms of sudden cardiac arrest who are unresponsive and not breathing. | Cardiac Science Corporation | 1 | Certain models of Powerheart G3 AEDs may not have received adequate electrical safety testing prior to being shipped from the factory. Affected models are: Powerheart 9300A; 9300C; 9300D; 9300E; 9300P; 9390A; 9390E; CardioVive 92532; 92533; NK 9200G; 9231; and Responder 2019198 and 2023440. The affected devices were manufactured between October 29; 2003 and February 10; 2009. | Cardiac Science issued "Urgent - Medical Device Corrective Action" notices dated October 2009. Customers were instructed to identify the affected product and arrange for product replacement with the firm. Where possible, end user information was requested for further follow-up by the firm. For further information, contact Cardiac Science at 1-877-402-4599 or by email at AED155@cardiacscience. com.
Z-1527-2010 | May-11-2010 | 2010 | GE Avance - GE; 1009-9002-000 Avance Anesthesia Machine and Monitor | Ge Healthcare; Llc | 1 | Unanticipated shut-down: A specific lot of GE Healthcare Aisys and Avance machines control board wiring harness have a defect; which can cause the machine to unexpectedly shut down; terminating ventilation; anesthetic delivery; and potentially patient monitoring. | Consignees were sent a GE Healthcare "Urgent Medical Device Correction" letter dated March 12, 2010. The letter was addressed to Chief of Anesthesia, Health Care Administrator/Risk Manager, Director of Biomedical/Clinical Engineering. The letter listed the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information.
Z-2625-2010 | Sep-29-2010 | 2010 | Philips M3002A IntelliVue X2 - Philips M3002A IntelliVue X2Model: 865039Indicated for use by health care professionals whenever there is a need for monitoring the physiological parameters of patients. Intended for monitoring; recording; and alarming of multiple physiological parameters of adults; pediatrics; and neonates in a hospital environment. ST segment monitoring is restricted to adult patients only. The M3002A and M8102A are used for monitoring; recording and alarming of multiple physiological parameters of adults; pediatrics; and neonates in health care facilities | Philips Healthcare Inc. | 2 | Speakers on the Philips IntelliVue X2 and MP2 patient monitors may fail; causing absence of an audible alarm and delaying patient treatment. | Philips issued a URGENT - MEDICAL DEVICE letter dated August 26, 2010, to each affected customer. The Urgent Medical Device Correction notification/Field Safety Notice informs customers of the issue and gives instructions on actions to take while awaiting their replacement speaker assemblies. The speaker replacement assembly is expected to be available in 8 weeks. Affected products may continue to e used while awaiting replacement of the speaker, provided that the user routinely checks for the display of the "Speaker Malfunct." INOP at power-up. If you experience this INOP or there is no sound from your X2/MP2 speaker while it is not connected to another audible alarm annunciation system, remove the device from use and contact your local Philips service representative at 1-800-722-9377.
Z-0024-2011 | Oct-08-2010 | 2010 | Sarns Ultrasonic Air Sensor - Air bubble detection cable assembly; catalog numbers 78-8067-7375-6 and 149892; Terumo Cardiovascular System; Ann Arbor; MI | Terumo Cardiovascular Systems Corporation | 2 | Malfunction of the air bubble detection cable assembly causes a false air alarm or a check sensor message during system set-up. In some instances; users were not able to reset the alarm to allow for further air bubble detection. . | The firm sent URGENT MEDICAL DEVICE CORRECTION notices dated September 17, 2010. The notices told consignees that a Terumo representative would visit, inspect the cables and replace the recalled cables. if the customer has a service contract, the visits would occur at the next scheduled preventative maintenance visit. Customers without a service contract will be contacted by Terumo representative to schedule an appointment. Customers can contact Terumo Customer Service at 800-521-2818.
Z-1135-2010 | Mar-15-2010 | 2010 | Philips HeartStart - Philips Switched Internal Defibrillator Paddles (used with Philips HeartStart XL and HeartStart MRx monitor/defibrillators) used during open heart surgery. Models:M4741A - Extra Large Switched internal PaddlesM4742A - Large Switched internal PaddlesM4743A - Medium Switched internal PaddlesM4744A - Small Switched internal Paddles | Philips Healthcare Inc. | 2 | Paddle shock switch may fail to actuate and delay or prevent delivery of defibrillation therapy | Philips issued Urgent Medical Device Recall notification to customers by letter on 1/25/10 via UPS advising of this defect and providing instructions on how to determine if they have a defective paddle. The customers will be assisted by field service engineers to remove and replace these defective units (those within the date range of September 1, 2008 to December 30, 2008). Customers are asked to follow the Actions to be taken by customer/user section of the Urgent Medical Device Recall Notification. Removed units will be replaced free of charge and will be destroyed by Philips. Distribution of the Urgent Medical Device Recall Notification will be managed by Philips representatives in each affected geography. Customers are asked to follow the Actions to be taken by customer/user section of the Urgent Medical Device Recall Notification.
Z-0141-2011 | Oct-25-2010 | 2010 | BV Endura R2.2 Mobile X-ray system - Philips BV Endura R2.2 Mobile X-ray system. The Type Product Code 718074 can be found under the X-ray trolley cover on the front lower right side of the module or under the stand cover. The device brochure is labeled in part: "Philips...Mobile C-Arm system BV Endura". | Philips Medical Systems North America Co. Phillips | 2 | Reliability of the Stand-Trolley cable may result in loss of images on the monitor during an examination. This intermittent problem can lead to System lock-up; where no radiation is possible and images are not visible on the monitor | On 10/04/2008, Philips began sending the URGENT-Device Correction BV Pulsera R2.2 and Endura R2.2 letters to their customers. On 09/03/2008, the firm issued the Field Change Order (FCO) 71800024 and the firm's representative will replace the stand-trolley cable free of charge. Customers should contact Philips Call Center at 1-800-722-9377, #5, #2, and reference "FCO 71800024" with any questions.
Z-2293-2010 | Aug-27-2010 | 2010 | HeartStart MRx Defibrillator/Monitor with Q-CPRTM - HeartStart MRx Defibrillator/Monitor with Q-CPR Meter option.The HeartStart MRx monitor/defibrillator has an option (Q-CPR Meter) to provide visual and audible feedback to the rescuer on the quality of CPR. The Q-CPR option offers real time measurement and corrective feedback on the rate; depth and complete release of compressions; ventilations rate and lack of CPR activity in accordance with current CPR guidelines. The Q-CPR option is contraindicated as follows: *The Q-CPR option is contraindicated for use on neonatal and pediatric patients (under 8 years of age or weighing less than 25 kg). *The Q-CPR option is not for use when CPR is contraindicated. | Philips Healthcare Inc. | 2 | Reduced sensitivity in the force measurement sensor can result in inaccurate feedback to the user on how well he or she is releasing pressure after each chest compression. This can result in incomplete release of chest compression; which will reduce CPR effectiveness. | Philips issued an "Urgent Medical Device Recall notification/Field Safety Notice" dated April 2010 to all consignees. The letter described the issue, how to identify affected devices, and required actions. The consignees were instructed to remove affected devices from service. Philips will contact consignees to arrange on-site replacement of affected meters with meters containing revised labeling and calibration software. The new meters will indicate that meter recalibration is required after approximately 50,000 compression cycles. Consignees returning devices to Philips for routine recalibration will receive free replacement. All CPR meters will be replaced with re-designed meters when they become available. For questions or assistance in notifying accounts about this recall, contact Philips at 1-800-722-9377.
Z-0006-2011 | Oct-13-2010 | 2010 | Alaris - CareFusion Alaris PC Unit Model 8015 (formerly MedleyPC Unit): Alaris PC Units (Model 8015) with logic board part number TC10005522 manufactured and serviced between December20; 2008 and May 6; 2009; and Alaris PC Units (with logic board part number TC10005572 manufactured and serviced between May 7; 2009 and September 8; 2009. For use with Alaris infusion and monitoring systems. | CareFusion Corporation | 1 | Under certain wireless network conditions a communication error can occur; which freezes the PC Unit screen; which may result in a delay of therapy. A delay of therapy may result in serious injury and/or death. CareFusion identified this potential risk associated with the Alaris PC unit model 8015 only (PC Unit) manufactured between December-2008 and September-2009. | CareFusion sent the customer notification letter to all customers using the affected Alaris PC Unit beginning August 24, 2010. The customer notification letter is addressed to the Directors of Nursing, Risk Management, and Biomedical Engineering. Customers will be required to confirm receipt of the notification by returning the Recall Response Card to CareFusion by postage-paid, self-addressed mail, fax or email. The recall notice informed the customers of the reason for the notice, the units affected, the potential risk, a step-by-step description of the issue and the required actions for users. The customers were also told that CareFusion will contact their facility through phone or in person within 60 days to initiate implementation of the corrective action. The corrective action will require a hardware update to all affected units. CareFusion does not require that you return your devices. In the interim, if you observe this issue, remove the device from service and contact CareFusion Recall Center immediately. Customers were also told that any adverse reactions experienced with the use of this product, and/or quality problems should also be reported to the FDA's MedWatch Program. The recall notice provided contact information for CareFusion's Recall Center, Customer Advocacy and Technical Support units.
Z-2176-2010 | Aug-13-2010 | 2010 | Ikaria INOMAX DS Drug Delivery System - Ikaria INOMAX DS Drug Delivery System; model 10003. Ikaria 6 Route 173; Clinton; NJ 08809 USA.The INOMAX DS Delivery System delivers INOMAX (nitric oxide for inhalation) therapy gas into the inspiratory limb of the patient breathing circuit in a way that provides a constant concentration of nitric oxide (NO); as set by the user; to the patient throughout the inspired breath. It uses a specially designed injector module; which enables tracking of the ventilator waveforms and the delivery of a synchronized and proportional dose of NO. It may be used with most ventilators. The INOMAX DS provides continuous integrated monitoring of inspired O2; NO2; and NO; and a comprehensive alarm system. | Ikaria Holdings | 1 | The pressure switch in the INOMAX DS drug-delivery system may fail; which may interrupt or delay the administration of INOMAX (nitric oxide) for inhalation to patients. Delay in administration of INOMAX therapy may cause: -worsening of systemic oxygenation indices (i.e.; hypoxemia; especially manifested as decreased arterial oxygenation saturation). -hypotension; and/or -increase in pulm | Ikaria sent an Urgent Medical Device Recall letter to customers dated July 21, 2010. This letter is also posted on the Inomax webpage. The letter described the issue, how to recognize it, what to do in the event of a suspected leak, and recommended maintaining back-up cylinders and depressurization of the regulator when not in use. The replacement plan for all INOMAX DS drug-delivery systems was described. Customers can contact Ikaria Customer Care at 1-877-566-9466.
Z-0143-2011 | Oct-26-2010 | 2010 | Scanpoint Communication Cradles - The label (which is affixed to ScanPoint Communications Cradle) states: "...VERATHON ScanPoint Docking Station...Part No. 0570-0168...Bothell; WA USA..." | Verathon; Inc. | 3 | Incorrect programming causes QuickPrint to fail to recognize the communication Cradles. As a result; ScanPoint Communications Cradles could not communicate with ScanPoint Online Tool while using the QuickPrint interface. This will affect the transfer of exams and device calibration through QuickPrint interface. | On 09/20/10, Verathon issued an URGENT: Medical Device Accessory Recall Notification letter dated September 20, 2010 to consignees, identifying the affected product, the issue prompting the recall, and actions to be taken by consignees. The consignees are instructed to return any cradles within the serial number range between CC012000-CC012084, and the Part Number 0570-0168. The firm will provide a replacement communication cradle at no cost. Consignees can contact Verathon Customer Care at 425-867-1348 or 800-2331-2313 or send e-mail to cservice@verathon.com.
Z-1788-2010 | Jun-11-2010 | 2010 | Triage Total Controls 5; Level 1 - Triage Total Controls 5; Level 1: Five 0.25 mL tubes. Storage Temp <-20C. Intended use: Assayed materials used with the Triage Profiler S.O.B. Panel; Triage CardioProfilER Panel; Triage BNP Test; Triage D-Dimer Test and Triage Meters to assist in monitoring test performance. | Biosite Inc | 2 | Incidence of low recovery results has exceeded post market surveillance thresholds of customer complaints in the field. | The firm, Biosite Incorporated, sent an "URGENT MEDICAL DEVICE RECALL" dated February 6, 2009, to all customers. The letter describes the product, problem and action to be taken by the customers. The customers were instructed to immediately discontinue all use and/or sale of the listed products, discard all affected product in accordance with their local regulations, complete and fax the enclosed Verification Form within 10 days to confirm their receipt of the notice, to indicate the number of kits discarded from their inventory and the product replacement will be sent to them. Contact Inverness Medical Technical Service at 1-877-441-7440 and select option 2 from the phone prompt and a representative will confirm the number of kits discarded from your inventory and arrange for replacement product to be sent to you. Should you have any questions about the information contained in this notification, please contact (877) 441-7440 or email: Recalls@biosite.com.
Z-1995-2010 | Jul-12-2010 | 2010 | Access AccuTnI Assay - Access Immunoassay Systems enhanced AccuTnl Reagent Kit; Part Number: A78803The Access AccuTnl assay is a paramagnetic particle; chemiluminescent immunoassay for the quantitative determination of cardiac troponin I (cTnI) levels in human serum and plasma using the Access Immunoassay Systems to aid in the diagnosis and treatment of myocardial infarction and cardiac muscle damage. Cardiac Troponin I determination aids in the risk stratification of patients with unstable angina or non-ST segment elevation acute coronary syndromes with respect to relative risk of mortality; myocardial infarction; or increased probability of ischemic events requiring urgent revascularization procedures. | Beckman Coulter Inc. | 2 | The recall was initiated because as a follow up to PCA # 12605 dated December 16; 2009; Beckman Coulter has confirmed additional reports that each of the following assays may produce lower than expected results when run in conjunction with Access AccuTnl (REF A78803):(1) Access Testosterone (REF 33560);(2) Access Thyroid Uptake (REF 33810);(3) Access Total T4 (REF 33800).This issue can occ | The firm, Beckman Coulter, issued a "URGENT: PRODUCT CORRECTIVE ACTION" (PCA) letter with attached Customer Response form dated March 24, 2010 to all customers. The letter describes the issue, impact, action and resolution. In the PCA beckman Coulter recommends the following mitigation to reduce the risk of generating erroneous results: (1) Verify the APF software on the immunoassay system. If the updated APF software listed in the table above was already loaded , skip the remainder of this section and proceed to the Resolution section; (2) If the updated APF software has not yet been loaded, but is available, load the software appropriate for the system, skip the remainder of this section and proceed to the Resolution section; (3) If the updated APF software is not immediately available, please contact Technical Support at 1-800-854-3633 in the United States and Canada. Consignees outside the United States and Canada, contact their local Beckman Coulter Representative and follow the instructions below until the software has been loaded on their system. (a) Access and Access 2 users should batch process Access Testosterone, Thyroid Uptake, or Total T4 samples; (b) For UniCel DxI users, assign Access AccuTnI (REF A78803) to a different reagent pipettor(s) than Access Testosterone, Thyroid Uptake, or Total T4; (c) The Access Testosterone, Thyroid Uptake, or Total T4 results should be interpreted in light of the total clinical presentation of the patient, including: symptoms, clinical history, data from additional tests, and other appropriate information. Abnormally low Access Testosterone, Thyroid Uptake, or Total T4 results from samples tested in conjunction with Access AccuTnI (REF A78803) should be reviewed in the context of other diagnostic tests. Customers were instructed that after loading the updated APF software, the Access AccuTnI assay (REF A78803, test code TropI, test ID 275) must be recalibrated. Recalibrate of the Access Testosterone, T.
Z-1631-2010 | May-20-2010 | 2010 | LATITUDE¿ - Boston Scientific; LATITUDE¿ Communicator; Distributed by Boston Scientific; 4100 Hamline Avenue North; St. Paul; MN 55112. Models 6476 (USA) and 6468 (OUS) | Boston Scientific CRM Corp | 2 | Boston Scientific has determined that alert conditions for a limited subset of out-of-range Daily Measurements related to leads may not be recognized by Model 6476 (United States) or Model 6468 (Europe) LATITUDE in-home communicators. If a communicator fails to recognize a valid alert situation; notification for this or subsequent alerts for the same test will not be displayed for physician review | Consignees were sent on 9/10/09 a Boston Scientific "Urgent Medical Device Information" letter dated September 9, 2009. The letter was addressed to "Dear Doctor". The letter provided the information related to the LATITUDE Patient Management system recall as to Background, Description, Rate of Occurrence, Patient management, Devices Affected and Further Information. Patients were also sent a Boston Scientific letter dated September 9, 2009. The letter addressed Background information, Why are we communicating, What You Should Do and Questions.
Z-1591-2010 | May-10-2010 | 2010 | VNS Therapy System - VNS Therapy¿ Demipulse¿ Generator; Model 103 and VNS Therapy Demipulse Duo¿ Generator; Model 104VNS Therapy Programming System; Model 250 Version 7.0 and 7.1 Software | Cyberonics; Inc | 2 | Battery life projection is inaccurate. | Safety Alert sent April 6, 2010 to all physicians of record for Model 103 or Model 104 generators and to all hospitals with non-implanted affected product, that explains the event, risks presented and recommended actions to take. The generator can be reprogrammed by the physician the next time the patient is seen in the clinic. The issue can also be corrected during the initial implant procedure with the instructions provided. Additional questions should be directed to the firm; United States customers can contact Clinical Technical Support at 866-882-8804.
Z-0631-2010 | Jan-13-2010 | 2010 | GE ApexPro Oximeter - GE ApexPro Oximeter systems (ApexPro CH or ApexPro FH with ApexPro Telemetry version 3.9 software) | General Electric Medical Systems Information Technology | 2 | Potential misuse of the ApexPro Telemetry System and incorrect message and alarm information in the ApexPro Operator's Manual may impact patient safety. When ApexPro Telemetry systems are used with SpO2 oximeters:1. ApexPro FH Transceiver can operate with the ApexPro FH Xpod connected; but without ECG lead wires attached to the transceiver. This is unintended use of the device that is not do | Consignees were sent on 10/27/09 a GE Healthcare "Urgent Medical Device Correction" letter dated September 25, 2009. The letter was addressed to Risk Manger, Chief of Nursing, director of Biomedical Engineering. The letter described the Safety Issues, Affected Product Details, Product Correction, Safety Instructions and Contact Information. "ApexPro Telemetry System Operator's Manual Supplement for SpO2 was enclosed with the letter.
Z-0650-2010 | Jan-15-2010 | 2010 | Smiths Medical Medfusion infusion pump - Smiths Medical Medfusion model 3500 syringe Infusion pump; software version V3.0.6. Model 3500. Rx Only. Smiths Medical MD; Inc.; St. Paul; Minnesota 55112 (USA). | Smiths Medical ASD; Inc. | 2 | Smiths Medical has become aware of an increased trend in reports of Motor Not Running (MNR) and Motor Rate Error (MRE) alarm message events with certain Medfusion? Syringe Infusion Pumps; Models 3010A and 3500 (Pump). MNR/ MRE are high priority alarm messages which can occur for a variety of reasons as part of the usual self-test functionality of the Pump. The Pump is designed so that when its so | An Urgent Medical Device Recall letter, dated November 16, 2009, was sent to customers beginning November 16, 2009. The letter described the issue, identified affected devices, and asked that affected devices be returned to Smiths Medical to be inspected, repaired, if necessary, and returned back to medical facility. A Confirmation Form is also asked to be completed and returned to Smiths Medical. To return Pumps or for questions regarding this Urgent Recall Notice, please contact Smiths Medical Customer Service Department at 1~866-270-7060. Please report any issues with these Pumps please contact Smiths Medical Technical Customer Service Department at 1866-216-8806.
Z-1246-2010 | Apr-06-2010 | 2010 | Powerheart AED - Powerheart 9390A automated external defibrillator.This fully automatic model does not require the user to press a shock button in order for the device to deliver therapy (when appropriate).Brand name: Cardiac Science.Product name: Powerheart.Device Operation: Fully automatic.Color: Yellow.Voice Prompt Level: Extensive.(Extensive voice prompts provide more detailed instruction on device use; pad placement and CPR performance; as compared to the basic prompts.)CPR Metronome: Yes.Rescue Ready Indicator: Yes.Text Display: Yes.ECG Display: No. | Cardiac Science Corporation | 2 | A component of the AED device may fail during a rescue attempt and the AED will not be able to deliver therapy. | The firm began initial notification on 2/3/2010 by sending letters to all affected consignees via USPS certified mail (domestic) and FedEx (international). Also on 2/3/2010 the firm issued a global press release and posted recall information on their website. Customers are instructed to remove their affected units from service, and have been given instructions for returning their affected devices. Replacement AED(s) will be sent as they are available. For questions, please contact the firm's Cardiac Science Support at 1.888.402.2484 or. AED195@cardiacscience.com to arrange for replacement(s).
Z-0648-2010 | Jan-19-2010 | 2010 | Ultraview SL 2400 - Spacelabs Model 91369-24 SL 2400 Compact MonitorSpacelabs Medical patient monitors passively display data generated from Spacelabs Medical patient parameter modules; Flexport interfaces; gas analyzers; and other SDLC based products. | Spacelabs Healthcare; Incorporated | 2 | Potential for monitor to shut down and not being able to power on again. | On 11/6/09 the firm sent Urgent - Device Correction notifications to customers by certified mail informing them of the monitor failure and that a field service engineer will contact the customer to fix the affected device at no cost. On 11/13/2009 the firm's Global Tech Support sent notification via email to all foreign subsidiaries and distributors informing them of the monitor failure and that a field service engineer will contact the customer to fix the affected device at no cost. Questions are to be directed to Spacelabs Medical's Technical Support at 1-800-522-7025, and 1-425-657-7200 x5089 for those customers outside of the US.
Z-0369-2011 | Nov-15-2010 | 2010 | GE BrightSpeed CT - GE BrightSpeed Excel/Edge/Elite Select / Elite CT Scanner System; models 2377708-560; 5377708-540; 5377708-560. Composed of a gantry; patient table; operator console; computer; and PDU and includes image acquisition hardware; image acquisition and reconstruction software; associated accessories and connections/interfaces to accessories. Is intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes; including Axial; Cine; Helical (Volumetric); Cardiac; and Gated acquisitions. | GE Healthcare; LLC | 2 | GE has become aware of a potential set of circumstances that could cause x-ray continuation during an unexpected table stop on certain LightSpeed VCT and BrightSpeed systems. | An Urgent Medical Device Correction letter dated September 2, 2010 was sent to consignees beginning September 3, 2010. This letter addressed the LightSpeed VCT systems. Another Urgent Medical Device Correction letter, dated September 24, 2010 was sent September 24, 2010 for the BrightSpeed Systems. The letters described the issue and provided safety instructions for the health care provider to be aware if the table should stop unexpectedly and the x-ray remains on to manually stop the scan until the devices software is upgraded. If you have any questions, please contact the US (800) 437-1171, Japan 0120-055-919, Korea 1544-6119, Australia/New Zealand 800 659 465, China 800-810-8188; for other countries, please contact your local GE Healthcare Services Representative.
Z-2275-2010 | Aug-25-2010 | 2010 | CARESCAPE MONITOR B850 - CARESCAPE MONITOR B850Usage: The CARESCAPE TM Monitor B850 is a multi-parameter high acuity patient monitor intended for use in multiple areas within a professional healthcare facility. | GE Healthcare; LLC | 2 | GE Healthcare has become aware of four potential safety issues associated with the CARESCAPE" Monitor B850:1. Loss of user input when using certain displays or the USB remote control.2. Potential for a delay in treatment or missed alarm when the alarm light is used as a primary alarm source and audible alarms are turned down or off.3. The CARESCAPE Monitor B850 will not use user entered hemo | Consignees were sent on 7/16/10, a GE "Urgent Medical Device Correction" letter dated July 12, 2010. The letter was addressed to Healthcare Administrator/Risk Manger, Chief of Nursing, Director of Biomedical Engineering. They stated the Safety Issues, Affected Product Details, Product Correction, Safety Instructions and Contact Information. If you have any questions, please contact Technical Support at 1-800-558-7044.
Z-2476-2010 | Sep-22-2010 | 2010 | Allura XPER FD20 - Philips Allura XPER FD20 Angiographic X-ray system. This system with Velara Generator can be identified as type number 9890 000 PRODUCTS 6501 * (number located on the generator cabinet).Intended use for Angiographic X-ray system with and without a solid state X-ray imaging device for cardiovascular and vascular diagnostic and interventional procedures. | Philips Medical Systems North America Co. Phillips | 2 | A high voltage reliability issue resulting in the loss of the x-ray generation and this can cause an intermittent stop of fluoro acquisition and exposure during an examination without warning and hence no image on the display. | The firm, Philips, sent "URGENT-Device Correction" letters dated November 10, 2008, to their customers and informed them about the reliability issue in the Allura XPER and Integris X-ray systems with the High Voltage Velera (HV) Generator. This reliability issue can lead to a system error causing intermittent loss of fluro and/or exposure acquisition, potentially resulting in the loss of the X-ray image on the display. The customers are advised if they experience High Voltage Generator reliability issue and the HV Generator stops working generating an O3EA error, a cold restart is required and reboot the system when the X-Ray image disappears. The firm will implement multiple mandatory Field Change Orders (FCOs) to solve the Generator reliability by updating the software for each X-ray system. The customers could contact Philips Care Center at 800-722-9377, #5, #4, #1 for any questions regarding this recall or 425-487-7469.
Z-1551-2010 | May-07-2010 | 2010 | Peloris Rapid Tissue Processor - Peloris Rapid Tissue Processor; Leica Biosystems Melbourne Pty Ltd; Mount Waverly; 3149 VIC; Australia; catalog number 26.0001 and 26.0005 | Leica Microsystems; Inc. | 2 | Confusion during manually changing reagents; general users running unvalidated protocols; and wax valve leakage have contributed to tissue damage during processing; leading to rebiopsies. | Leica Microsystems sent Medical Device Correction letters dated 2/9/10 to the affected customers via first class mail on the same date, informing them of the potential for tissue damage during processing and outlining the software upgrade being made to minimize the risk of tissue damage. Customers were informed that their local Leica Microsystems representative will contact them to arrange the software upgrade and to provide training on the use of the new software. The letters also provided the necessary steps required to avoid tissue damage during processing while they are awaiting the software upgrade. The accounts were asked to complete the attached Acknowledgement Form, listing the serial numbers of the units at their facility and fax it back to Leica at 1-847-607-3147. Additional questions should be addressed to the firm's office of Regulatory Affairs and Quality Assurance.
Z-0871-2010 | Mar-01-2010 | 2010 | PTS Panels Cholesterol test strips - PTS Panels Cholesterol test strips for self-testing or professional use with CardioChek analyzers; Manufactured by Polymer Technology Systems; Inc.; Indianapolis; IN.Catalog # 1711.Intended to provide a quantitative measurement of total cholesterol in whole blood. | Polymer Technology Systems; Inc. | 2 | The memory chip was coded with incorrect product information. | Polymer Technology Systems, Inc. issued an "Urgent - Field Correction" letter dated April 2, 2009. Consignees were advised of the problem and asked to identify how many kits were in inventory so that the firm could replace product. For further information, contact Polymer Technology Systems, Inc. Customer Service at 1-877-870-5610.
Z-0393-2011 | Nov-16-2010 | 2010 | Innova 2000 - Innova 2000 The principle system components include a C-arm; image acquisition; processing and archiving capabilities.The Innova systems are indicated for use in generating fluoroscopic images of human anatomy for diagnostic and interventional cardiac angiography procedures. These system can be operated in a mobile or fixed site environment. | GE Healthcare; LLC | 2 | GE Healthcare has become aware of a potential video splitter failure with certain Innova 2000 and 4100 X-ray systems that may impact patient safety.A potential video splitter failure inside the GE equipment cabinet could occur with no advance warning and thereby prevent x-ray image display. Should this occur; the systemcannot be recovered without service intervention.GE Healthcare has obse | GE Healthcare sent an " Urgent Medical Device Correction" letter dated July 9, 2009, to Hospital Administrators / Risk mangers, Managers of Radiology / Cardiology and Radiologists / Cardiologists. The letter identified the product, described the Safety Issue, Affected Product Details, Safety Instructions Product correction and Contact Information. Hospitals were instructed to establish emergency workflow procedures should a component stop operating during a patient examination. A GE Healthcare service representative would contact each customer and correct all affected systems. For questions regarding this recall call (262) 548-2731.
Z-0714-2011 | Dec-16-2010 | 2010 | CyberKnife Robotic Radiosurgery System - CyberKnife Robotic Radiosurgery System; medical charged-particle radiation therapy system; manufactured by Accuray Inc.; Sunnyvale; CA.Indicated for treatment planning and image-guided stereotactic radiosurgery and precision radiotherapy of lesions; tumors and conditions anywhere in the body when radiation treatment is indicated. | Accuray Inc | 2 | If users modify the dose calculation box to cover just the contoured anatomy and not the entire CT volume; dosage may not be properly calculated. | Accuray Inc., sent an "URGENT DEVICE CORRECTION" letter dated October 21, 2010, to all customers. The letter described the product, problem and actions to be taken by the customers. The customers were instructed to increase the size of the dose box to include CT volume, not just the contoured anatomy; recalculate dose at high resolution, and examine the treatment plan to ensure that applicable organs at risk are accounted for and all portions of the patient receive the desired level of doses. If you have any questions or concerns regarding this issue, please do not hesitate to contact Accuray Customer Support at +1-877-668-8667 (USA), +1-408-716-4700 (non-USA) or email: customersupport@accuray.com.
Z-0847-2010 | Mar-23-2010 | 2010 | DigitalDiagnost - DigitalDiagnost. General radiographic examinations and applications wherever a solid state X-ray imaging device can be used. | Philips Medical Systems North America Co. Phillips | 2 | Device failed to display appropriate labels indicating Radiation Symbol along with the required wording. | Urgent - Field Safety Notice letters, dated January 20 & 21, 2010, were sent to customers. The letter identified the affected products, gave a description of the problem & the hazard involved, stated how to identify affected products, actions to be taken by the customer or user, and actions planned by the firm. Customers are to make sure that the relevant safety precautions are observed and check their operators console (EEC) if the labeling is present as shown in the drawing. If the labeling is missing, a Philips rep should be contacted. For further information or support concerning the issue, a Philips rep should be contacted at 800-722-9377.
Z-2241-2010 | Aug-20-2010 | 2010 | Dental chair headrest - The magnetic headrest is designed to fit the following dental chairs (THE HEADREST IS THE RECALLED ITEM; NOT THE CHAIRS):The dental chairs are sold under the brand names Marus and DCI Equipment. The affected models are:Marus MaxStar-Agency Model # DC1490;Marus MaxStar-Agency Model # DC1690;Marus NuStar-Agency Model # DC1700;Marus NuStar-Agency Model # DC1702;Marus ProStar-Agency Model # DC1540;Marus ProStar-Agency Model # DC1535;DCI Equipment DC1235 Hydraulic-Agency Model # DC1235; and DCI Equipment DC1335 Electromechanical-Agency Model # DC1335. | Dental Equipment LLC | 2 | The magnetic headrest on certain dental chairs may affect the function/programming of some implantable pacemakers or defibrillators if the implanted device is programmed to respond to a magnet. | DCI Equipment sent an "Urgent Field Safety Alert" letter dated March 15, 2010 to consignees advising them of the potential problem. The letter instructed that patients with implantable pacemakers or defibrillators should avoid dental chairs with the magnetic headrest. A warning label was sent with the letter for insertion into the product Use and Care manual for operator warning. Consignees were asked to complete a Field Safety Alert Acknowledgement/Return Form and fax to the firm. Consignees were requested to notify customers.Consignees may contact the firm at 503 537-3602.
Z-1173-2010 | Mar-19-2010 | 2010 | Xario - Xario; SSA-660A; software version 1.0 and later | Toshiba American Medical Systems Inc | 2 | Toshiba America Medical System Inc initiated a field corrective action on Aplio 50; SSA-700A; software version 5.5r002 and later; Aplio 80; SSA-700A; software version 5.5r002 and later; and Xario; SSA-660A; software version 1.0 and later; because the product does not meet the Safety Standard requirements. | Toshiba America Medical Systems (TAMS) has issued a Field Modification Instruction notification to all of its customers about the affected devices. The correct software will be administered through TAMS' representatives. Reimbursement for the modification will be given upon receipt of the reimbursement claim invoice.
Z-1405-2010 | Apr-16-2010 | 2010 | Elekta Synergy XVI Medical Linear Accelerator - Elekta Synergy XVI Medical Linear Accelerator. Model Numbers R3.5; R4.0 and R4.2. Elekta; Inc. Norcross; GA 30092.Intended for use as a radiation therapy treatment of malignant neoplastic diseases; as determined by a licensed medical practitioner. | Elekta; Inc. | 2 | It is possible that two (or more) CT reference data sets are necessary for one patient; for example; if you have to scan a patient during breath hold and then again during free breathing. | Elekta, Inc. issued an "Important Notice" dated December 12, 2008. Customers were informed of the affected device and instructed on proper steps until an update is provided. For further information, contact Elekta, Inc. at 1-770-300-9725.
Z-0365-2011 | Nov-16-2010 | 2010 | SoftLab GUI Version 4.0.1.0 - 4.0.1.16; 4.0.2.0 - 4.0.2.10; 4.0.3.0 - 4.0.3.13; and 4.0.4.0 - 4.0.4. - SoftLab GUI Version 4.0.1.0 - 4.0.1.16; 4.0.2.0 - 4.0.2.10; 4.0.3.0 - 4.0.3.13; and 4.0.4.0 - 4.0.4.5.SCC Soft Computer 5400 Tech Data Drive Clearwater; FL 33760. Manufacture/Distribution Dates:4.0.1.0 - 08/06/04; 4.0.1.4 - 10/18/04; 4.0.1.5 - 01/18/04; 4.0.1.6 - 12/15/04; 4.0.1.7 - 01/10/05; 4.0.1.8 - 03/29/05; 4.0.1.9 - 05/06/05; 4.0.1.13 - 01/25/06; 4.0.1.14 - 03/26/07 4.0.1.15 - 05/02/07; 4.0.1.16 - 08/22/08; 4.0.2.0 - 04/08/05; 4.0.2.1 - 08/13/08; 4.0.2.2 - 06/06/05; 4.0.2.3 - 09/19/05; 4.0.2.4 - 09/19/05; 4.0.2.5 - 09/19/05; 4.0.2.7 - 02/13/06; 4.0.2.8 - 12/12/06; 4.0.2.9 - 02/20/08; 04.2.10 - 08/13/08; 4.0.3.2 - 05/17/06; 4.0.3.3 - 09/18/06; 4.0.3.4 - 09/27/06; 4.0.3.5 - 11/29/06; 4.0.3.6 - 06/22/07; 4.0.3.7 - 03/28/07; 4.0.3.8 - 06/22/07; 4.0.3.9 - 08/03/07;4.0.3.10 - 08/24/07; 4.0.3.11 - 09/24/07; 4.0.3.12 - 03/07/08; 4.0.3.13 - 06/13/08; 4.0.4.2 - 02/22/08; 4.0.4.3 - 05/22/08; 4.0.4.4 - 07/28/08; 4.0.4.5 - 05/04/09.Intends to provide clients with an interface that links all clinical laboratory departments throughout the enterprise. | SCC Soft Computer | 2 | Clients using Patient Maintenance to perform Moves of Stays may send the wrong billing number during an ADT update to SoftWeb; SoftPath or SoftMedia. The wrong billing number can come from a different patient in Order Entry. Only Moves of Stays are affected. | Clients were notified via the proprietary communication software on 06/01/2009. This is a MANDATORY CORRECTION. The Risk-To-Health Notification Task gives the background on the issue, the potential hazard, root cause, corrective action, alternative working solution, and requested action. Clients were instructed to acknowledge receipt of the notification. If the client wishes to receive the correction, they should document a request to load the software.
Z-0077-2011 | Oct-18-2010 | 2010 | Sentry - semi automatic external defibrillator (AED) - DDC-6-AA Data Card; an accessory to the Sentry DDU-1 00 semiautomatic external defibrillator (AED).The DDC-6-AA data card is an accessory to the AED and does not have its own intended use. | Defibtech; LLC | 2 | Data cards may include an incorrect format which may cause the AED to not function correctly. | Defibtech Initially notified customers by telephone on 8/27/10. Customers were notified of the product, the problem and the action to be taken. Customers were instructed to locate affected data cards and send them back to Defibtech for evaluation. If the data card was inserted into the AED, then the customer was instructed to check the AED software version. If the software version was at 2.005, the customer was instructed to keep the AED out of service until Defibtech could evaluate the returned data card. (Any other software version indicated that the device was not affected by a faulty data card). If the software version was any other version other than 2.005, then the customer was instructed to keep the AED in service unless the original service inquiry required it to remain out of service. If the data card was not inserted into the AED, then the customer was instructed to return the data card for evaluation. For any questions customers were to contact Customer Service at (866) 333-4241 from 8:30-5pm EST, M-F.
Z-0865-2010 | Feb-26-2010 | 2010 | Electronic Data Interchange - i-STAT HL7 InterfaceThe HL7 Interface product is an accessory to the analyzer | Abbott Point of Care Inc. | 2 | Abbott Point of Care has determined that it is possible under certain circumstances that some test records could have their status update messages mis-matched with another record; resulting in the incorrect patient name; order number of comment being included in the CDS test record only. | Urgent Recall Notification letters were sent by Fed Ex to all US customers on September 29, 2009 and worldwide on October 1, 2009. Questions should be directed to Abbott Point of Care Technical Support at 1-800-366-8010, Option 1.
Z-2381-2010 | Sep-08-2010 | 2010 | Symbiq Infuser - Hospira Inc. Symbiq One- and Two- Channel Infuser; List Numbers 16026 and 16027 respectively. | Hospira Inc | 1 | If the administrator set is removed prior to the cassette carriage fully opening without closing the slide/roller clamp; unrestricted flow and delivery may result. | Hospitra Urgent Device Correction Notification mailed on 03/26/2010 via federal express. Tracking done by delivery tracking. Distribution of product continues and has not stopped. Additional notification was issued in 09/2010. A software correction is planned.
Z-1410-2010 | Apr-21-2010 | 2010 | LIFEPAK 15 Monitor/Defibrillator - Biphasic LIFEPAK 15 Monitor/Defibrillator. Manufactured by Physio-Control Inc.; a division of Medtronic; Inc. Redmond; Washington USA.The LIFEPAK 15 monitor/defibrillator is a complete acute cardiac care response system designed for basic life support (BLS) and advanced life support (ALS) patient protocols. | Physio Control; Inc. | 1 | Potential for the device to power off then on by itself; or to power off by itself and requiring the operator to turn it back on; or the device doesn't turn off. | Physio-Control Inc. issued an "Urgent Medical Device Correction" notification dated March 2010 to domestic consignees. Foreign consignees were notified by overseas representatives. Consignees were advised to keep the affected device in service and to test the units in accordance with operating instructions. They were further advised that a local service representative will call and schedule a service visit within 60 days. Consignees were asked to forward the notification to all of their sites and to notify the recalling firm if they no longer own the device. For further information, contact Physio-Control Inc. Technical Support at 1-800-442-1142, Option 5, 6:00 am to 4:00 pm (Pacific), Monday-Friday.
Z-1930-2010 | Jul-13-2010 | 2010 | Constellation Vision System - Constellation Vision System; Model: Constellation LX; Catalogue Number: 8065751551The Constellation Vision System is an ophthalmic microsurgical system that is indicated for both anterior segment (i.e. phacoemulsification and removal of cataracts) and posterior segment (i.e.; vitreoretinal) ophthalmic surgery in addition to the indications included with the optional Next Generation laser. | Alcon Research LTD dba Alcon Laboratories; Inc. | 1 | Alcon identified system performance and machine settings that may impact the infusion performance of the CONSTELLATION Vision System; causing unexpected shut down; unresponsive touch screens; and fluidics issues. | The firm, Alcon, sent a "URGENT: MEDICAL DEVICE RECALL" letter dated July 2, 2010, along with training material to all customers. The letter informed the customers of the problem and indicated that Alcon Surgical Sales representative will be contacting them in the near future to follow-up on this training material and to answer any questions you may have. Alcon informed the customers that this action will not require removal of the CONSTELLATION Vision System from their facility. The letter also informed the customers that their Alcon Technical Customer Service Representative will be contacting them when we are authorized to provide you with the software and hardware updates for your system. For assistance and questions regarding the CONSTELLATION Vision System, customers were told they may contact their Alcon Surgical Sales Representative or Alcon Technical Consumer Affairs at 1-800-561-6466, Option 8.
Z-0824-2010 | Mar-04-2010 | 2010 | Circulatory Support System - Abiomed AB5000 Circulatory Support System; Catalog Number: 0015-0000. Manufactured by Abiomed; Inc. Danvers; MA.Intended to provide complete short-term support of the left and/or right sides of the heart to patients suffering from potentially reversible ventricular dysfunction. | Abiomed; Inc. | 1 | Console shut down with no audible alarm | Abiomed initiated the recall by telephone contact and followed with a Recall Letter dated January 27, 2010. Consignees were informed that a Field Service Engineer will be in contact and schedule a time to visit and correct the affected device. For further information, contact Abiomed Quality Assurance at 1-978-656-1543 or Abiomed Field Service at 1-800-554-8666.
Z-1657-2010 | May-26-2010 | 2010 | Coulter¿ LH 500 Series System - Coulter¿ LH 500 Series System. Beckman Coulter; Inc. Brea; CA.Quantitative; automated hematology analyzer and leukocyte differential cell counter for in vitro diagnostic use in clinical laboratories. | Beckman Coulter Inc. | 3 | Issue 1: The system will substitute or omit the characters # @ [\] ' {I) when the system is configured in a language other than English or Chinese. Example: Using the character @; a substitution in the string '123@456' would result in '123#456' and an omission would result in '123456'Issue 2: The system will omit the character ~ regardless of language.Impact: There is a potential for Specimen | Beckman Coulter, Inc. issued "Urgent: Product Corrective Action" notifications dated February 12, 2010. Consignees were informed of the affected product and instructed on proper actions. The firm also asked that the enclosed response forms be returned. For further information, contact your local Beckman Coulter Representative or call 1-800-526-7694.
Z-0719-2011 | Dec-21-2010 | 2010 | Bard¿ CritiCore¿ - Bard¿ CritiCore¿ Monitor; Software Version 2.1; Catalog number 000002N; C.R. Bard; Inc.; Covington; GA 30014.Intended for multiple uses as it is an electronic battery-operated monitor used in conjunction with a single-use disposable urine collection container and/or a temperature sensing Foley catheter | C.R. Bard; Inc.; Urological Division | 3 | Error can occur in the reading of the Present Interval Output (PIO) reading when the device is subjected to bumping or forceful replacement of the fluid container. | Consignees were notified by an Urgent: Product Advisory Letter and Medica Device Recall letter on/about November 18, 2010. The letter identified the affected product, stated the reason for recall, the interim precautions, and actions to be taken. The letter requested acknowledgement of the recall and to return the entire affected product. Each consignee that is a distributor was instructed to forward the letter to their consignees to retrieve relevant product from those locations. Affected product received by Bard Medical will be reworked with the software upgrade and place back into distribution. Consignees will either receive a software upgrade on site by a third party or replacement product with the software upgrade. For international distribution, the hospital will be informed of this issue through Bard affiliate and Notified Body Bard Europe and managed according to the requirement. If there are questions or assistance is needed, customers can contact BD at 1-770-784-6969.
Z-1437-2010 | Apr-26-2010 | 2010 | Enfant Pediatric VEP Vision Testing System - Diopsys Inc.; Enfant Pediatric VEP Vision Testing System; Input: 120 Vac; 60Hz/2.5A IEC 60601-1 Class I; Type BF. | Diopsys; Inc. | 3 | The computer based system may lose test counts when the time in the device was changed during a changeover from standard time to daylight savings time in March 2007. | Field Service personnel were sent an e-mail notification on 4/2/2007 which included a TZedit file to be downloaded. E-mails instructed reps to install the patch on their customers devices immediately. Field correction was completed 7/31/2007.
Z-0870-2010 | Feb-26-2010 | 2010 | ERIC (Electronic RapID Compendium) Version 1.0.77 - ERIC (Electronic RapID Compendium) Version 1.0.77; Catalog Number R8323600; used in conjunction with: remel RapID Yeast Plus System is a qualitative micromethod employing conventional and chromogenic substrates for the identification of medically important yeast; yeast-like; and related organisms isolated from human clinical specimens. A complete listing of the organisms addressed by the RapID Yeast Plus System is given in the RapID Yeast Plus Differential Chart.RapID" YEAST PLUS System; Catalog No. R8311007. 4 hour identification of yeast based on enzyme technology. Database includes over 40 taxa. Requires Inoculation Fluid; REF R8325106. CE; 20 Panels/Kit The responsible firm name on the product labeling is Remel; 12076 Santa Fe Dr.; Lenexa; KS 66215-3594 | Remel; Inc | 3 | There is an error in Version 1.0.77 of the electronic code compendium (ERIC) software. The error is limited to the database corresponding to the R8311007 RapID Yeast Plus System. The database erroneously includes Candida dubliniensis as well as displays an incorrect temperature range as an accessory test for codes with an overlap between Candida albicans and Candida dubliniensis. When using the | The firm issued an "IMPORTANT MEDICAL DEVICE PRODUCT RECALL NOTICE" dated October 29, 2009. After completion of the firm's investigation and risk assessment, Remel decided to notify customers of this situation and quarantine any remaining units on hand. The notification provided the consignees with instructions for obtaining and downloading the software patch. A response form was included with the recall letter and the consignee was asked to complete the form and return it to Remel. The consignees should contact the Technical Services Department at 800-255-6730 (USA) or 913-888-0939 (International), if they have inquiries concerning this advisory.
Z-2438-2010 | Sep-20-2010 | 2010 | Integrated Data Management System - Version 7.1; 7.2 and 7.3; - Integrated Data Management System - Version 7.1; 7.2 and 7.3; ITC Corporation; 8 Olsen Place South; Piscataway; NJ 08854 | International Technidyne Corporation | 3 | A tabulation error in the Integrated Data Management Systems (IDMS) data management software that results in incorrect Electronic Quality Control (EQC) Summary Reports: The Operator and the Device Performance Summary Reports. The IDMS software is not used for the clinical management of individual patients. | ITC sent an URGENT MEDICAL DEVICE CORRECTION letter dated March 10, 2010 to Distributors and First Consignee End Users identifying the affected product and affected summary reports, the reason for recall, and the actions to be taken by the distributors and customers. New Errata Instructions were enclosed with each letter and are to be put in the idms user guide for reference. The new instructions are to be followed when running reports for the ESV or TVT tests run by the Hemochron Response. The enclosed Customer Account Tracking Form is to be completed and returned to ITC indicating that the letter has been received and appropriate questions answered as to the use of the software system. Distributors are asked to check inventory for recalled product. If affected product is found, shipment should be stopped and placed on hold. Distributors should notify customers of the correction and provide them with a copy of the correction letter and updated instructions. This should be done by traceable mail. Distributors should return the account tracking form and return affected software disks for replacement. Questions should be directed to ITC Technical Support at 800-631-5945.
Z-0478-2011 | Nov-26-2010 | 2010 | NanoMaxx ultrasound system - The software is ARM version 60.80.101.025 (referred to as Nano 1.0.5).The software is used with the SonoSite NanoMaxx ultrasound system. | Sonosite; Inc. | 3 | Software ARM version 60.80.101.025 (referred to as Nano1.0.5) on the NanoMaxx ultrasound system results in errors when the "print" command is selected. As a result; the displayed MI or TI values may be incorrect. | Sonosite issued Urgent Medical Device Correction letters dated August 9, 2010 to consignees. Consignees were advised of the software problem and given directions to verify the software version they are using. They are advised to contact SonoSite at 877-657-8118 to obtain a free software upgrade.
Z-1658-2010 | May-26-2010 | 2010 | Unicel¿ DxH 800 Coulter¿ Cellular Analysis System - Unicel¿ DxH 800 Coulter¿ Cellular Analysis System. Beckman Coulter; Inc. Brea; CA.Quantitative; automated hematology analyzer and for in vitro diagnostic use in screening patient populations found in clinical laboratories. | Beckman Coulter Inc. | 3 | Issue 1: The system will substitute or omit the characters # @ 1 \ ] ' ( I ) - when the instrument is configured in a language other than English; Chinese or Japanese. Example: Using the character @; a substitution in the string '123@456' would result in '123#456' and an omission would result in '123456'.Issue2: The characters 7 " and * (question mark; quotation mark; and asterisk) used as part | Beckman Coulter, Inc. issued "Urgent: Product Corrective Action" notifications dated February 12, 2010. Consignees were informed of the affected product and instructed on proper actions. The firm also asked that the enclosed response forms be returned. For further information, contact your local Beckman Coulter Representative or call 1-800-526-7694.
Z-1560-2010 | May-08-2010 | 2010 | One Touch Ping Insulin Pump - One Touch Ping Insulin Pump using ezManager Max diabetes management software | Animas Corporation | 3 | Software malfunction which disallows users to download; view and print information from their pump (including blood glucose; insulin delivery; carbohydrate and pump events). Additionally; users are unable to customize the built-in food database and pump settings using the software. | The recalling firm telephoned the patients/users to inform them of the problem. A letter dated 12/22/09 was also sent to users. Users will have a replacement unit sent to them so that they will be able to fully utilize the pump and the software. As a courtesy a letter dated 12/22/09 was issued to the health care provider/prescribing physician for the affected users.
Z-0106-2011 | Oct-20-2010 | 2010 | TrueBeam and TrueBeam STx - Varian brand TrueBeam and TrueBeam STx; Model Number:TMX- H19; Distributed by and Manufactured by: Varian Medical Systems Inc.; Palo Alto; CA | Varian Medical Systems Oncology Systems | 2 | Due to a software anomaly; guidance-based couch shift values may not be applied as expected when the operator selects "Apply Shift" and presses Motion Enable buttons. This can result in the actual patient position differing from that position indicated by the imaging application affecting poistion accuracy and precision. | Varian Medical Systems issued an Urgent Medical Device Correction letter date dated May 24, 2010 to users identifying the affected device, the issue prompting the correction, and the actions to be taken by the user. Customers were asked to set couch tolerance values to small numbers, and for radiosurgical treatment tolerance values to be set to zero. A Varian representative will contact customers to schedule installation of a software version correcting the issue. US Customers can contact Varian at 1-888-827-4265,. European customers can contact Varian at +41 41 749 8844.
Z-1901-2010 | Jun-25-2010 | 2010 | VARIANT II Hemoglobin Testing System with CDM 3.5 Software Version 3.5 - VARIANT II Hemoglobin Testing System with CDM 3.5 Software Version 3.5; Model Number: 270-2000; Product is manufactured and distributed by Bio-Rad Laboratories; Diagnostics Group; Hercules; CA.Provides an integrated method for sample preparation; separation and determination of the relative percent of specific hemoglobins in whole blood. | Bio-Rad Laboratories Inc | 2 | Sample test results for the first run appears to be correct but all subsequent results during that run were assigned incorrect barcode and tube position. | Bio-Rad Laboratories, Inc. issued and "Urgent: Medical Device Correction" dated January 22, 2009 to all affected consignees by FED EX, followed up by telephone calls for those not responding. Consignees were instructed on proper steps to take in dealing with the affected product. Further communication will follow to provide a solution. For further information, contact Bio-Rad Laboratories at 1-510-724-7000.
Z-0111-2011 | Oct-21-2010 | 2010 | Consulta CRT-D - Medtronic Consulta CRT-D; model D224TRK and Model D234TRK not available in the US. Digital implantable cardioverter defibrillator with cardiac resynchronization therapy (DDE-DDDR). Sterilized using ethylene oxide. Medtronic; Inc.; Minneapolis; MN 55432 USA. Manufactured in: Juncos; Puerto Rico; USA.The Dual chamber implantable cardioverter defibrillator with cardiac resynchronization therapy (CRT-D) is a multiprogrammable cardiac device that monitors and regulates the patient's heart rate by providing single or dual chamber rate-responsive bradycardia pacing; sequential biventricular pacing; ventricular tachyarrhythmia therapies; and atrial tachyarrhythmia therapies. The device automatically detects ventricular tachyarrhythmias (VT/VF) and provides treatment with defibrillation; cardioversion; and antitachycardia pacing therapies. The device also automatically detects atrial tachyarrhythmias (AT/AF) and provides treatment with cardioversion and antitachycardia pacing therapies. Simultaneous or sequential biventricular pacing is used to provide patients with cardiac resynchronization therapy. The device responds to bradyarrhythmias by providing bradycardia pacing therapies. | Medtronic Inc. Cardiac Rhythm Disease Managment | 2 | Medtronic is informing you of a rare device software issue in certain ICD and CRT-D devices. A software update will be available to correct his issue. There have been 5 reports of this issue occurring with no reported patient injuries or deaths. Medtronic has identified the root cause to be a rare and specific sequence of events that must occur within a few milliseconds of each other. In the u | Medtronic, Inc. sent an IMPORTANT MEDICAL DEVICE CORRECTION LETTER, dated May 2010 to physicians beginning May 05, 2010. The letter informed doctors of the issue and informed them that a software update would be available to correct this issue. The software was to be installed on the programmer, with a Medtronic representatives assistance, so that on a patient's next in-clinic interrogation the device would receive the upgrade automatically and their device would be corrected. For any questions regarding this recall physicians were to contact their local Medtronic Representative or Medtronic Technical Services at 800-723-4636.
Z-1040-2010 | Mar-09-2010 | 2010 | Concha Therm Neptune Heated Humidifier - Teleflex Medical; Hudson RCI; Concha Therm Neptune; Heated Humidifier. Product numbers: 425-00; 425-10 and 525-30. | Telefelx Medical | 2 | When turned on; the Neptune heater may cause an artifact on the patient's heart monitor. | Teleflex consignees were notified via Urgent Medical Device Recall letter via next day delivery on/about 01/12/2010. They were asked to return all of the affected product. Each consignee that is a distributor was instructed to forward the letter to their consignees to retrieve relevant product from those locations. A second mailing to non-responders on the consignee list will be made approximately four to six weeks after the first mailing. A third mailing to non-responding customers on the consignee list will be made approximately four to six weeks after the second mailing or as an alternative, they may be contacted by phone, e-mail or fax. A record of these contacts will be maintained. Affected products received by Teleflex Medical will be reworked with software upgrade and placed back into distribution. Consignees will either receive a software upgrade on site by a third party or replacement product with the software upgrade. Records will be maintained accordingly.
Z-1590-2010 | May-10-2010 | 2010 | Stryker Endoscopy SDC HD and SCD HDi - An Imaging device for medical procedures; to provide surgeons or operating room personnel the ability to capture still images and to record video during the surgical procedures in various file formats. Images can be transferred to external media. | Stryker Endoscopy | 2 | Video device is not able to assign the correct time/date stamp. | Stryker Endoscopy has notified the 2865 domestic consignees affected by the device correction. The notification was mailed to affected consignees and their corresponding sales representatives. It included two mailings of the Urgent: Device Correction notice, SDC HD 3.0z correction software, and a SDC HD upgrade dongle. 100% of consignees have been contacted and notified of the device correction needed. The effectiveness of the correction has been tracked via acknowledgement cards mailed, faxed, or emailed. The local Stryker Sales Representative will install a software upgrade to fix this issue. Additional questions should be directed to the firm's Customer Service at (800)624-4422 or (408)754-2039.
Z-1214-2010 | Mar-29-2010 | 2010 | N'Vision software Application Card - Medtronic N'Vision 8870 Application Card for use in the Medtronic model 8840 Physician Programmer. Manufactured for: Medtronic; Inc.; Minneapolis; MN 55432-5604 USA. | Medtronic Neuromodulation | 2 | Routine monitoring of Medtronic's vigilance program identified user related programming errors. Medtronic has received reports of both overdose and underdose that were the result of programming use of SynchroMed implantable infusion pumps. These errors were most commonly associated with: --Priming Bolus Programming; and --Bridge Bolus Programming. Pump programming errors that result in overdos | Medtronic is exchanging 8870 versions AAG, AAJ and BBG and replacing it with the updated versions AAH01 and BBH01. The updated versions contain user interface improvements to reduce the chance of user related programming errors when programming SynchroMed II and SynchroMed EL Infusion Pump system. A Medtronic "Medical Device Correction" letter was left behind for consignees. The letter was addressed to "Healthcare Provider" and was dated February 2009. The letter provided information regarding changes to the programming software used in SynchorMed Infusion System. For additional information, please contact your Medtronic Representative or Medtronic Neuromodulation Technical Services at 1-800-707-0933.
Z-1256-2010 | Apr-07-2010 | 2010 | NicoletOne vEEG System - NicoletOne vEEG System; VIASYS Healthcare; Inc.; Madison; Wisconsin 53711. | CareFusion | 2 | CareFusion NeuroCare is voluntarily implementing a field correction to delete certain NicoletOne software protocols; which users may interpret incorrectly when conducting patient evaluations. This may result in the misinterpretation of the left versus right side of the brain.Users are being provided with instructions for removingthese protocols and verifying that the deletion has beenprop | Consignees were sent a CareFusion "Urgent Medical Device Correction" letter dated March 3, 2010. The letter described the problem and product involved. The letter included "Action to be taken by Users," "Instructions for deleting protocols," and "Photos-Amplifier label and Headbox insert to be discarded". It also included a "Verification Form" to be completed by recipient. For questions, notify CareFusion-NeuroCare, Attention: Sue Niesen, at (800) 356-0007 ext. 5129, or via email at susan.niesen@carefusion.com.
Z-2121-2010 | Aug-02-2010 | 2010 | MEDFUSION¿ SYRINGE INFUSION PUMP; - MEDFUSION¿ SYRINGE INFUSION PUMP; MODELS 3010 AND 3010a WITH SOFTWARE VERSIONS 2.0.2; 2.0.3; AND 2.0.4 Medfusion¿ Syringe Infusion Pumps are designed for use in critical care areas for the administration of fluids requiring precisely controlled infusion rates. | Smiths Medical ASD; Inc. | 2 | Smiths Medical is conducting a correction of Medfusion? Syringe Infusion Pumps; Models 3010 and 3010a; with obsolete software versions 2.0.2; 2.0.3; and 2.0.4. Smiths Medical has found that pumps with obsolete software may continue to run past the set volume limit. This can result in an over-delivery  if the syringe is over-filled. Smiths Medical has received no reports of serious patient in | Consignees were sent on 6/22/10 a Smith Medical "Urgent Medical Device Correction Notice', dated June 11, 2010. The letter was addressed to Dear Risk/ Safety Managers, Distributors, Clinical/ Biomedical Engineering, Clinicians, Critical Care Physicians, Pediatricians, Pediatric Intensivists, Neonatologists; and Nursing, Pharmacy,and Anesthesia Professionals. The letter described the problem and product involved. It provided Advise on Action to be Taken by the User which included to 1) Determine the software version of their Medfusion pump 2. Complete and return the Confirmation Form 3. Return the completed Software Update Test Form 4. If unable or unwilling to perform the software upgrade they provided information for returning the pumps to Smiths Medical for the upgrade. For further information please contact Smiths Medical at 651-628-7280.
Z-1020-2010 | Mar-16-2010 | 2010 | Siemens Rapidpoint 405 Test System - RapidPoint¿ 405 System Analyzers The RP405 is capable of testing all of the analytes listed below including Co-Oximetry.The Siemens 400 series system is intended for point of care and laboratory testing of blood gases; electrolytes and metabolites in arterial; venous and capillary whole blood samples. | Siemens Healthcare Diagnostics Inc | 2 | Software: Reference sensor failure may result in inaccurate pH reuslts | An "Urgent Field Safety Notice" dated September 2009 was sent via overnight mail to customers and included software Version 3.5.2 with instructions for install that will detect and prevent the reporting of potentially errant test results. The customers are requested to please install the included software update immediately and return the attached Fax back form. Siemens will provide the mandatory update to all customers with the Rapidpoint 400 /405 test system. If you have any questions, please contact your local Technical Support at 508-359-3877.
Z-0363-2011 | Nov-15-2010 | 2010 | Solar 8000i with Patient Data Module / Transport Pro Monitor - GE Healthcare Transport Pro¿ Monitor with the CARESCAPE" Patient Data Module. | GE Healthcare; LLC | 2 | Transport Pro Monitor stops communication with the CARESCAPE Patient Data Module (PDM) after 414 days of continuous run time. This time will equate to different amounts of real time depending on how much the units is actually in service per day. Transport Pro contains an internal timer that is used to control the software and remind users to perform preventive maintenance. When this internal timer | GE Healthcare issued an "Urgent Medical Device Correction" letter dated September 17, 2010 to consignees. The letter was addressed to Hospital Administrator, Head of Biomedical Engineering and Nursing Manager. The letter described the product, Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information. Service representatives will updates all of the affected Transport Pros with PDM that were distributed. Customers may contact GE at 262-548-2731 about this correction.
Z-1412-2010 | Apr-20-2010 | 2010 | Eclipse Treatment Planning System - Varian Medical Systems; Eclipse Treatment Planning System; with v client versions 7.3 (build 7.3.10); 8.0 (builds 7.5.x); 8.1 (builds 8.1.x) and 8.5( builds 8.2.x). For radiation therapy. | Varian Medical Systems Oncology Systems | 2 | Potential for misdelivery of treatment-- When preparing a treatment plan involving a split; the device may calculate excessive or insufficient monitor units for the treatment plan. | Recall initiated on 12/09/2009. Varian issued a notification letter to its consignees, including a description and corrective action step. The firm is in the process of preparing a software fix for the issue.
Z-2211-2010 | Aug-12-2010 | 2010 | DATEX-OHMEDA NETWORK AND ICENTRAL - DATEX-OHMEDA NETWORK AND ICENTRAL '05; SALES REVISION 4.3 ( Datek-Ohmeda Network and iCentral (including iCentral Client)) The Datex-Ohmeda Network and iCentral (including iCentral Client) transfers information between networked Datex-Ohmeda devices in the Monitor Network as well as between Monitor Network and networked devices in GE Unity NetworkTM. It also allows information transfer between several iCentrals. Within one Monitor Network it allows a networked device to display; store; print and otherwise process information received from other networked devices. The iCentral maintains the network connections between the Datex-Ohmeda bedside monitors and other networked devices in Monitor Network. Network connections consist of hardwired network cables and/or Wireless LAN (WLAN) connections. Furthermore; it coordinates the transfer of information between devices in the Datex-Ohmeda Network; between the Datex-Ohmeda Network and devices in GE Unity Network as well as between the Datex-Ohmeda Network and Hospital Information Systems (HIS). The iCentral can be used for remote monitor management; storing; printing; viewing; reviewing or otherwise processing of information from several bedside monitors or other networked devices. The iCentral Client can be used for remote monitor management; printing; viewing; reviewing or otherwise processing of information from several bedside monitors or other networked devices. The Datex-Ohmeda Network will be used for patients in the hospital and it is meant for continuous use. The device is for use by qualified personnel only. | Ge Healthcare; Llc | 2 | The iCentral or iCentral client v5.0.2 with HP xw 4400 and 4600 may spontaneously shutdown; causing monitoring loss at the central monitoring station. Monitoring loss may result in prolonged delay of treatment of potentially lethal arrhythmia. | GE Healthcare sent a "Urgent Medical Device Correction letter" dated 6/18/10 on 6/25/10 to consignees. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information. The letter instructs consignees to follow specific detailed instructions for safe use of the device pending provision of a software upgrade by the firm. Consignees will be contacted by a GE Healthcare representative to schedule the software upgrade.
Z-0476-2011 | Nov-26-2010 | 2010 | Seno Advantage 1.x; Seno Advantage 2.0; Seno Advantage 2.1 - GE Seno Advantage 1.x; Seno Advantage 2.0; Seno Advantage 2.1 | GE Healthcare; LLC | 2 | The measurement values provided by the GE Workstation may be incorrect when applied to magnification images acquired on non-GE digital mammography systems.If not noticed by the caregiver; this could lead to an overestimate of the size of the breast lesion that may impact patient treatment and safety. The measurement values for GE images are not affected. | GE Healthcare issued "Urgent Medical Device Correction" letters dated August 28, 2009 or August 31, 2009 to either Seno Advantage Customers (Hospital Administrators/Risk Mangers, Radiology Department Managers), and Image Diagnostic Customers (Hospital Administrators/Risk Managers, Radiology Department Managers). The letters described the Safety Issue, Affected Product Details, safety Instructions, Product correction and Contact Information. GE will linstall new software with the release of Field Modification Instruction 12129. Customers can contact GE Healthcare about this issue at 800-437-1171.
Z-0455-2011 | Nov-24-2010 | 2010 | Varian brand Eclipse/VARiS Vision - Varian Medical Systems; Inc. Eclipse/VARiS Vision; an Eclipse radiotherapy Treatment Planning System (Eclipse TPS); Eclipse Version 7.3.10 (multiple workstations) VARiS Vision System 7. x;Model Numbers: H48/H83. | Varian Medical Systems Oncology Systems | 2 | Incorrect treatment plan or record: A software design deficiency has the potential for corrupted treatment plan records or treatment records in cases where multiple users editing a record simultaneously can result in a corrupted (but accessible) treatment plan record. | On May 13, 2010, the firm issued letters to all its direct consignees, informing them of the affected products and providing instructions on the recall.
Z-1131-2010 | Mar-15-2010 | 2010 | Optovue RTVue Optical Coherence Tomography with Normative Database - Optovue RTVue Optical Coherence Tomography (OCT); Model-RT100 with software versions 3.0.x.x and higher; with Normative Database; manufactured by Optovue; Fremont; CAThe RTVue is an optical coherence tomography system indicated for the in vivo imaging and measurement of the retina; retinal nerve fiber layer; optic disk; and for cornea and anterior eye scans as an aid in the diagnosis and management of retinal disease. | Optovue Inc. | 2 | Marketed without proper 510(k)-- No 510(k) clearance for Optovue RTVue Model-RT100 with Software Versions 3.0.x.x and higher which includes the Normative Database. | The firm is preparing a software "Roll Back" version of software to disable the NDB and other features, ensuring that this version will work properly. A customer notification letter is provided as well as a correction verification form, installation instructions and user manual. If there are any questions, please call 1-866-344-8948.
Z-0624-2011 | Dec-16-2010 | 2010 | Cmax Surgical Table - Cmax Surgical Tables; Model #'s: Cmax110; Cmax 220 manufactured by Steris Corporation in Montgomery; AL | Steris Corporation | 2 | Cmax Surgical Tables may lose calibration if the table is operated only on battery power and if the table is in a low battery charge condition. This could result in the table moving into an unexpected tilt position when the Level button is depressed on the hand control keypad. | Steris issued Urgent Field Correction Notice letters dated August 26, 2010 to each affected consignee notifying them of the voluntary field correction. A Steris Representative will contact customers to arrange for STERIS to upgrade the software in the table. For further information about this corrective action, contact Steris at 1 800 333-8828.
Z-0462-2011 | Nov-24-2010 | 2010 | Signa¿ Excite" 1.5 T Magnetic Resonance Systems; GE Healthcare; P.O. Box 44 Milwaukee; WI 53201. - Signa¿ Excite" 1.5 T Magnetic Resonance Systems; manufactured by GE Healthcare; P.O. Box 44 Milwaukee; WI 53201. | GE Healthcare; LLC | 2 | Signa? MR Excite" 1.5T system containing 11.1_M4_0818.a software; associated with 2D Fast Spin Echo based pulse sequences (2D FSE-XL; 2D FRFSE-XL; 2D FSE-IR; 2D T2FLAIR and 2D T1FLAIR) may experience an image flip along the phase encode direction and incorrect annotation; which may result in incorrect localization of anatomy in oblique axial images with the pulse sequences. | GE Healthcare issued an "Urgent Medical Device Correction" letter dated April 21, 2009. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information. GE Field Engineers will contact customers to schedule a field modification of the affected products to address this issue. Customers can contact GE Healthcare at 262 521-6100 about this issue.
Z-0004-2010 | Jan-20-2010 | 2010 | GE OEC 9900ESP Elite - GE OEC 9900 Elite ESP To produce mobile fluoroscopic images of human anatomy. | Ge Healthcare | 2 | A software defect - may result in a false indication and warning on the OEC 9900 Workstation and C-arm display that the X-Ray tube Anode is over-heated. This will also cause the displayed air kerma rate and cumulative air kerma to be inaccurate. | An Important Electronic Product Radiation Warning letter will be sent to customers identifying the issues, interim solution, and permanent solution. Questions or concerns are to be directed to the GEHC OEC service team at 1-800-874-7378.
Z-1547-2010 | May-07-2010 | 2010 | ZOLL E Series Defibrillator/Pacemaker/Monitors - ZOLL E Series Defibrillator/Pacemaker/Monitors; BLS Model; with software versions 4.XX; 6.XX and 7.0X. Zoll Medical Corporation; Chelmsford; MA.Intended for the purpose of converting ventricular fibrillation to sinus rhythm or other cardiac rhythms capable of producing hemodynamically significant heart beats. | ZOLL Medical Corporation; World Wide Headquarters | 2 | Device issued Shock Advised message but failed to auto-charge the defibrillator. | Zoll Medical Corporation issued an "Urgent Medical Device Correction" notification dated March 26, 2010 via certified mail to consignees who have received affected product. Customers will be contacted and given the option to have a software upgrade kit to update their devices or return devices to the firm to have the device upgraded. For further information, contact Zoll Medical Corporation 24/7 Technical Support at 1-800-348-9011 or 1-978-421-9460.
Z-1803-2010 | Jun-15-2010 | 2010 | SwitchPoint Infinity 2 (SPI2) and SwitchPoint Element (SPE) - The SwitchPoint Infinity 2 (SPI2) and SwitchPoint Element (SPE); an integrated voice; video; data router and teleconferencing interface for the operating room.The SwitchPoint Infinity 2 (SPI2) and SwitchPoint Element (SPE) are designed as an integrated voice; video; data router and teleconferencing interface for the operating room. The primary intent of the SPI2 and SPE product lines is to allow operating room personnel a center point for controlling all equipment and communication during surgery. The SPI2 and SPE route surgical and non-surgical images throughout the operating room to facilitate operating procedures. All video inputs and video outputs of the router require the use of various signal and cables types. Digital and analog video signals transmit from the auxiliary devices in the room through the SPI2 and SPE router to designated monitors. The primary difference between the SPI2 and SPE systems is that the SPE is a more compact version of the SPI2 offering less video inputs and outputs; and is intended to be used in smaller clinics. | Stryker Communications Corp | 2 | Users may experience issues with the functionality of the SwitchPoint Infinity 2 and SwitchPoint Element that may be exhibited as a "Page Not Found" error message; missing drop down options; or a frozen touch panel. | The firm, Stryker, sent an "URGENT: Medical Device Correction" letter dated May 4, 2010, to all customers. The letter describes the product, problem and action to be taken by the customers. The firm recommended that the customer continue to use their SwitchPoint Infinity 2 or SwitchPoint Element. The customer is instructed to refer to their user manual, in the event they experience any failures; drive on module failure and/or product failure during a procedure. A Stryker representative will contact the customers facility to coordinate a service event to replace the existing component with a new component at no cost. Should the customer experience any of the aforementioned failures and require urgent technical support, please call Stryker Technical Support at 1-866-841-5663. If you have any questions regarding this letter or the implementation of this correction, please contact Shanta McKinzie at (972) 410-7319 or via email at StrykerCommunicationsRA@stryker.com.
Z-2193-2010 | Aug-09-2010 | 2010 | Crossfire Console - Crossfire Console; model number 475-000-000; Manufactured by Stryker Endoscopy San JoseMedical device indicated for use in orthopedic and arthroscopic procedures for the knee; shoulder; ankle; elbow; wrist and hip. The system provides abrasion; resection; debridement; and removal of bone and soft tissue through its shaver blade and ablation and coagulation of soft tissue; as well as hemostasis of blood vessels through its electrosurgical probe. Examples of use include resection; ablation and coagulation of torn knee cartilage; subacromial decompression and resection of synovial tissue in other joints. The probe is contraindicated for use in procedures where a nonconductive irrigant is used or with patients having cardiac pacemakers or other electronic implants. | Stryker Endoscopy | 2 | RF energy or powered shaver handpieces may continue to operate when not intended when used with the iSwitch Wireless Foot switch. | Stryker Endoscopy sent an "Urgent: Device Correction" letter dated June 28, 2010, via FedEX Mailer to US customers. A letter dated July 9, 2010, was sent to International sites.The letter describes the product, problem, and action to be taken by customers. The Customers were instructed to follow the included instructions and use the USB dongle to install the correction and to complete and sign the enclosed "acknowledgement of receipt" form,upon completion of the software correction, and fax to (480) 754-8378 or scan in and email to crossfireconsole@stryker.com. The firm noted that-It is important to send the self addressed confirmation forms back to Stryker RAQA. If you have any questions regarding this letter, please contact us by phone at 1-800-624-4422 or via email at crossfireconsole@stryker.com.
Z-1801-2010 | Jun-11-2010 | 2010 | The CARESCAPE" Monitor B850 version 1.0.1.13 and 1.0.2.1 - GE Healthcare; CARESCAPE" Monitor B850; Software version 1.0.1.13 and 1.0.2.1The CARESCAPE " Monitor B850 is a multi-parameter high acuity patient monitor intended for use in multiple areas within a professional healthcare facility. The CARESCAPE Monitor B850 is intended for use on adult; pediatric; and neonatal patients and on one patient at a time. The CARESCAPE Monitor B850 system is indicated for monitoring of Hemodynamic (including ECG; ST Segment; Arrhythmia Detection; ECG Diagnostic Analysis and Measurement; Invasive Pressure; Noninvasive Blood Pressure; Pulse Oximetry; Cardiac Output; Temperature; Impedance Respiration and SvO2 (Mixed Venous Oxygen Saturation)); Airway Gases (Fi/Et C02; 02; N20 and Anesthetic Agent); Spirometry; Gas Exchange (02 Consumption (VO2); C02 production (VCO2); energy expenditure (EE); and respiratory quotient (RQ)) and neurophysiological (including electroencephalography (EEG); Entropy; Bispectral Index (BIS) and Neuromuscular Transmission (NMT) Monitoring) status. The CARESCAPE Monitor B850 provides alarms; trends; snapshots and events; and calculations and can be connected to displays; printers and recording devices. The CARESCAPE Monitor B850 can be a stand-alone monitor or interfaced to other devices. It can also be connected to other monitors for bed to bed viewing and to data management software devices via a network. The CARESCAPE Monitor B850 is intended for use under the direct supervision of a licensed healthcare practitioner; or by personnel trained in proper use of the equipment in a professional healthcare facility. In addition to the healthcare practitioner; the CARESCAPE Monitor B850 is designed to provide configuration and troubleshooting capabilities to qualified service personnel. The CARESCAPE Monitor B850 is not intended for use during MRI. | Ge Healthcare; Llc | 2 | GE Healthcare has become aware of a safety issue associated with the CARESCAPE" MonitorB850 when monitoring ECG. The safety issue includes the potential for a delay in treatment. | The customers were sent on May 6, 2010, a GE Healthcare "Urgent Medical Device Correction" letter dated April 22, 2010. The letter was addressed to Healthcare Administrator / Risk Manager, Chief of Nursing, Director of Biomedical Engineering. The letter listed the Safety Issue, Affected Product Details, Product Correction, Safety Instructions and Contact Information. The customers were instructed to immediately discontinue use of the "Update Lead Set" function until the device software has been upgraded. If you have any questions regarding this medical device correction or identification of affected items, please contact Technical Support at 1-800-558-7044, or your local service representative.
Z-0076-2011 | Oct-18-2010 | 2010 | Sequencer - Sequencer Verify & Record System; for use with radiotherapy treatment machines; Class IIb; manufactured by Impac Medical Systems; Sunnyvale; CA. | Impac Medical Systems Inc | 2 | Software issue-- matching criteria can be configured that permits an invalid patient match; resulting in overwriting of patient records with information from a different patient. | Elekta Impac Software issued a Important Safety Notice letter dated September 2010 to all customers identifying the software issue and clinical impact. Customer Support representatives will contact each site and modify site configuration or upgrade existing software to the version containing the corrective fix. Customers can contact Impac Software at 800 488-4672.
Z-1394-2010 | Apr-15-2010 | 2010 | GE Centricity PACS-IW Software - GE Centricity PACS-IW software; GE Healthcare Integrated IT Solutions; Barrington; IL 60010. The intended use:The Centricity PACS-IW is used to receive medical images (including mammograms) and data from various imaging sources. Images and data can be stored; communicated; processed and displayed within the system or across computer networks at distributed locations. | Ge Healthcare It | 2 | There is a potential safety issue associated with the use of GE Centricity PACS-IW and PACS Web Diagnostic software related to use of built-in reconstruction functions which may result in incorrect anatomic orientation markers. | An "Urgent Medical Device Correction" letters dated February 18, 2010 to their Centricity PACS-IW software version 3.x customers and Centricity PACS Web Diagnostic software version 3.7.4.x customers, The letters advised the users of the patient safety issues associated with the use of the software when using built-in reconstruction functions, which may result in incorrect anatomic orientation markers. The separate letters provided the users with specific safety instructions for that software version to follow until the software is updated. The customers were requested to complete the enclosed customer reply form, supplying the workstation software or build number, and fax the completed form to 847-939-1479. Any questions should be directed to the Remote Online Center, Field Engineer or the GE Customer Care Center at 800-437-1171.
Z-0138-2011 | Oct-25-2010 | 2010 | Instrumentation Laboratory - ACL TOP CTSPN 000280020Bench top; fully automated; random access analyzer designed specifically for in vitro diagnostic clinical use in the hemostasis laboratory for coagulation and/or fibrinolysis testing in the assessment of thrombosis and/or hemostasis. The systems provide results for both direct hemostasis measurements and calculated parameters. | Instrumentation Laboratory Co. | 2 | Software: New data check; "Multiple Threshold Check"; was added to identify and fail abnormal (bimodal) clot curves | Instrumentaion Laboratory notified the US Distributor/Canada via BCI Service Modification Notice (#10793) on August 13, 2010. The Mandatory software upgrade to Verison V4.3.0 in the United States and Canada will be conducted by BCI service representatives and tracked through the distributor, Beckman Coulter (BCI). International Affiliates/Dealers: Worldwide Regulatory Notices were e-mailed to IL international affiliates and dealers, notifying them of the mandatory software upgrade, providing issue description and a risk summary and the support. The upgrades were to be handled through the local representatives.
Z-0645-2010 | Jan-21-2010 | 2010 | MAC 800 RESTING ECG ANALYSIS SYSTEM - GE Healthcare MAC 800 RESTING ECG ANALYSIS SYSTEM.The MAC 800 is a portable EGG acquisition; analysis and recording system. The MAC 800 is intended to acquire; analyze; display and record information from adult and pediatric populations. Pediatric population is defined as patients between the ages of 0 and 15 years. The MAC 800 is intended to beused by trained operators in a hospital or medical professional's facility environment to record EGG signals from surface electrodes. The basic system shall provide 2 modes of operation: (I) Resting EGG mode and (2) Arrhythmia mode. The basic systems shall print 3; 6-leads of EGG. The device shall be upgradeable to provide software options such as 12-lead ECG measurement and interpretive analysis. | GE Medical Systems; LLC | 2 | GE Healthcare has become aware of an issue with MAC800 laser printouts. When more than one ECG report in the file manager is batch printed in laser print; the potential exists for an incorrect patient identification or name to be displayed on the print out. Misidentification of patient data potentially contributing to delayed or incorrect treatment is then possible. | Consignees were sent a GE Healthcare "Urgent Medical Device Correction" letter dated December 10, 2009. The letter was addressed to Healthcare Administrator/Risk Manager, Chief of Nursing and Director of Biomedical Engineering. The letter described the Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information. GE Healthcare will provide a software upgrade to resolve the issue. A GE Healthcare representative will contact customers to schedule implementation of the correction when it is available for installation at the customer facility. Direct questions to the GE Healthcare Technologies Support Center at 1-800-558-7044, enter Option 2 and ask for Technical Support or your local GE Field Service Representative.
Z-0411-2011 | Nov-18-2010 | 2010 | TRIGEN SURESHOT Distal Targeting System - smith & nephew TRIGEN SURESHOT Distal Targeting System; Trauma Interface; V2.0.2; Trauma Launcher V1.01.1Product is intended to be an intraoperative image-guided localization system. It is a computer-assisted orthopedic surgery tool to aid the surgeon with drill positioning for screws during intramedullary nail implantation. | Smith & Nephew Inc | 2 | Targeting software has been found to provide erroneous targeting when used with the 13 mm TAN/FAN Left nails. | Smith & Nephew Inc. sent an URGENT MEDICAL DEVICE CORRECTION letter by FedEx on October 15, 2010, to all customers. The letter was dated October 4, 2010, and identified the product, the problem, and the action to be taken by the customers. Customers were instructed to perform a software upgrade and complete the form provided and return by fax to (901) 566-7835. For questions regarding this recall call (901) 399-5441.
Z-1322-2010 | Apr-12-2010 | 2010 | Baxter 1550 Hemodialysis Machine - Baxter 1550 Single Patient System Hemodialysis Machines; Baxter Healthcare Corporation; Deerfield; IL 60015; product codes 5M5538; 5M5538R; 5M5575; 5M5551 and 5M5551R. | Baxter Healthcare Renal Div | 2 | Some of the 1550 Hemodialysis Instruments may have replacement circuit boards with incorrect versions of software. | An "Urgent Device Correction" letter dated January 29, 2010 was sent to the direct accounts via first class mail on the same date, to the attention of the Hemodialysis Administrator. The letter informed the accounts that subsequent to a prior device field correction in 2003-2004, some 1550 hemodialysis instruments may have replacement circuit boards with incorrect versions of software. The accounts were instructed to confirm the software version in their systems and replacement boards. Instructions for checking the software version were attached to the letter. A complementary software upgrade will be made available to the account if their 1550 instrument and/or spare circuit board inventory has a software version other than 3.0A or 4.1B. The accounts were also requested to complete and fax back to Baxter the enclosed reply sheet confirming receipt of the letter and listing the serial numbers and software versions of those instruments that do nit have software 3.0A or 4.1B. Home patients were not notified directly, the information was sent to their Dialysis Center. The 1550 operator's manual was revised during the 2003 device field correction. The manual correctly warns the operator to discontinue treatment if the 1550 experiences a FL04, FL07, or FLOg fault code. To make arrangements for the software upgrade or questions regarding the communication were directed to Baxter Global Technical Services at 1-800-553-6898, prompt 3. For additional questions, contact the company Monday through Friday from 8:00 AM to 5:00 PM CST at (Toll Free): 800-422-9837.
Z-1167-2010 | Mar-19-2010 | 2010 | Aquilion 64 CT System - Aquilion 64 CT System; Model number TSX-101 A/H The product is an x-ray device that uses digital geometry to generate a three-dimensional image of the inside of an object from a large series of two-dimensional X-ray images take around a single axis of rotation | Toshiba American Medical Systems Inc | 2 | Toshiba America Medical Systems Inc (TAMS) has initiated a field correction on the Aquilion 32 CT Scanners due to the potential for the bolus tracking software (SureStart) to cause delay of the helical scan; up to 10 seconds; when large volumes of data are being reconstructed on the secondary console. | Customer notification was sent and a software upgrade will be issued to all customers affected by the field correction. If you have questions, please call (800) 521-1968.
Z-0856-2010 | Feb-25-2010 | 2010 | Access Immunoassay Systems - UniCel DxI 800 and 600 Access Immunoassay Systems; UniCel DxC 880i; 860i; 660i; and 680i SYNCHRON Access Clinical Systems with Dxl System Software versions 4.2 or 4.2.1.In vitro diagnostic device used for the quantitative; semi-quantitive or qualitative determination of various analyte concentrations found in human body fluids. | Beckman Coulter Inc. | 2 | This recall was initiated because of software issue was identified in System Software 4.2 and 4.2.1; which may result in the incorrect handling of a wash arm dispense plate motion failure. Instead of generating a red warning event and stopping sample processing; the system will post a yellow caution event and continue to operate with the dispense plate in an elevated position until the instrument | Beckman Coulter issued an "Urgent: Product Corrective Action" letter dated October 28, 2009. Customers were notified of the affected product and how the issue will be resolved. For further information, contact Beckman Coulter Inc. Technical Support at 1-800-854-3633.
Z-2191-2010 | Aug-09-2010 | 2010 | Transesophegeal Ultrasound Transducer - SonoSite TEE; 8-3 MHz Transducer labeled: "Tee Transducer... SonoSite; Inc. 21919 30th Drive SE Bothell; WA 98021".The TEE/8-3 MHz Transducer is designed for 2D; M Mode; color Doppler; pulse wave (PW) Doppler and continuous wave (CW) Doppler by applying ultrasound energy through the esophagus or stomach of the patient into the heart. The TEE transducer is intended to be used on adults only. Backscattered ultrasound energy from the patient's heart is used to form images of the heart to detect abnormalities in structure or motion; to evaluate the velocity of blood flowing within the heart; and to obtain a color depiction of the velocities of blood flowing in the heart. | Sonosite; Inc. | 2 | The TEE transesophageal transducer; when connected to a SonoSite M-Turbo Ultrasound System; may exhibit erratic operation of the Scanplane Orientation Indicator on the system display. The indicator may move erratically and not accurately represent the orientation of the TEE transducer array. | On August 21, 2008, SonoSite began calling 15 consignees who have purchased the TEE/8-3 MHz Transducer and an M-Turbo 1.1 system. The firm provided instruction to discontinue use the system to 11 consignees and left voice messages to the remaining four consignees. On August 21, 2008, SonoSite sent out the URGENT-MEDICAL DEVICE CORRECTION letter and the Urgent Field Safety Notice to all affected consignees in the US and in foreign countries. The URGENT-MEDICAL DEVICE CORRECTION letter describes the recall product as TEE/8-5 MHz Transducer, SonoSite Part Numbers P05183 and P06841. The letter informed the consignees that the TEE transesophageal transducers may exhibit erratic operation of the Scanplane Orientation Indicator on the system display when it connected to the SonoSite M-Turbo Ultrasound System. The letter instructs consignees to discontinue use of the TEE/8-3MHz Transducer with the M-Turbo (or Micromaxx system) immediately, while the firm is working to create a software solution to resolve the issue and offers customers a software upgrade within 6-8 weeks. The consignees are also advised to forward the recall notice to others who need to be aware of the issue or where the potentially affected devices have been distributed. Customers should call the Technical Support at 1-1-425-951-1330 or 1-877-657-8118 from Monday to Friday, 6 am-5 pm PST or e-mail to service@sonosite.com for any questions regarding the recall.
Z-0003-2011 | Oct-04-2010 | 2010 | BladderScan BVI 9600 with Aorta Scan Mode - BVI 9600 Bladder Volume Instrument; Aorta Scan ModeThe BladderScan BVI 9600 with AortaScan Mode is a user-selectable; dual-function ultrasound device that projects ultrasound energy either into the lower abdomen to obtain an image of the bladder for measuring bladder volume; or into the mid-abdomen to obtain an image of the abdominal aorta for aortic diameter measurements. | Verathon; Inc. | 2 | The BladderScan BVI 9600 with AortaScan devices may experience the loss of a factory installed software configuration file which results in a loss of the aorta measurement function. | Verathon sent an URGENT Medical Device Recall Notification letter dated August 12, 2010 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Verathon provided an addendum to the User Manual with instructions on how to reset the instrument. Customers were advised to: " Add the enclosed addendum to their current Users Manual. " Update the application software of the device to v2.7.0 or higher. If the customers had questions or needed support for the software update process, they were to contact Verathon Medical Customer Care at (425) 867-1348 or (800) 331-2313 or cservice@verathon.com.
Z-0479-2011 | Nov-26-2010 | 2010 | Philips Easy Diagnost Eleva Conventional R/F system - Product's marketing brochure is labeled in part: "Philips Easy Diagnost Eleva Conventional R/F system...PHILIPS". manufactured by Philips Medical Systems Gmbh; Hamburg; Germany. | Philips Medical Systems North America Co. Phillips | 2 | Patient and image data can be exchanged when a patient is selected in the patient list and the examination tab is pressed with background query running at the same time. There is a risk that the data may then be linked to the incorrect patient; which may adversely affect diagnosis and treatment. | Philips Medical Systems issued URGENT-Device Correction letters dated January 27, 2009 to consignees describing the issue with the Easy Diagnost Eleva GXR RF systems. The consignees were informed of the potential for incorrect diagnosis or treatment if an incorrect image is used for diagnosis. Consignees should contact the Philips Healthcare Call Center at 800-722-9377, #5, #2 and reference "FC0 70600030" with any communication.
Z-0005-2011 | Oct-04-2010 | 2010 | BladderScan BVM 9500 - BladderScan BVM 9500The BladderScan BVM 9500 is intended to project ultrasound energy through the lower abdomen to obtain an image of the bladder and measure urinary bladder volume and mass noninvasively. | Verathon; Inc. | 2 | The BladderScan BVM 9500 devices may experience the loss of a factory installed software configuration file which results in a loss of the aorta measurement function. | Verathon sent an URGENT Medical Device Recall Notification letter dated August 12, 2010 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Verathon provided an addendum to the User Manual with instructions on how to reset the instrument. Customers were advised to: " Add the enclosed addendum to their current Users Manual. " Update the application software of the device to v2.7.0 or higher. If the customers had questions or needed support for the software update process, they were to contact Verathon Medical Customer Care at (425) 867-1348 or (800) 331-2313 or cservice@verathon.com.
Z-1998-2010 | Jul-13-2010 | 2010 | O-Arm Imaging system - O-Arm Imaging SystemThe O-Arm imaging system is a mobile X-Ray system designed for medical applications in surgical theaters; focusing on 3D imaging but also offering 2D fluoroscopic Imaging. | Medtronic Navigation; Inc. | 2 | The possibility that the real time image display can freeze up during use of a 2D continuous fluoroscopy. | The firm, Medtronic, Inc., sent an "Urgent Field Safety Notice" letter dated May 10, 2010, to all custmers. The letter describes the product, problem and actions to be taken by customer. The firm recommended that the users of the O-Arm Imaging System reveiw the information provided to ensure patient safety remains uncompromised during use of the product. The customers were instructed to discontinue use and contact your local Medtronic Navigation representative or call our Technical Services Group at 800-595-9709 or (702) 890-3200, if the image is not updating as expected. A Medtronic Navigation representative will be contacting you within the coming months to schedule a time to evaluate and service the potentially affected components within your system. A service visit will also be scheduled starting November 2010 to update your system with a software based solution that will prevent this issue from occuring. If you have any questions about this notification, please contact your local Medtronic Navigation representative or call our Technical Services Group at 800-595-9709 or (702) 890-3200.
Z-0836-2011 | Dec-29-2010 | 2010 | FastPlan Versions 5.5 and 5.5.1 - Varian Medical Systems FastPlan software; Versions 5.5 and 5.5.1; Model Number HZ3; manufactured by Varian Medical Systems; Palo Alto; CATreatment Planning System is intended as aid in quickly and precisely planning the radiation treatment of small cranial lesions such as an asteriovenous malformations; pituitar tumors; pincalomas; acoustic neuromas and malignant neoplasms. | Varian Medical Systems Oncology Systems | 2 | Due to software Issue; an error occurs when performing cone planning using MRI; which may result in the patient receiving a higher radiation dose than expected; potentially to healthy tissue. | Varian sent Urgent Medical Device Correction letters on March 23, 2010, by Federal Express to all customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to run test cases or otherwise determine whether the MR images that they receive have the 'Series Date' parameter populated. Customers should contact their MR system manufacturer to get a confirmation that the 'Series Date' field is populated. Customers were instructed to confirm by visual inspection that the fusion results are loaded appropriately in Cone Planning before starting to create or edit a plan. For questions regarding this recall contact your local Varian Customer Support Distirct or Regional Manager. Phone: USA and Canada - 1-888-VARIANS (888-827-4265) Europe -+41 41 749 8844 E-mail: North America: support-america@varian.com Australia/New Zealand: support-anz@varian.com Europe: support-emea@varian.com South East Asia: seasia.apps.helpdesk@varian.com China/Asia: china.apps.helpdesk@varian.com Japan: Japan.Apps.Helpdesk@varian.com Latin America: soporte.al@varian.com Internet: Oncology Systems customer site - www.myvarian.com Varian Medical System public site - www.varian.com.
Z-2579-2010 | Sep-27-2010 | 2010 | Paradym Defibrillators - Sorin Paradym CRT-D Model 8750 and Sorin Group Paradym sonR CRT-D Model 8770.Made in Italy by Sorin Biomedical CRM S.r.l. Via Crescentino s.n. ; 13040 Saluggia (VC) Italy | ela Medical Llc | 2 | The Paradym CRT 8750 and CRT sonR 8770 device models have a software anomaly; which will cause the device to lose the ability to sense/pace and to deliver further therapy. The software anomaly could only occurs under a specific sequence of events. | Sorin Group issued a "Medical Device Correction" letter dated June 7, 2010 to consignees. The letter described the product, the problem, and actions to be taken by customers. A Sorin Group representative will contact customers to update the software. Prior to receiving the software update, customers should program the exercise AV delay to 40 ms, which will trigger a reset in the event of an anomaly. This will restore normal operation after shock delivery. If the anolmaly occurs on an implanted unit, a RV coil or SVC coil continuity test should be started, which will ensure that the device recovers proper operation. Corin Group can be contacted at 1-615-788-1807.
Z-0414-2011 | Nov-18-2010 | 2010 | Philips IntelliVue Clinical Information Portfolio - Philips IntelliVue Clinical Information Portfolio (ICIP); Clinical Management SystemSoftware Version D.00 through D.03. | Philips Healthcare Inc. | 2 | Software: Scheduling function of the Philips Intellivue Clinical Information Portfolio Releases D.00 through D.03. Edits to a scheduled even frequency order may result in pending interventions not being generated after the first 24 hours. | Philips Healthcare issued an Urgent Medical Device Correction letter dated October 15, 2010 to customers. Customers are asked to follow the guidelines in the "Action to be taken by Customer/User" section of the Urgent Medical Device Correction Notice until they receive their software correction. The correction will be in the form of a software upgrade and will be performed by Philips service personnel. Contact Philips Customer Care Center with questions or concerns about this correction at 1-800-722-9377.
Z-0797-2011 | Dec-21-2010 | 2010 | Vital; Vitrea fX and Vitrea Enterprise Suite fX - Vital; Vitrea fX and Vitrea Enterprise Suite fXModel number: Vitrea fX 2.1; fX 3.0; and fX 3.1Vitrea Enterprise Suite (VES) fX 1.2 and VES fX 1.3The Vitrea 4D CT Brain Perfusion option is intended for post processing based on dynamic CT images continuously acquired during the infection of contrast; for the visualization of apparent blood flow in brain tissue and pictorial illustration of perfusion-related parameters to aid in the assessment of the type and extent of cerebral perfusion disturbances. | Vital Images; Inc. | 2 | Vital Images Inc.; reported that a software defect exists with Vitrea fX 2.1; 3.0 and 3.1 CT Brain Analysis whereby the annotated numeric ROI values on a batched perfusion map are incorrect when created using the auto batching function (from the right click menu) with the auto ROI templates enabled. | Vital Images, Inc sent a Vital "Urgent Software Update Alert" letter dated 14 October 2010 to all Consignees/Customers. The letter described the product and problem. They advised, until new software is installed, that the customers can avoid obtaining incorrect values by tuning off ROI templates prior to selecting the batch option from the right click menu in the CT Brain Analysis application; and recommended that the customers re-examine any ROI perfusion values that were previously generated using automatic collage batching with ROI templates. Please direct any questions or comments to Customer Support at support@vitalimages.com or at 1-800-208-3005. Any help needed to install this software update can also be directed to Customer Support.
Z-2270-2010 | Aug-25-2010 | 2010 | Intesys Clinical Suite (ICS) G2 92810 Smart Disclosure. - Intesys Clinical Suite (ICS) G2 92810 Smart Disclosure. The Smart Disclosure product collects patient waveforms; alarms; vital signs and 12-lead reports. A 24-hour waveform acquisition is standard; with an upgrade option to 72 hours. The product provides basic trends in both tabular and graphical format; along with a retrospective ECG analysis. The analysis may be defined by patient; and the data may be presented in multiple formats; including a rhythm; a histogram; and a summary view. Standard reports are available from Smart Disclosure; including saved events; disclosure; trends histogram and summary. | Spacelabs Healthcare; Incorporated | 2 | Potential for device to print out the wrong waveform data for the time indicated in the printout or display. | On 07/17/2008, Spacelabs Healthcare began sending the MEDICAL DEVICE CORRECTION letter to their consignees. The letter informs the consignees that there is an issue involving ICS G2 Smart Disclosure (Model 92810) which causes the display or printing of an incorrect waveform for the time indicated in the printout or display. The consignees are advised whether to discontinue to use the device until the firm upgrades the system. If the consignees decide to continue to use the device, then they should discontinue any use of the "Refresh" button. When they print or view waveform reports,t hey should utilize the Saved Event view functionality of the product rather than Arrhythmia, Alarm or Waveform Views. If the printing from the Arrhythmia, Alarm, or Waveform views, the consignees must visually inspect the print out and verify that the waveform contains what you expected to see printed. The consignees could contact Spacelabs Healthcare at 1-800-522-7052 or 1-425-657-7200, x5089 for any questions about the corrective actions of the affected device. The firm's field service engineer will make a schedule a mutually agreeable time with the consignees to update the ICS G2 system at free of charge.
Z-0100-2011 | Oct-20-2010 | 2010 | SoftLab version 4.01 ASCII Lab/Mic Addon - SoftLab version 4.01 ASCII; ASCII Base 3.1.3.0-3.1.3.32; 3.1.4-3.1.4.27; 3.1.6-3.1.6.4; 4.0.0.0-4.0.0.9; LAB/MIC Add-on 4.0.1; 4.0.1.0-4.0.1.13; 4.0.2.0 in 2006. SCC Soft Computer; 5400 Tech Data Drive; Clearwater; FL 33760. | SCC Soft Computer | 2 | For SoftLab version 4.01 ASCII; the qualifying test result does not reach the Call List because the order is locked by either another application; or two Hosparams (HISRR_RBS_sync and Auto_sync_RBS). This creates a lock on the RBS server. The problem was identified in the RBS "C" rule which sends qualifying results to the Call List for SoftLab. All affected clients received the upgrade for Soft | SCC Soft Computer issued a Risk-to-Health notification with Client Notification Verbiage to the affected client using SCC Soft Computer's proprietary Task Management System in June 2006. SCC recommended SoftLab version 4.01 ASCII users change program settings to prevent locking of the RBS server. A software upgrade with corrective patch will be delivered electronically to all clients. Client was asked to acknowledge receipt of the task and grant permission to load the required software.
Z-1436-2010 | Apr-26-2010 | 2010 | iView GT - iView GT Electronic Portal Imaging Device; used with radiation therapy treatments. | Elekta; Inc. | 2 | Image problems that have an effect on patient position corrections and on the accuracy of the radiation treatment. | Elekta sent IMPORTANT NOTICE number A329, dated 31 July, 2009 to all customers. This corresponds with Field Change Order - Documentation 764 00 551 060 "An Important Notice (A329) has been issued for replacement of iViewGT R3.4 Service Pack 2 (SP2) installation CD". Mandatory Field Change Order 764 00 551 059 "Replacement of iViewGT R3.4 installer disk"supersedes FCO 764 00 551 053.
Z-2422-2010 | Sep-14-2010 | 2010 | DigitalDiagnost X-ray system - DigitalDiagnost X-ray System is a stationary Diagnostic X-ray using Digital Solid State Imaging Detector. It is identified with Part Number 4512 201 04752 Software Version 1.5.1 and Part Number 4512 201 04753 Software Version 1.5.2. The device is manufactured by Philips Medical Systems DMC GmbH; RontgenstraBe 24; 22335 Hamburg; Germany. | Philips Medical Systems North America Co. Phillips | 2 | When using DICOM print functionality; the printed images may contain data for the wrong patient. | Philips Healthcare issued an "URGENT PRODUCT CORRECTION " letter dated March 10, 2008 to their consignees. The letter described the product, problem and action to be taken by the customer. The letter instructed the customer to: 1) Validate that image and patient information are correct on hard copy prints. 2) Reprint images that have mixed patient information, and re-check that image and patient information are correct. 3) Do not use printed images if the information is incorrect. The firm will issue the Field Change Order (FCO) 71200025 and the field service engineers will visit each consignee and conduct a software upgrade to solves the printing problem. The consignees may contact the Philips Call Center at 1-800-772-9377, #5 #2 and reference "FCO 71200025".
Z-1316-2010 | Apr-09-2010 | 2010 | ABBOTT ARCHITECT System Software - ARCHITECT System Software v5.00 and v5.10; List Number 5F48-22 and 5F48-23; Part Number 7-204717-01 and 7-204717-02 when used with PM Algorithm LLS Brd; part number 7-94255-01 and PM LLS kit; U10 Chip and Ejectors; part number 7-200788-01.The Abbott ARCHITECT System is designed to perform automated: Chemistry tests utilizing photometry and potentiometric technology; and immunoassay tests utilizing CMIA detection technology. | Abbott Laboratories; Inc | 2 | ARCHITECT System Software v5.00 and v5.10 are incompatible with ARCHITECT LLS Board. Software patch re-establishes correct parameters. | Technical Service Bulletins dated September 11, 2009 were issued to affected consignees. The correction was issued via the Technical Service Bulletins (TSB) to provide a software patch to re-establish the correct parameters on the LLS board. Direct questions to Abbott Laboratories, Inc. by calling 1-972-518-6000.
Z-1735-2010 | May-28-2010 | 2010 | Horizon Medical Imaging Systems - Voiceware clip software update to certain Horizon Medical Imaging Systems. McKesson Medical Imaging Company. Richmond; Canada V6X 3G5.Intended to receive; transmit; store; retrieve; display; print and process digital medical images; digital medical video and associated medical information from various medical imaging systems. | Mckesson Medical Immaging | 2 | Voice Clip issue: When the user is recording voice clips in Horizon Rad Station; there is a potential for voice clips to be saved to the incorrect study of the same patient. | McKesson Medical Imaging Company issued an "Advisory Notice" dated May 7, 2010. Consignees were notified of the affected product and advised to contact McKesson Customer Support department to review their sites current configuration and install the appropriate product software update(s). For further information, contact McKesson Medical Imaging Customer Support at 1-800-663-2533. On June 19, 2010 the recall was amended to include the following actions to follow: Instead of highlighting the anchor study in the Patient Portfolio Sidebar before saving t he voice clip in the Patient Portfolio Sidebar, McKesson recommends that you save the voice clip to the desired study first, before browsing through other documents. A software update that prevents the problem from occurring is currently under development for all affected releases and will be available in approximately three months. Consignees will be notified when the software update is ready to be installed on their system.
Z-0126-2011 | Oct-22-2010 | 2010 | EndoTool Drug Dose Calculator - EndoTool Drug Dose Calculator-Software Version 7.2.1800.3 and 7.2.1825.5 | Hospira; Inc. | 2 | The EndoTool software did not always provide an audible alert when a blood glucose measurement was due. This could result in a patient not receiving an appropriate dose of insulin. | Hospira issued an Urgent Device Field Correction letter dated July 23, 2010 to consignees, providing details of the potential issue and the associated patient risk. Hospira is finalizing software enhancements which will resolve this situation within the next 90 days ,and will contact each consignee when it is available to be downloaded from the Hospira website. The notification letter also included a clinical bulletin, outlining mitigation instructions for clinicians to use until the software enhancement is available. Consignees were requested to complete and return an attached Reply Form by fax. Customers can contact the Hospira EndoTool support line at 1-877-349-4582 for clinical inquiries about this recall.
Z-1600-2010 | May-13-2010 | 2010 | Radiological Image Processing Software - Vitrea 5.1. Vitrea fX 3.0; Vitrea Enterprise Suite 1.2 with the licensing of CT lung or CT Colon. | Vital Images; Inc. | 2 | Vital images; Inc is issuing a recall notice on their medical device software in the use of CT Lung and CT Colon in Vitrea 5.1; Vitrea fX 3.0 or Vitre Enterprise Suite 1.2. because incorrect measurements may result. Questions or comments are directed to the firm's Customer Support at support@vitalimages.com or at 1-800-208-3005. | Consignees were sent a Vital "Urgent Medical Device Software Recall Notice" letter dated 23 February 2010. The letter was addressed to "Dear Customer." The letter described the problem, product affected and provided recommendation to their customers. Questions or comments are directed to the firm's Customer Support at support@vitalimages.com or at 1-800-208-3005.
Z-1661-2010 | May-26-2010 | 2010 | GE Healthcare Innova - GE Healthcare; Innova 2100IQ (The Digital Fluoroscopic Imaging System consists of an a monoplane positioner; a vascular or cardiac table; an X-ray system and a digital detector) | Ge Healthcare; Llc | 2 | A potential software failure related to the Worklist functionality has been identified on a subset of Innova 2100IQ; 3100IQ and 4100IQ systems. When using the Worklist function to import patient data to Innova; there are two scenarios thatwill cause the next exam to fail and the inability to recall acquired sequences. | Consignees were sent on 3/24/10, a GE Healthcare " Urgent Medical Device Correction" letter dated March 23, 2010. The letter included the Safety Issue, Affected Product Details, Product Correction and Contact Information.
Z-0067-2011 | Oct-15-2010 | 2010 | SoftPath GUI Releases 3.17 through 4.1.0. - SoftPath GUI Releases 3.17 through 4.1.0.SCC Soft Computer 5400 Tech Data Drive Clearwater; FL 33760.Manufacture/Distribution Dates:Version 3.17.5 10/14/2003. Version 3.17.6 03/28/2003. Version 3.17.7 07/02/2003. Version 3.17.8 03/03/2004. Version 4.1.0 07/03/2003. Version 4.1.1 10/03/2003. Intended for use in the Pathology Lab for receiving and accessioning specimens; documenting tissue processing; producing material labels; generating final diagnostic reports; and creating billing and management reports for surgical; GYN; NON-GYN; and Autopsy cases. | SCC Soft Computer | 2 | SCC Soft Computer; Clearwater; FL issued a Correction for SoftPath GUI Releases 3.17 and 4.1 in March 2007 for clients with high volume output of final reports who use the Batch Print for Final Reports in the Scheduler; there is a potential for some reports not printing. This issue was corrected for the one affected user in October 2003. | SCC Soft Computer sent a SAFETY NOTIFICATION TASK VERBIAGE dated January 5, 2004, to one client. The Notice identified the product, the problem, and the action to be taken by the client. SCC recommended the client set the Search Date Range to a small enough range to assure that the number of cases scanned during the job would be less than 10,000. The client was asked to acknowledge receipt of the task and to indicate if the Workflow or functionality was used/not used and Agree to use the alternative working solution. For any questions regarding this recall call (727) 780-0100, ext. 4235.
Z-2113-2010 | Jul-29-2010 | 2010 | BleaseSirius Anaesthesia Machine - BleaseSirius Anaesthesia System; gas machine. Blease Beech House; Chiltern Court Asheridge Road; Chesham; Bucks HP5 2PX; EnglandIntended for use in the hospital environment and operating room. | Spacelabs Healthcare; Llc | 2 | Potential for the touch screen on the ventilator portion of the device to go blank when touched. Possibility that this failure may also cause mechanical ventilation to stop. | On 06/24/2010, Spacelabs began mailing the URGENT-MEDICAL DEVICE CORRECTION letter to their US consignees and e-mailed the same letter to their foreign consignees. The letter describes the touch screen of the BleaseSirius Anesthesia machine goes blank when touched and the possibility that this failure may cause the mechanical ventilation to stop. The consignees are asked to weigh the benefits versus the risk when deciding to use the device or not until the firm upgrades the device. The letter also instructs the consignees to notify their staff about the possibility of the display failure to prevent any adverse event that may occur. The letter states that their field service engineer will contact consignees to schedule a time to visit and upgrade all affected devices at no cost. The US consignees are advised to call 1-800-522-7025, Select 2, and the foreign consignees are advised to call 1-425-657-7200 x5089 for technical support or questions regarding the recall.
Z-0168-2011 | Nov-01-2010 | 2010 | Volcano s5/s5i imaging system software version 3.2 - Volcano s5/s5i imaging system Family Software version 3.2; Part Numbers 435-0602.01 to 05; including ChromaFlo and Virtual Histology; manufactured by Volcano Corporation of Rancho Cordova; CA. | Volcano Corporation | 2 | Software defect results in improper image brightness; which may result in introduction of artifacts; which may cause misclassification of results and affect classification of tissue types. This may result in misdiagnosis and inappropriate treatment. | Volcano issued a Voluntary Recall of Volcano s5 Family Software version 3.2 letter dated June 30, 2010, identifying the software defect and actions to be taken by Volcano. Field Service Engineers will visit each site and reinstall the previous software version, which will resolve the issue. Customers can contact Volcano at 916 281-2790.
Z-0004-2011 | Oct-04-2010 | 2010 | AortaScan AMI 9700 - AortaScan AMI 9700 The AortaScan AMI 9700 is an ultrasound device that projects ultrasound energy into the mid-abdomen to obtain an image of the abdominal aorta for aortic diameter measurements | Verathon; Inc. | 2 | The AortaScan AMI 9700 devices may experience the loss of a factory installed software configuration file which results in a loss of the aorta measurement function. | Verathon sent an URGENT Medical Device Recall Notification letter dated August 12, 2010 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Verathon provided an addendum to the User Manual with instructions on how to reset the instrument. Customers were advised to: " Add the enclosed addendum to their current Users Manual. " Update the application software of the device to v2.7.0 or higher. If the customers had questions or needed support for the software update process, they were to contact Verathon Medical Customer Care at (425) 867-1348 or (800) 331-2313 or cservice@verathon.com.
Z-1747-2010 | Jun-02-2010 | 2010 | Seno Advantage 2.2 workstations with software version 22_01 and 22_02 - GE Healthcare; Seno Advantage 2.2 workstations with software version 22_01 and 22_02Seno Advantage is a medical image review station that allows easy selection; processing; filming and media interchange of multi-modality images from a variety of diagnosis imaging systems. When interpreted by a trained physician; mammographic images displayed on the high-resolution monitors may be used as an element for diagnosis. Furthermore filmed images from all modalities may also be used as an element for diagnosis. | Ge Healthcare; Llc | 2 | GE Healthcare has recently become aware of inaccurate Scale factor annotation provided by SenoAdvantage 2.2 workstations that may impact patient safety.When using the Print Image or Print Screen feature on SenoAdvantage 2.2; the Scale factor annotation printed on hard copy images is inaccurate. Secondary Capture (SCPT) images displayed on the screen and Secondary Capture hard copy images may a | The firm, GE Healthcare, sent an "Urgent Medical Device Correction" letter dated February 8, 2010 to customers. The letter described the Safety Issue, Affected product details, Safety Instructions to prevent further inaccurate scale factor annotation, Product Correction and Contact Information. GE Healthcare will correct all affected workstations by providing a software upgrade and an updated Operator Manual. The GE service representative will contact customer to arrange for this correction. If you have any questions or concerns regarding this notification, please call one of the following phone numbers: United States: 800-437-1171, Japan: -0120-055-919. For other countries, please contact your local GE Healthcare Service Representative.
Z-1604-2010 | May-14-2010 | 2010 | LEKSELL GAMMAPLAN - Leksell GammaPlan 8.0 image fusionLeksell GammaPlan is designed for use with the Leksell Gamma Knife manufactured by Elekta Instrument AB. the Leksell GammaPlan is intended to be used for planning the dosimetry of treatments in stereotactic radiosurgery and stereotactic radiation therapy. It processes the inputs of the health professions (Neurosurgeons; Radiation therapists; Radiation Physicists) such that the desired radiation does is proved by the Leksell Gamma Knife to a precisely defined target area within the cranium. | Elekta; Inc. | 2 | The precision of the calculation used to create the fused study in LGP 8.0 is too low and should not be used until the system is upgraded to LGP 8.2. | Phase 1 - Important Notice and Technical Note #200 058 - "Accuracy problems in Leksell GammaPlan 8.0 image fusion", dated December 10.2007 and Field Change Order #200 058, "Technical Note: Accuracy problems in Leksell Gamma Plan 8.0 image fusion", dated January 7, 2008 were sent to affected customers and all field service engineers as notification of the problem. The notice describes the product, problem and actions to be taken by customers. The firm recommended that the customers not use the fusion function until their Leksell GammaPlan(R) is upgraded to version 8.2. For further question, do not hesitate to contact your Elekta support respresentative at.
Z-0852-2010 | Feb-25-2010 | 2010 | Cybernet Medical MedStar Telemedicine System - Cybernet Medical MedStar Telemedicine System; Manufactured by Cybernet Medical; Ann Arbor; MI; Model 533-210. The device is for a remote medical data collection device for the monitoring of patients by healthcare providers. It is intended to work with a weight scale; spirometer; blood pressure unit and pulse oximeter and to transmit data over a phone line. | Cybernet Systems Corp. | 2 | A software problem renders these devices incapable of capturing lung function; oxygen saturation and pulse data collected from multiple tests in accordance with the instructions for use. | The firm sent out a recall letter dated 10/27/09. Consignees were asked to return the devices in prepaid shipping containers. The recalling firm would do a software upgrade and return the device to the consignee. The firm expects it to take 3-6 months to upgrade all devices. Contact Cybernet Medically immediately, at 1-800-CYBERNET, with the serial number and code versions of the devices in question and Cybernet will start the RMA process for returning the devices for the updates, free of cost.
Z-1428-2010 | Apr-22-2010 | 2010 | Leksell SurgiPlan - Leksell SurgiPlan | Elekta; Inc. | 2 | Images from a different patient can be imported without a warning for the user. | Phase I - A technical note for LGP and SPS describing the risk for importing wrong image files and including recommendations on how to reduce this risk has been written and distributed to users of both Leksell GammaPlan and Leksell SurgiPlan on May 23, 2005 via FCO DOC 05-0001. Phase 2 - A maintenance release was created of the Scanner Interface (SIF) to correct the problem. This was distributed to Leksell SurgiPlan users on November 30, 2005 via FCO0000153-1. Additional questions are directed to the firm to Product Manager lrobert.aslund AT elekta.com, or After Sales gustaf.piehl AT elekta.com..
Z-0102-2011 | Oct-20-2010 | 2010 | SoftPath GUI versions 4.2.3.4; 4.2.4.1; 403.0.3; and 4.3.0.4. - SoftPath GUI version 4.2.3.4 2008-06-30 (Manuf Date). SoftPath GUI version 4.2.4.1 2008-09-30 (Manuf Date). SoftPath GUI version 4.3.0.3; 2008-07-23 (Manuf Date). SoftPath GUI version 4.3.0.4 2008-09-16 (Manuf Date). SCC Soft Computer; 5400 Tech Data Drive; Clearwater; FL 33760. | SCC Soft Computer | 2 | For SoftPath GUI versions 4.2.3.4; 4.2.4.1; 403.0.3; and 4.3.0.4.; the Previous Patient History and some QA Management reports may have missing or incorrect diagnosis text. However; the patient final report; HIS resulting and Query display are all correct. All affected clients have been provided corrected software.. | SCC Soft Computer issued a Risk-to-Health notification with Client Notification Verbiage to affected clients using SCC Soft Computer's proprietary Task Management System on January 26, 2009. SCC issued a software upgrade with corrective patch electronically to all clients. Client was asked to acknowledge receipt of the task and grant permission to load the required software. SCC will contact each client to coordinate the running of the utility and the loading and validation of the correct patch.
Z-0124-2011 | Oct-22-2010 | 2010 | Allura Xper FD10 - Allura Xper FD10 Monoplane X-ray system.Intended for use in cardiovascular and vascular x-ray imaging applications; including diagnostic; interventional procedures (such as PTCA and stent placement and atherectomies); pacemaker implantations and electrophysiology. It is compatible with specified magnetic navigation systems. | Philips Medical Systems North America Co. Phillips | 2 | Due to a potential geometry software lock up during a Digital Rotational Angiography (DRA) run; the stand and table movement may no longer be possible. | Philips Medical Systems North America Co. sent an URGENT - Device Correction, dated September 15, 2010, to their consignees. The letter identified the product, the problem, and the action to be taken. The consignees were informed that there was a potential geometry software lock up during a Digital Rotational Angiography (DRA) run, the stand and table movement may no longer be possible. The consignees were advised that as geometry software errors happen intermittently, no actions can be performed by the consignees to minimize the risk except for not performing DRA procedures until the firm upgrades the software. For question consignees were to contact the Philips Call Center at 800-722-9377 #5, #4, #1 and reference FCO 72200111 or FCO 72200125.
Z-1897-2010 | Jun-24-2010 | 2010 | Lenstar LS 900 - Lenstar LS 900; Biomicroscope; Slit-Lamp; AC-powered. HAAG-STREIT AG; Switzerland.Used for obtaining ocular measurements and performing calculations to assist in determination of the appropriate power and type of Intraocular Lens for implantation after removal of the natural crystalline lens following cataract removal. | Haag-Streit USA Inc | 2 | During the creation of the Lenstar LS 900 software; the ability to perform a manual adjustment of the retinal thickness measurement was not disabled for products targeted for US distribution. The current software installed allows the user to manually alter the retinal thickness measurement; which could modify the automatic axial eye length measurement resulting in the selection of an incorrect in | HAAG - Streit USA, Inc. issued notification letters to their customers informing them of the problem. They instructed their customer to refrain from adjusting the gates for the measurement of retinal thickness until after the system has been updated. For further information, contact HAAG - Streit USA, Inc. at 1-513-336-7255.
Z-1534-2010 | May-04-2010 | 2010 | Hitachi Oasis Magnetic Resonance Scanners Software - Hitachi Oasis Magnetic Resonance Scanners Software V1.0CThe intended use: The MR system produces transverse; coronal; sagittal; oblique; and curved cross-sectional images that display the internal structure of the head; body; or extremities. The images produced by the MR system reflect the spatial distribution of protons (hydrogen nuclei) exhibiting magnetic resonance. | Hitachi Medical Systems America Inc | 2 | The device has a software error that could potentially lead to a misdiagnosis. | Sites have already been corrected. Software updated from V1.0C to V2.0B.
Z-0360-2011 | Nov-15-2010 | 2010 | SoftID.NET versions 2.0.0.6 - 2.0.0.19. - SoftID.NET versions 2.0.0.6 - 2.0.0.19. SCC Soft Computer 5400 Tech Data Drive Clearwater; FL 33760. SoftID.NET is to be used in a medical or clinical laboratory setting by knowledgeable; trained; and experienced professionals. SoftID.NET supports bedside collection of patient specimens used in diagnostic testing. This product is integrated with Soft Computer's SoftLab; SoftScape and can also interface with all other SCC laboratory systems and Hospital Information Systems (HIS). | SCC Soft Computer | 2 | SCC Soft Computer issued a Correction for SoftID.NET versions 2.0.0.6 - 2.0.0.19. When using the Cancel functionality; if the collection list refreshes during the transaction and changes the position of the patient on the list; the wrong patient order can be cancelled. All affected clients have been notified. | SCC Soft Computer notified Customers on August 28, 2009, via the proprietary communication software entitled "Risk-to-Health Notification Task Verbiage" dated August 17, 2009. The notification described the product, problem and action to be taken by customer. The customers were instructed to acknowledge receipt of the notification; document a request to load the software, if they wish to receive the correction; document that the Alternative Working Solution will be used by the customers facility until the correction has been taken live. If you have any questions, contact (727) 789-0100.
Z-0364-2011 | Nov-15-2010 | 2010 | SA BASE with SoftLab Version 3.1.4.x. - SA BASE with SoftLab Version 3.1.4.x. Manufactured 03/2003. SCC Soft Computer 5400 Tech Data Drive Clearwater; FL 33760.Intends to provide clients with an interface that links all clinical laboratory departments throughout the enterprise. | SCC Soft Computer | 2 | Functionality issue: For clients using the instrument to generate results on the interface menu; two different result records could be displayed under the same order and sequence numbers. This could cause incorrect results to be populated to a patient record. | Clients were notified via the proprietary communication software on 06/01/2010. The communication discussed the risk-to-health notification task, potential hazard, root cause, corrective action, and requested action. Clients were instructed to acknowledge receipt of the notification. If the client wishes to receive the correction, they should document a request to load the software.
Z-1406-2010 | Apr-16-2010 | 2010 | PrecisePLAN 2.10 Treatment Planning System - PrecisePLAN 2.10 Treatment Planning System | Elekta; Inc. | 2 | Error in dose rates which were calculated at exactly twice the expected values. | As described in a prior related submission, Phase I - "Important Notice A285, Scrolling Selected Frame Dose Calculator Error" dated March 15, 2006 was distributed to affected customers to inform them of the potential and what steps should be taken. Phase 2 -- Software release PrecisePLAN 2.11 was released July 11, 2006 to correct this problem. All customers with PrecisePLAN 2.10 will be upgraded to the revised software PrecisePLAN 2.11 through Field Change Order #772 00 540 052 dated June 5, 2006. Phase 3 -- All customers with PrecisePLAN 2.10 will be upgraded to the revised software PrecisePLAN 2.11 through Field Change Order #772 00 540 052 dated June 5, 2006. The company is directing questions to the local Electa representative or to their offices in the U.K. at telephone 44 (0)-1293-654200.
Z-2209-2010 | Aug-11-2010 | 2010 | Medtronic Carelink Monitor Model 2490C - Medtronic Carelink Monitor Model 2490C. The Medtronic CareLink Monitor Model 2490C is designed to automatically gather information from the implanted heart device. This automatic wireless communication between the heart device and the monitor takes place at times scheduled by the doctor or clinic. The gathered heart device information is sent over a telephone line to the clinic using a toll-free number. The monitor can also send wireless heart device alerts to the clinic automatically. | Medtronic Inc. Cardiac Rhythm Disease Managment | 2 | Medtronic has identified that a subset of 2490C CareLink Monitors recently received an incorrect software update. Patients reported that their monitor was making a "ticking sound" and they were unable to transmit device data via CareLink. These monitors are now non-functional. No patient injuries have been reported as a result of this issue. | Starting on 6/11/2010, Medtronic began sending Australian and Canadian consignees "Medical Device Recall" letters dated 6/10/2010. The letters describe the problem and product involved. The letters instructs consignees to order replacement monitors, notify affected patients, and instruct each patient to perform a manual transmission upon receipt of the new monitor. Medtronic will provide affected patients with a prepaid return kit to faciliate the return of the monitor to Medtronic. Australian consignees should contact Scott Dunlop at 61-7 3025-3547. Canadian consignees should contact Medtronic Patient Services at 1-800-268-5346.
Z-0069-2011 | Oct-15-2010 | 2010 | Hospital Information System Results Reporting Interface Versions 4.0.2.38.1; 4.0.2.39.2; 4.02.40.1; - Hospital Information System Results Reporting Interface Versions 4.0.2.38.1; 4.0.2.39.2; 4.02.40.1; and 4.0.2.41.1.SCC Soft Computer; 5400 Tech Data Drive; Clearwater; FL 33760. Manufacture Dates:4.0.2.38.1 06/15/07 4.02.39.2 06/29/07; 4.02.40.1 07/19/07; 4.02.41.1 08/20/07. The Result Reporting Interface is a highly versatile interface designed to electronically convey clinical laboratory orders and results for the SCC applications to other recipient information systems. | SCC Soft Computer | 2 | SCC Soft Computer; Clearwater; FL issued a Correction for Hospital Information System Results Reporting Interface Versions 4.0.2.38.1; 4.0.2.39.2; 4.0.2.40.1; and 4.0.2.41.1 on 10/05/2007. A software anomaly in the result reporting interface caused result comments entered after an empty line <CR> not to be sent to the HIS system. This included interpretative comments; results; and normal range c | SCC Soft Computer sent a RISK-TO-HEALTH NOTIFICATION TASK VERBIAGE dated October 5, 2007, through their Task Management System to all affected clients on October 17, 2007. The notice identified the product, the problem and the action clients should take to correct the problem. Clients were to acknowledge receipt of the task notice and to identify if the Workflow or functionality is not used and if they Agree to use alternative working solutions. Each client was asked to grant permission to load the required software to correct the issue. For any questions regarding this recall call (727) 789-0100.
Z-0088-2011 | Oct-19-2010 | 2010 | SoftMic Versions 4.0.4.7 and 4.5.1.3 - SoftMic Versions 4.0.4.7; 4.5.1.3. Version 4.0.4.7 Manufactured 02/03/2010. Version 4.5.1.3 Manufactured 03/18/2010. SCC Soft Computer 5400 Tech Data Drive Clearwater; FL 33760. | SCC Soft Computer | 2 | Softmic system may delay or omit reporting of clinically significant results including organisms and drug sensitivities. Incomplete culture results may appear complete. SCC Soft Computer is issuing a Correction for SoftMic Versions 4.0.4.7-4.5.1.3 software due to this potential hazard. | SCC Soft Computer issued Risk-to-Health notifications to clients on July 19, 2010 via proprietary communication software. Corrective actions will be via programming correction via hot fix (software modification) or such so that Global variable is not used. Clients are asked to acknowledge receipt of this task and update the task with permission to load the required correction.
Z-1655-2010 | May-26-2010 | 2010 | Elekta XiO Radiation Treatment Planning System - Elekta XiO Radiation Treatment Planning System | Computerized Medical Systems Inc | 2 | Under certain circumstances; XiO will over-estimate dose in the region where the backup jaw is inside the MLC opening | Elekta issued an undated "Important Safety Notice" to all affected sites via e-mail on 4/20/2010. The notice was also sent to the sites on 4/6/10 and 4/12/10 (for 4 sites) before the firm made the decision to recall. The letter explains the correct functioning of the system, the problem, when and why it occurs, the clinical impact, and the work around. Although not stated in the Safety Notice, a self-addressed postcard was enclosed for the customer to return as evidence they received and understood the notice. Sites that do not return the card will be contacted by phone.
Z-0362-2011 | Nov-15-2010 | 2010 | SoftBank Versions 23.1.2.x; 23.2.0.x; and 25.0.0.x with SoftReports. - SoftBank Versions 23.1.2.x; 23.2.0.x; and 25.0.0.x with SoftReports.SCC Soft Computer 5400 Tech Data Drive Clearwater; FL 33760.23.1:BK040048 23.2:BK080020 25.0:BK090017The SoftBank II application; using SoftScape user interface; is a decision support software device that requires knowledgeable user intervention to document steps and events in a transfusion service. The application provides single and multiple site facilities the ability to manage their transfusion service by integrating patient and unit information. SoftBank uses an SCC proprietary interface called SoftScape that facilitates access to the application. | SCC Soft Computer | 2 | SCC Soft Computer is issuing a Correction for SoftBank II Versions: 23.1.2.x; 23.2.0.x; and 25.0.0.x software applications. Functionality issue - For clients generating patient history reports using Patient>Reports>Batch Reports>History Report option in SoftBank with SoftReports activated for patient reports; when the report is generated for more than one patient ALL of the results for the qualif | SCC Soft Computer notified Customers on June 22, 2010, via the proprietary communication software entitled "Risk-to-Health Notification Task Verbiage" dated June 11, 2010. The notification described the product, problem and action to be taken by customer. The customers were instructed to acknowledge receipt of the notification; to grant permission to load the required hot fix, and to make the manual correction. For the manual correction to the datapath for patient report layouts, your SoftBank Support Specialist will Coordinate those changes and your validation of affected environments. If you have any questions, contact (727) 789-0100.
Z-1402-2010 | Apr-16-2010 | 2010 | GE Centricity PACS-IW Software - GE Centricity PACS-IW software; GE Healthcare Integrated IT Solutions; Barrington; IL 60010. Device that receives; stores and communicates medical images and data from various imaging sources. | Ge Healthcare It | 2 | There is a potential safety issue associated with the use of GE Centricity PACS-IW and PACS Web Diagnostic software when dragging-and-dropping images from the Navigator causing flipping of the images that may impact patient safety. | GE Healthcare issued "Urgent Medical Device Correction" notifications dated February 11, 2010. The initial notification advised users of the patient safety issues associated with the affected product. Separate letters provided the users with specific safety instructions. Accounts were requested to complete the enclosed customer reply form, supplying the workstation software or build number, and fax the completed form to 847-939-1479. For additional information, contact the Remote Online Center, Field Engineer or the GE Customer Care Center at 800-437-1171.
Z-2072-2010 | Jul-22-2010 | 2010 | TomoTherapy Hi-Art System¿; ver. 4.0.0; 4.0.1; 4.0.2. - TomoTherapy Hi-Art System¿; ver. 4.0.0; 4.0.1; 4.0.2; ThomoTherapy Incorporated 1240 Deming Way; Madison; WI 53717. The TomoTherapy HI-ART System¿ is intended to be used as an integrated system for the planning and delivery of intensity modulated radiation therapy (IMRT). The HI-ART System provides precise delivery of radiation to tumors or other targeted tissues while minimizing the delivery of radiation to vital healthy tissue. | TomoTherapy Incorporated | 2 | An issue has been identified; as a result of input received from one TomoTherapy Hi-Art Treatment System customer.With a particular workflow; the Planning Station's Draft and/or final Plan Report can provide incorrect dose statistics. | Consignees were sent on 6/23/10 a TomoTherapy "Urgent Field Safety Notice Medical device Correction" letter dated June 22, 1010. The letter addressed the Issue, Product Affected, Recommended Action and Resolution. Prior to generating a draft or final Plan Report for a given plan, the plan must be closed and reopened at least once. This issue will be remedied by Hi-Art software version 4.0.3, scheduled for release in late 2010. Customers with questions were provided with a contact in the letter at the TomoTherapy Customer Interaction Center. The firm can be contacted at 608-824-2800.
Z-0125-2011 | Oct-22-2010 | 2010 | Allura Xper FD10 - Allura Xper FD10/10 Biplane X-ray system.Intended for use in cardiovascular and vascular x-ray imaging applications; including diagnostic; interventional procedures (such as PTCA and stent placement and atherectomies); pacemaker implantations and electrophysiology. It is compatible with specified magnetic navigation systems. | Philips Medical Systems North America Co. Phillips | 2 | Potential for table and stand not to move due to a software error. | Philips Medical Systems North America Co. sent an URGENT - Device Correction, dated September 15, 2010, to their consignees. The letter identified the product, the problem, and the action to be taken. The consignees were informed that there was a potential geometry software lock up during a Digital Rotational Angiography (DRA) run, the stand and table movement may no longer be possible. The consignees were advised that as geometry software errors happen intermittently, no actions can be performed by the consignees to minimize the risk except for not performing DRA procedures until the firm upgrades the software. For question consignees were to contact the Philips Call Center at 800-722-9377 #5, #4, #1 and reference FCO 72200111 or FCO 72200125.
Z-1974-2010 | Jul-09-2010 | 2010 | Horizon Cardiology Hemo Component Schiller Argus Pro PB-1000 Monitor unit - Horizon Cardiology Hemo Component Schiller Argus Pro PB-1000 Monitor unit.Argus Pro PB 1000 firmware versions 2.41+M NIBP and 2.42.01+M.A programmable diagnostic computer intended for acquiring; digitizing; storing; displaying and reviewing hemodynamic data for use in hospital cardiac catheterization laboratories. The system is cleared for us with a component physiological data collection unit - the Schiller Argus Pro PB 1000 Monitor unit. | McKesson Provider Technologies - Medical Imaging Group | 2 | McKesson has recently discovered that certain Horizon Cardiology Hemo systems were running specific Schiller Argus-Pro PB 1000 firmware versions that might cause a delay in the display of physiological parameters on the Horizon Cariology Hemo Real Time Monitor Screen. | The firm, McKesson, sent an "Advisory Notice" letter dated April 15, 2010, to all customers. The letter describes the problem, products affected, and actions to be taken by customers. The customers were instructed not to use the Emergency Replacement Hemo kit until they receive a replacement Schiller Argus Pro PB-1000 and they have contacted McKesson Support department at 1-866-777-0202. If you have any questions regarding this notice, please call our Customer Support department at 1-0866-777-0202 and ask to speak to your support Manager.
Z-1209-2010 | Mar-25-2010 | 2010 | XiO Radiation Treatment Planning System - XiO Radiation Treatment Planning System; Release 4.51.00. | Computerized Medical Systems Inc | 2 | Inconsistent MLC positions; dose; and monitor units | Elekta contacted customers by phone and mail beginning February 22, 2010, explaining the affected product and providing software update information. For further information, contact Elekta at 1-408-830-8023.
Z-1252-2010 | Apr-06-2010 | 2010 | CareSuite - CareSuite" - Critical Care Manager; PACU Manager and Anesthesia ManagerSoftware: CareSuite 5.1 and later | Picis Inc. | 2 | Under specific timing of conditions and in configuration with 3rd party infusion pumps; an error within the clinical application causes the manual documentation of clinical orders to be changed to an automatic update status. At this point; the order documentation becomes frozen at the last dose and/or rate and the user is unable to modify; stop or delete it. This error could impact the fluid bala | Picis issued letter notification beginning on March 8, 2010 with software correction information. The letter communicates details of the problem, along with recommendations for mitigating the potential to experience this error condition. Use of the product should be discontinued until the update described in the letter is performed. Customers with questions should call the company at 781-557-3000.
Z-0012-2011 | Oct-06-2010 | 2010 | Essenta DR X-ray system - Essenta DR is a digital multifunctional X-ray systems; suitable for all routine radiographic exams; including specialist areas like trauma or pediatric work; excluding mammography.It is designed for radiographic examination of the standing or seated patient or the recumbent patient in combination with a mobile X-ray table (trolley).The intended use is from direct digital imaging using the built in flat panel detector and in addition for free exposures on radiographic cassettes. | Philips Medical Systems North America Co. Phillips | 2 | When a patient is selected in the patient list and examination tab pressed with background query running at the same time; the wrong patient may be selected. | Philip Medical System sent an "URGENT-Device Correction" letter dated August 13, 2008 to their customers. The letter describes the recalled product as Essenta DR Patient and Image Data Software (SW) release 1.0. The firm informs the customers that when patient is selected in patient list and examination tab pressed with background query running at the same time, there is a risk that the wrong patient is selected on the Essenta DR Computed Radiography System. Customers are advised to verify the selected patient's information before releasing an exposure while the firm works on upgrading the Software. All affected units will receive a software upgrade to correct the issue. Customers should contact Philips Customer Care Service at 800-772-9377, #5, #4, #1 and reference FCO 712000031 for any questions or support concerning the recall.
Z-1955-2010 | Jul-06-2010 | 2010 | SSH-880CV Aplio Artida System - SSH-880CV Aplio Artida System | Toshiba American Medical Systems Inc | 2 | Toshiba America Medical Systems (TAMS) is initiating a field correction on the SSH-880CV Aplio Artida System due to a problem related to a wrong measurement value display in using with DICOM viewer.When an image is transferred from this system in DICOM transfer and a measurement is performed on a Doppler image in the DICOM viewer; measured values such as VTI (Velocity Time Integral) and PI/RI | The software will be replaced. To accomplish this, Toshiba America Medical Systems (TAMS) will notify consignees via USPS return receipt mail. The letter contains a return reply form that is to be faxed to TAMS for retention. For questions or concerns, please contact your Toshiba representative or the firm at (714) 730-5000.
Z-0123-2011 | Oct-22-2010 | 2010 | GE Centricity Laboratory Instrument Interface Software - GE Centricity Laboratory Instrument Interface software; GE Healthcare Integrated IT Solutions; Barrington; IL 60010. The intended use: the Centricity Laboratory is used to facilitate the general clinical; anatomic pathology and cytology laboratory workflow such as order entry; results entry; instrument interfacing; results reporting and patient record retrieval. | Ge Healthcare It | 2 | Instrument may incorrectly report the result of an antimicrobial sensitivity test. | The firm, GE Healthcare, sent "Urgent Medical Device Correction" letters dated August 27, 2010 to GE Centricity Laboratory software customers on the same date. The letters advised the users of the patient safety issue associated with the use of GE Centricity Laboratory instrument interface to Biomerieux's Vitek analyzer where results may be incorrectly reported. The letters provided the users with specific safety instructions to follow until the software is updated. The customers were requested to complete the enclosed customer reply form and fax the completed form to 847-620-2780. Any questions should be directed to their GE Service Representative or the GE Help desk at 888-778-3375.
Z-2590-2010 | Sep-30-2010 | 2010 | Beckman Coulter UniCel DxC 880i; 860i; and 600i Systems - UniCel DxC 880i; 860i and 600i Systems; Software Part Number: A84500; A86646; Softwareversion: 4.9.01 The UniCel Synchron Clinical Systems are fully automated; computer-controlled clinical chemistry analyzers designed for the in vitro determination of a variety of general chemistries; therapeutic drugs; and other chemistries. Analysis can be performed on serum; plasma; urine; or cerebrospinal fluid (CSF) and whole blood (sample type is chemistry dependent) | Beckman Coulter Inc. | 2 | Beckman Coulter is initiating a recall on their UniCel DxC Clinical Systems (880i; 860i;680i and 660i Systems with serial number <2760 that have software version 4.9.01 installed and have the VME Motorola ICS board) because they may exhibit issues with stat sample rack loading; causing a possible delay in results. | Beckman Coulter sent an URGENT: FIELD SAFETY NOTICE letter, dated March 25, 2010, to all customers. The letter identified the affected product, the problem, and the actions customers take. Customers were instructed to: if the UniCel DxC system serial number was less than 2760 the customer should contact their local representative to arrange for an upgrade to software version 4.9 build 02. Complete and return the enclosed response form within 10 days of the notice. Share the information with their laboratory staff and retain the notification as part of the Quality System documentation. If affected products were forwarded to another laboratory, provide a copy of the notification letter to them. Customers may contact the Customer Support Center (Hotline) at 1-800-854-3633. For questions or comments please contact the firm at (714) 961-3634.
Z-1177-2010 | Mar-22-2010 | 2010 | Radrex-I DRAD-3000E; - Toshiba America Medical Systems Inc (TAMS) Digital Radiography System Radrex-i DRAD-3000E; Electrostatic x-ray imaging system | Toshiba American Medical Systems Inc | 2 | Toshiba America Medical Systems Inc (TAMS) has initiated a field correction on the DRAD-3000E; Radrex-I; Electrostatic x-ray imaging system due to a software bug that results in a loss of image. | An Urgent - Medical Device Correction letter was sent to customers. The letter described the issue, the corrective action, and the avoidance. Toshiba America Medical Systems informed customers that software was being prepared to address the issue and planned to send representatives to each customer affected by the field correction to provide a software upgrade to the DRAD-3000E; Radrex-I, Electrostatic x-ray imaging system. Customers were to complete the attached customer reply form. Questions were to be directed to Aaron Ybarra, X-ray Product Manager at 800-421-1968, ext. 2492.
Z-1176-2010 | Mar-22-2010 | 2010 | Image Consultant (IC) Software - Image Consultant (IC) Software | R.O. Gulden & Co.; Inc. | 2 | Software not validated; inaccurate results may occur. | On 2/5/10, the recalling firm issued letters via certified mail to its consignees informing them of the problem and to cease use of the software for diagnosis.
Z-2237-2010 | Aug-19-2010 | 2010 | NeuViz Dual series Computed Tomography (CT) Scanner System - NeuViz Dual series Computed Tomography (CT) Scanner System; X-Ray System. Part Number: 989605651321. Intended to produce cross-section images of head and body bycomputer reconstruction of X-ray transmission data taken at differentangles. | Neuisys; LLC | 2 | The reference lines for image generated for a Surview scan may appear in the incorrect position in the "FILM" display mode. | On 07/21/2008, NEUISYS Imaging Systems Solutions distributor in Greensboro, NC began sending the URGENT DEVICE CORRECTION letter on behalf of Philips and NeuSoft Medical Systems (PNMS), Co., Ltd in China. The letter informs the consignees that there is potential safety problem in the NeuViz Dual series CT scanners due to the malfunction of "Incorrect Reference Line in the Film". When a "Head first, Couch out" or a "Feet first, Couch in" Surview scan is started then paused midstream, the image generated is what the doctor expects, but the reference lines appear in the incorrect position in "FILM" display mode. The acquired images accurately represent the exact positions planned on the original view. The incorrect/unmatched reference line in film may result in incorrect diagnosis. The consignees are advised not to stop or pause during Surview acquisition while PNMS in China will provide the updated software patch to all affected systems free of charge. The Field Change Order #FCOP-08-CT-NMS-008 SP4 will be implemented by NEUISYS Imaging Systems Solutions in Greensboro, NC., who is a distributor for PNMS. Consignees are advised to e-mail Service Support Department of PNMS at helpdesk@pnms.neusoft.com regarding the recall.
Z-0159-2011 | Oct-28-2010 | 2010 | Elekta Impac Software Sequencer - Elekta Impac Software Sequencer Verify & Record System; for MOSAIQ; manufactured by Elekta Impac Software; Sunnyvale; CA. The intended use of Sequencer (generically referred to as a verification system) is to assist in the process of process of patient treatment on a radiotherapy treatment machine. | Impac Medical Systems Inc | 2 | Mistreatment-- The database conversion utility used during the upgrade errantly changes the start gantry angle to zero. Any affected treatment fields; if not already treated; would start delivering from zero instead of the original prescribed value. | ELEKTA IMPAC SOFTWARE sent a USER NOTICE dated February 19, 2010, to all customers. The notice identified the product, the problem, and the action to be taken by the customers. Customers were urged to verify the start angle of the gantry against the treatment field definition prior to delivery. Review all treatment plans and charts with conformal arcs. If a treatment plan with an errant gantry start angle was identified, customers were to cease using that plan immediately. A workaround was to be re-imported into their current version of MOSAIQ. IMPAC Support was to follow-up with each site and execute a script to detect and correct all affected treatment plans. For questions, concerns, and requests for upgrade, customers were to email support@impac.com or call their support center.
Z-1231-2010 | Apr-05-2010 | 2010 | Radionuclide Radiation Therapy System - Leksell Gamma Knife Perfexion. Radionuclide radiation therapy system. Elekta; Inc. Norcross; GA.Indicated for use in the stereotactic irradiation of intracranial structures. | Elekta; Inc. | 2 | There has been an issue with "Image Fushing" where low precision calculation caused images to become inaccurate. | Elekta, Inc. issued a Field Change Order, Reference Number 0000946 dated March 14, 2008. Consignees were informed of the affected product and instructed on necessary steps to take. For further information, contact Elekta, Inc. at 1-770-300-9725.
Z-0133-2011 | Oct-22-2010 | 2010 | Computed Tomography X-Ray System MX 8000 IDT 16 - Computed Tomography X-Ray System MX 8000 IDT 16; Model # 728120; Serial #'s: 2513; 2518; 3147; 3217; 3014; 3050; 3068; 3188; 3091; 3051; 3099; 3126; 3183; 3303; 2539 & 3056.Produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. | Philips Medical Systems (Cleveland) Inc | 2 | Philips Healthcare has decided to recall the Computed Tomography X-Ray System MX 8000 IDT due to the fact that the software upgrade for the MX 8000 IDT 16 v4.2.1 may not provide the optional collimation of 4x4.5 (18mm). If the customer does not verify the increment and thickness coincide with the newly selected collimation suggested by the conversion report; there is potential for a 6mm gap to ap | Philips Medical Systems (Cleveland), Inc. sent an URGENT - Field Safety Notice dated July 29, 2010 and August 23, 2010 to all customers. The notices identified the product, the problem, and the steps to be taken. Customers were instructed to locate and correct the error. Field Service Engineers were to visit each site to update the software For further information customers were to contact their local Philips Customer Care Center at 1-800-722-9377.
Z-1609-2010 | May-18-2010 | 2010 | Odyssey - Odyssey Intended use: Radiation treatment planning system | PerMedics; Inc. | 2 | Anomaly 1; Invalid Characters in Patient ID Error:Currently; Odyssey considers a valid ID to be comprised of numbers and/or letters. If a patient ID in a study contains a character other than a letter or number (including spaces). Odyssey removes the invalid character when moving it to a patient folder. Due to this; it is possible that two different patient studies could be listed within Odyssey | A letter dated March 23, 2010 was mailed to customers acknowledging the anomalies and informed the reader that the mandatory version correcting the software will be provided by May 19, 2010. Customers with version 4.4 will receive 4.4.04:t hose with version 4.5 will receive 4.5.04; and those with version 4.6 will receive 4.6.01 containing the corrections to these issues. In the meantime, the following recommendations are provided. For Anomaly I . Invalid Characters in Patient ID Error, it 1s recommended that the users do not create patient IDS that contain characters other than letters or numbers. For Anomaly 2.0dysse.v Muchitle Library Collimator Type Error, it is recommended that the customer contact PerMedics ifany modifications to the library have been performed. The primary view should also be examined and used to detect and avoid the conditions causing the anomaly. Finally, independent field shape and monitor unit checks should be performed as part of a routine quality assurance program. For Anomaly 3, OOdyssey MLC Block Error, it is recommended to avoid using the MLC Block feature until the mandatory version correcting this anomaly is installed. If you need to use this feature, then it is recommended that a physical block tray is used during treatment for each beam that contains an Odyssey MLC Block. If customers believe a patient treatment has been affected by any of these issues, they should notify Customer Support as soon as possible at 877-4PerMed (877-473-7633) or support@permedics.com.
Z-0412-2011 | Nov-18-2010 | 2010 | Philips EasyVision MM workstation with software - Philips EasyVision MM workstation with software R10.2 and R11.1The EasyVision MM Workstation is intended for the manipulation and display of medical images; including mammograms. It can show images from different modalities and interfaces to various image storage and printing devices using DICOM or similar interface standards. Device options enable mammography reading; teleradiology; fast demonstration; 3D; etc. Device accessories enable orthopedic pre-operative planning; angiography/cardiology imaging capabilities; and bone densitometry capabilities. Typical users are trained professionals such as physicians; radiologists; nurses; medical technicians and assistants. | Philips Medical Systems North America Co. Phillips | 2 | Review by firm of Philips EasyVision MM workstations found that to assure compliance with 21 CFR 900.12(c)(5) corrected software was needed so that all sites are in compliance. | Philips Medical Systems North America Co. sent an URGENT-Field Safety Notice dated January 30, 2009, to all affected customers. The notice identified the product, the problem, and the action to be taken by the customer. Customers were advised to ascertain the quality of a hard or soft copy explicity before further distribution to make sure that no part of the tissue is covered by image identification information. Philips will issue the Field Change Order (FCO) 83000129 and release a level on EasyVision MM R10.2 L6 and EasyVision MM R11.1 L3 to resolve the issue. Customers could contact Philips Healthcare Call Center at 1-800-722-9377, #2, #3 and Reference FCO 83000129 for any questions.
Z-1143-2010 | Mar-16-2010 | 2010 | FHC microTargeting Guideline 4000 - FHC microTargeting Guideline 4000 ver 1.4 (GL4K) Software;Catalog # MT-GL4K.Intended to assist in functional neurological procedures. | FHC; Inc. | 2 | The raster tab for the GL4K version 1.4 software may assign incorrect depths for waveform recordings and if used to identify the implantation target for the DBS lead; the implant may be placed in the wrong location. | FHC notified customers by telephone and by an Urgent: Medical Device Correction email, dated December 2009, beginning December 28th. Customers have been instructed to keep the instruction with the GL4K until FHC can provide a revision to the software; and, to use only "distance from zero" or "distance traveled" modes (not "distance from target" mode) during procedures using the raster tab. Customers are to notify their FHC representative to confirm receipt of the notification. A software patch has been verified and validated and will be installed on all affected units. The new version of the corrected software is Guideline 4000 version 1.41. Questions should be directed to Jason Moyer, product manager or Keri Seitz, recall coordinator at 207-666-8190.
Z-2115-2010 | Aug-02-2010 | 2010 | TomoTherapy Hi-Art System¿; Version 4.0.0 & 4.0.1. - TomoTherapy Hi-Art System¿; Version 4.0.0 & 4.0.1. TomoTherapy Incorporated 1240 Deming Way; Madison; WI 53717. Intended to be used as an integrated system for planning and precise delivery of radiation therapy; stereotactic radiotherapy; or stereotactic radiosurgery to tumors or other targeted tissues while minimizing the delivery of radiation to vital healthy tissue. | TomoTherapy Incorporated | 2 | In some cases; the patient's diagnostic CT image is narrower than the Hi-Art radiotherapy couch image. | Consignees were sent on 2/22/10 a TomoTherapy Incorporated "Urgent Medical Device Correction Field Safety Notice" letter dated Monday, February 22, 2010. The letter was addressed to TomoTherapy Customer and identified the the affected product, and included a description of the Issue, Product Affected, Recommendation Actions and Resolution. The issue will be remedied by software versions 4.0.2 and later. Questions should be directed towards TomoTherapy Customer Interaction Center by e-mail or telephone.
Z-2194-2010 | Aug-10-2010 | 2010 | EasyDiagnost Eleva - The Philips EasyDiagnost Eleva with DRF. Philips Medical Systems multifunctional x-ray system. | Philips Medical Systems North America Co. Phillips | 2 | Potential for bright artifacts to appear on the patient image from a previous exposure. | On July 9, 2008, Philip Medical System began sending out the URGENT-Device Correction letter to their consignees. This letter describes the recalled product as EasyDiagnost Eleva with DRF system with SW release 1.0. The firm informs the consignees that under special circumstances there is a small possibility that bright artifacts appear on a patient image from a previous exposure. The consignees are advised to take care about the X-ray field limitation. If the detector is not affected by direct radiation, no problem will occur. If direct radiation reaches the detector, the collimation should not be too tight, such that the area of direct radiation is larger than 1.2 inch x 1.2 inch (or 30 mm x 30 mm). The minimum area is needed to detect and prevent memory artifacts. The firm will upgrade any affected products with the Field Change Order (FCO)70600031 and it will be implemented before the end of December 2008. The upgrade software is 1.0.3. The correction would be free of charge and the consignees should contact Philips Call Center at 800-772-9377, #5, #2.
Z-1423-2010 | Apr-20-2010 | 2010 | Syngo Imaging XS - Syngo Imaging XS model number 10496279The intended use: image processing radiological | Siemens Medical Solutions USA; Inc | 2 | images can be overwritten | A "Customer (Safety) Advisory notice" dated February 16, 2010, was sent all affected customers via Update Instructions IM019/10/S. This notice describes the product, problem and the action to be taken by the customers. The customers should immediately instruct their employees of the advisor notice and maintain awareness. If the customer sold the device/equipment and is no longer in possession of it, they should forward the notice and inform Siemens of the new owner. A software update to correct this issue will be distributed and installed by a Siemens Service Engineer on all affected systems under Update Instructions IM020/10/S. If you have any further questions regarding this issue, do not hesitate to contact your local SIEMENS Uptime Service Center (michael.mib.braun@siemens.com).
Z-0074-2011 | Oct-18-2010 | 2010 | Soft Query Results Reporting Versions 2.2.0.6.3 and 2.2.0.7.1 - Soft Query Results Reporting Versions 2.2.0.6.3 and 2.2.0.7.1.SCC Soft Computer; 5400 Tech Data Drive; Clearwater; FL 33760.Soft Query Results Reporting version 2.2.0.6.3 was released for General Availability on 09/11/03.Soft Query Results Reporting version 2.2.0.7.1 was released for General Availability on 01/08/04. | SCC Soft Computer | 2 | Soft Query Results Reporting Versions 2.2.0.6.3 and 2.2.0.7.1 HIS may display incorrect collection data for a procedure or test result. Clients using HBO Star RR interface to query for results on a canceled test may see the previous occurrence of the test with the current date and time displayed as the collection date and time instead of the correct collection date and time sent by the SCC interf | Client Notification Task Verbiage was sent through the Task Management System on December 3, 2003, identifying the software issue, the clinical impact, and corrective action. Customers were informed that the HIS may display incorrect collection data. Customers received the SQRR Service Pack with revised software, which will will not produce inaccurate information if there are no verified results. Customers were to acknowledge receipt of the task notification and indicate whether they would like the corrected software. SCC Soft Computer proprietary Task Management System is used to manage client communication. Clients use this system to request services, receive and provide software updates for any issue. Customers can contact SCC Soft Computer at 727-789-0100.
Z-0089-2011 | Oct-19-2010 | 2010 | SoftPath GUI Versions 3.17 and 4.1 - SoftPath GUI Versions 3.17 and 4.1.SCC Soft Computer 5400 Tech Data Drive; Clearwater; FL 33760.Manufactured: Version 3.17.7.6 10/23/2003 | SCC Soft Computer | 2 | SoftPath Cases with Word Perfect documents that were not signed out in ASCII are printing in GUI as a converted document from Word Perfect to PDF document instead of the current MS Word Document. This affects clients upgrading from ASCII to SoftPath GUI 3.17.7.6 Version. All affected clients have been notified and corrections have been made. | SCC Soft Computer issued Risk-to-Health notifications with Client Notification Verbiage to affected clients using SCC Soft Computer's proprietary Task Management System. beginning on December 30, 2003. Clients were instructed to sign out all Word Perfect documents before upgrading to GUI. The correction for this anomaly is to upgrade to the latest release of version 3.17.7 or 3.17.8. Clients were asked to acknowledge receipt of this Task and indicate whether they would like to take the correction at this time. When the client decides to accept the correction a new task will be created to initiate the patch using the Task Management System.
Z-1892-2010 | Jun-24-2010 | 2010 | Syngo Imaging - Syngo Imaging Version Lines VB20; VB30; VB35 in combination with a RIS. | Siemens Medical Solutions USA; Inc | 2 | Siemens has become aware of a potential malfunction and hazard to patients when using Syngo Imaging version lines VB20; VB30; VB35A in combination with a RIS that is violating DICOM Standard by creating non-unique study instance UIDs. | The recalling firm sent a Customer Safety Advisory Notice, dated 04-06-2010, to customers explaining when the malfunction occurs and the potential risks, steps to avoid potential risk of the issue by ensuring that the RIS works according to DICOM Standard and does not create double study instance UIDs. Siemens Service Reps will contact the consignees in a timely manner to check if there are symptoms for the reported issues. The issue will be resolved with a Software-Update for each affected version, except for version VB20. All customers on VB20 will be upgraded.
Z-0090-2011 | Oct-19-2010 | 2010 | SoftPath GUI version 3.17.6.3.1. - SoftPath GUI version 3.17.6.3.1.SCC Soft Computer 34350 US Highway 19 N; Palm Harbor; FL 34684.Date of Manufacture:Version 3.17.6.3.1 06/03/2003. | SCC Soft Computer | 2 | For clients using the SoftPath GUI version 3.17.6.3.1 Results Reporting interface with parsed text and having the auto correct and auto formatting functions active in MS Word; there is a possibility for some ANSI characters/symbols to be translated incorrectly to ASCII. SCC Soft Computer initiated a Correction in March 2004 for all affected clients. | SCC Soft Computer issued a Risk-to-Health notification with Client Notification Verbiage to the affected client using SCC Soft Computer's proprietary Task Management System on March 11, 2007. SCC recommended turning off the autocorrect and autoformatting functions in MS Word if custom configuration parses text to an ASCII file (HIS) for the Result Reporting interface. If this function was turned on at the time of creating the canned message in the setup file, edit and re-save all canned messages in SoftPath. Client was asked to acknowledge receipt of the task and grant permission to load the required software correction.
Z-0099-2011 | Oct-20-2010 | 2010 | SoftWeb Version 4.0.4.10.1 - SoftWeb version 4.0.4.10.1.SCC Soft Computer; 5400 Tech Data Drive; Clearwater; FL 33760.Manufacture Date: 05/16/2006. | SCC Soft Computer | 2 | If a case number is re-used in SoftWeb with SoftPath; the SoftPath case is linked to the original MRN in SoftWeb. This causes SoftPath cases being linked to the wrong patient in SoftWeb. Incorrect results on the patient with original MRN and no results on the patient with the second MRN (and the reused case number) are displayed in SoftWeb Results Query. SCC Soft Computer issued a Correction for | SCC Soft Computer issued a Risk-to-Health notification with Client Notification Verbiage to the affected client using SCC Soft Computer's proprietary Task Management System on June 8, 2006. SCC recommended SoftPath/SoftWeb users NOT re-use SoftPath case numbers when cases are cancelled. The generic option setting "Specimen Registration - Standard Settings - Minimum Time to Reactivate Case Number" should be set to "0" so the user is not prompted to reuse the same case number. Client was asked to acknowledge receipt of the task and grant permission to load the required software correction.
Z-0307-2011 | Nov-05-2010 | 2010 | Volumetric Infusion Pump - WalkMed Triton Pole Mount Infusion Pump; Catalog 300000.Indicated for delivering infusions of medications; as well as providing total parenteral nutrition; and patient controlled analgesia. | Walkmed Infusion LLC | 1 | Pump door may be in a near shut position; but unlatched; and the "Door Open" alarm may not sound. | All customers were notified by letter on 06/17/2010, describing the affected product and advising them of the possible malfunctioning alarm in the case a user does not follow the instructions as specified in the Triton Operation Manual. The letter also reminded all users to follow the procedures in the IFUs. The firm stated that it was best if customers re-educated all nurses on the proper door closure procedure to help prevent the potential bypass issue from occurring. Attached to the letter were two pages explaining the proper process for closing the pump door. Follow-up phone calls were made to coordinate an upgrade to all affected pumps. Questions or concerns should be directed to Ross Kurz at 720-531-4933.
Z-0653-2010 | Jan-27-2010 | 2010 | EDWARDS AQUARIUS HEMOFILTRATION SYSTEM - Edwards Lifesciences Aquarius System Model GEF09800. | Edwards Lifesciences; Llc | 1 | There have been reports of clinically significant fluid imbalance when a certain level of fluid imbalance is detected the Aquarius will trigger an alarm; however; users are able to override this alarm and continue therapy. It is possible to remove too much fluid from or replace too much fluid to the patient. In extreme cases; this may result in serious injuries or death. | Edwards is continuing a field corrective action related to the possibility of fluid imbalance when using the Aquarius. In order to mitigate some of the risk associated with this problem, Edwards sent notifications to consignees advising of mitigation steps. Consignees were provided: a description of the affected product, problem, and potential hazard; advice on action to be taken; and revised device labeling. Additionally, consignees were notified of a planned software upgrade to prevent users from bypassing the fluid balance alarm more than five times in a 20-minute period. Although Edwards does not believe there is a need to stop using the Aquarius system at this time, clinicians should weigh the risks/benefits to the patients before continuing the use of the device Edwards first provided some of this information to consignees in a Field Safety Notice on March 16, 2009. On January 11, 2010, Edwards provided an Urgent Product Recall letter to consignees which updated and expanded the earlier communication. The FDA classified this as a Class I recall.
Z-1904-2011 | Apr-07-2011 | 2011 | Zoll Auto Pulse Resuscitation System - Zoll Auto Pulse Resuscitation System; Zoll Circulation Inc.; Sunnyvale; CA 94085Cardiac Compressor used as an adjunct to manual CPR. Use of AutoPulse is intended to reduce the impact of rescuer fatigue and enable rescuer to address other patient needs. | Zoll Circulation; Inc. | 3 | Complaints of batteries discharging prematurely and that during use the Auto Pulse Units would cease operating (providing compressions). | The firm, Zoll Circulation Inc., initiated a Corrective and Preventive Action Report on January 28, 2009. According to the firm, the customers were contacted by telephone calls and were provided with a Battery Maintenance program. For further questions you may call (408) 541-2140 ext.341.
Z-1378-2011 | Feb-18-2011 | 2011 | InfusO.R. - InfusO.R. Infusion Pump; an Rx syringe infusion pump; Baxter Healthcare Corporation; Deerfield; IL 60015 U.S.A.; product codes 2L3100; 2L3100R; 2L3100U; 6461500 and 6461500R. These product codes includes refurbished units and InfusO.R. pumps previously sold under the Bard label. Intended for the controlled rate delivery of small volume parental fluids as prescribed by a physician. | Baxter Healthcare Corp. | 3 | Should a battery with the minimum contact diameter (.51 inches) be placed on the battery spring (.53 inches); an electrical circuit will not be completed and the InfusOR will not operate. The reason being that the InfusOR (2L3100) Battery Spring has a flat contact that measures .53 inches ?.01 inches on the inner diameter of the spring coil. The IEC 60086-2 and ANSI 08 standards specify a minimum | Safety Alert letters dated January 11, 2011 were mailed via first class mail to the InfusOR direct accounts, to the attention of the Director of Biomedical Engineering, the Director of Anesthesia and the Director of Materials Management, on the same date, informing them that over-the-counter Energizer C size alkaline batteries should not be used with the InfusO.R. Pumps (Product codes 2L3100, 2L3100R, 2L3100R, 2L3100U, 2L3100N, 6461500 and 6461500R). That particular battery incorporates a feature that may inhibit electrical contact and affect the ability of the pump to power-up. The accounts were requested to check their inventory of InfusO.R. Pumps for the presence of over-the-counter Energizer C size alkaline batteries. The accounts were also requested to complete the attached customer reply form confirming their receipt of the letter and fax it to Baxter at the number provided on the form. Any questions regarding the communication were directed to The Center for One Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm CST. *** Baxter expanded the scope of their action to include the Mini-Infuser System and sent Urgent Device Correction letters dated 3/4/11 via First Class mail to all InfusOR and Mini-Infuser customers who had purchased either of the affected product lines. The letters informed the accounts that the Mini-Infusers use the same size battery and incorporate the same battery contact interface as the InfusOR pump, and that the same feature that inhibits electrical contact in over-the counter Energizer C size batteries may also be incorporated by other battery manufacturers. The accounts were requested to verify that the InfusOR and Mini-Infuser will power-on before use; if the device does not power on, replace the batteries with a different brand of size C batteries and verify that the device will power-on before use; and to place a copy of the letter with each operator's manual for the InfusOR and Mini-Infuser System for future refer.
Z-1722-2011 | Mar-18-2011 | 2011 | defibrillator/monitor - The LIFEPAK 20 defibrillator/monitor and LIFEPAK 20e. Product part numbers: 3202487-xxx; 3202488-000; U3202487-000; 3202488-015; and U3202488-000.The LIFEPAK 20 defibrillator/monitor and LIFEPAK 20e are acute cardiac care devices used by authorized healthcare providers in hospital and clinic settings. It is commonly found on "code" or "crash" cars in hospital hallways; plugged into AC power. In an emergency the user retrieves the cart; unplugging the monitor; and uses the device with the door closed for easy 1-2-3 operation. When the code team arrives the user will open the door and have access to all functions; including ECG monitoring. In addition to functioning as a semi-automatic defibrillator the LIFEPAK 20/20e provides options for: ECG monitoring; manual mode defibrillation; Non-invasive pacing; synchronized cardioversion; pulse oximetry (SpO2); hard paddles (adult and pediatric) and internal paddle accessories; printing; and backup batter power. The units are biphasic waveform devices | Physio Control; Inc. | 2 | A component failure on the AC power supply assembly results in "NO AC POWER" operation.A failure of AC power can result in a delay of defibrillation therapy if backup battery power is depleted. | PhysioControl, Inc. began sending an Urgent Medical Device Correction letter via certified receipt to all affected customers describing the potential issue. Customers were advised to keep the defibrillators in service and to follow the daily Operator's Checklist. All power supplies will be updated. For questions customers were instructed to contact Technical Support at 1-800-442-1142. For questions regarding this recall call 425-867-4000.
Z-1548-2011 | Mar-04-2011 | 2011 | LIFEPAK 12 Defibrillator/Monitor-AC Power Adapter - The LIFEPAK12 Defibrillator/Monitor.The LIFEPAK12 Defibrillator/Monitor can be used in AED mode; manual mode or advisory mode. In AED mode; the LIFEPAK 12 is intended for use by personnel who are authorized by a physician/medical director. In manual mode; the operator determines whether or not a defibrillation shock is needed by looking at the ECG rhythm display. Advisory mode can be useful during monitoring; CPSS (Continuous Patient Surveillance System) runs in the background and alerts the operator when a shockable rhythm is detected. | Physio Control; Inc. | 2 | L1FEPAK 12 Operating Instructions that describethe process for disconnecting the device from AC (wall) power are not always being followed. When theinstructions are not followed and the Power ON button is pushed within 2 seconds after disconnectingthe device from AC power; the service indicator will illuminate and the device may fail to charge the defibrillator and/or initiate pacing. | Physio Control, Inc. began mailing the URGENT-MEDICAL DEVICE INFORMATION letter to their consignees on October 1, 2010. The letter identified the product, the problem, and the action to be taken by the consignee. The consignees were advised to follow the current LIFEPAK 12 Operating Instructions-Section 7.5-AC Power Adapter Operation and wait at least 2 seconds to push the Power ON button after disconnecting the device from AC power. If the instructions are not followed and the Power ON button is pushed within 2 seconds after disconnecting the device from AC power and the service indicator illuminates, it's important to be aware that turning the device off, then back on again, will reset the device and restore full functionality. The complete copy of the current LIFEPAK 12 Operating Instructions can be downloaded from the company's website at www.physio-control.com/products/ defibrillators, Select LIFEPAK 12, Select Documents. The consignees can call the firm's Technical Support Team at 1-800-442-1142, option 5, between 6:00 A.M. to 4:00 P.M. (Pacific), Monday - Friday.
Z-2607-2011 | Jun-22-2011 | 2011 | Eon Mini Neurostimulation Sysem - Eon Mini Neurostimulation (IPG) System Model 3788System is indicated as an aid in the management of chronic intractable pain of the trunk and/or limbs including unilateral or bilateral pain associated with the following: failed back surgery syndrome; intractable low back and leg pain. | Advanced Neuromodulation Systems Inc. | 2 | Defective battery was identified as cause for reports of inability to communicate or recharge the product. | St. Jude Medical sent an "IMPORTANT MEDICAL DEVICE RECALL" letter dated May 24, 2011 to all affected customers. The letter describes the product, problem and actions to be taken by the customers. St. Jude Medical recommends no action be taken if the product serial numbers do not match the numbers listed in the letter. For product remaining in unimplanted inventory, customers are instructed to contact their SJM representative so that product may be returned and replacements be provided. For implanted product that matches the serial numbers, the Medical Advisory Board recommend not to unnecessarily explant the devices if they are functioning as intended. Replacement devices will be provided at no additional cost to the customer. A Product Reconciliation Form was attached to the letter for customers to complete and return any unused product to the St. Jude Medical Neuromodulation Division at 6901 Preston Road., Plano, TX 75024. Questions regarding this recall are directed to the customer's St. Jude Medical Neuromodulation Division Representative.
Z-1701-2011 | Mar-17-2011 | 2011 | Abbott CELL-DYN 3700 Hematology Analyzer CS/SL - Abbott CELL-DYN 3700 Hematology Analyzer CS (Closed Sampler); ; and CELL-DYN 3700 SL (Sample Loader); Manufactured by Abbott Diagnostics Division; Santa Clara; CaMedical Device; hematology analyzer for in vitro diagnostic use in clinical Laboratories. | Abbott Laboratories | 2 | FIRE HAZARD-- Some power supply modules may have incorrect fuses installed; which may result in fires and smoking. | Abbott Laboratories sent Product Correction letters, dated March 22, 2010, to its customers. The letter listed the additional products that are now included in the recall and gave an explanation of the reason for recall. The letter also discussed the safety impact and provided a list of necessary actions for customers. Customers are to check that the ampere rating of the fuse matches the voltage of operation that is currently being utilized with their CELL-DYN 3000 Analyzer. If the fuse is correct, then the steps provided should be followed. If the fuse is incorrect, then customers are to replace the fuse with the correct fuse and follow the instructions provided. Customers are to contact their local Customer Support Representative in the event that they do not have their replacement fuse that was provided with the Accessory Kit. Customers can refer to the Product Information Letter dated April 20, 2010 for additional labeling to assist with the correct fuse replacement. The Customer Reply form should be completed and the letter should be kept with the Operator's Manual. Questions should be directed to Customer Support at 1-877-4ABBOTT. The firm, Abbott Laboratories, sent a "Product Correction" letter dated March 27, 2009 to its customers. The letter included a Customer Reply form and CELL-DYN accessory kits to inform new customers of the necessity of replacing fuses correctly. The letter described the product, problem and actions to be taken. The customers were instructed to check that the ampere rating of the fuse matches the voltage of operation that is currently being utilized with their CELL-DYN 3700 System and follow the instructions provided, and complete and return the Customer Reply form via fax to 1-800-777-0051 or e-mail QAGCO@abbott.com, even if they no longer have the instrument. The firm intends to service items in the field. In April 2010, additional labeling was sent to customers providing additional information when replacing fuses. .
Z-2694-2011 | Jun-29-2011 | 2011 | Tobii C-Series Lithium Batteries - Tobii C-Series Lithium batteries Type TB8A; TB8B; TB12A used with the Tobii C8 and Tobii C12 Speech Generating Communication Devices.Communication device for the impaired and functions as a speech generating device; text or symbols can be used to generate synthesized or digitized speech. Perfect for asking or answering questions; having a conversation. | Tobii Assistive Technology; Inc. | 2 | Lithium Batteries may overheat and become combustible | Tobii Technology AB sent an "URGENT: MEDICAL DEVICE RECALL" letter dated May 25, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. Customers were advised to stop using the batteries and to remove them from their devices. They were instructed to dispose of the batteries in an environmentally responsible manner as hazardous materials. Tobii will replace the affected batteries. Customers were asked to complete and return an acknowledgement form via fax at 781-461-8213 or e-mail at recall@tobiiati.com. Contact the firm by E-mail: recall@tobiiati.com or telephone: 1-800-793-9227.
Z-2752-2011 | Mar-17-2011 | 2011 | Abbott CELL-DYN 3000CS (Closed Sampler); CELL-DYN 3000SL (Sample Loader) - Abbott CELL-DYN 1700 Analyzer; CELL-DYN 1700 CS (Closed Sample).Manufactured by Abbott Diagnostics Division; Santa Clara; CaMedical Device; hematology analyzer for in vitro diagnostic use in clinical Laboratories. | Abbott Laboratories | 2 | Expanded Recall to include these additional products: FIRE HAZARD-- Some power supply modules may have incorrect fuses installed; which may result in fires and smoking. | Abbott Laboratories sent Product Correction letters, dated March 22, 2010, to its customers. The letter listed the additional products that are now included in the recall and gave an explanation of the reason for recall. The letter also discussed the safety impact and provided a list of necessary actions for customers. Customers are to check that the ampere rating of the fuse matches the voltage of operation that is currently being utilized with their CELL-DYN 3000 Analyzer. If the fuse is correct, then the steps provided should be followed. If the fuse is incorrect, then customers are to replace the fuse with the correct fuse and follow the instructions provided. Customers are to contact their local Customer Support Representative in the event that they do not have their replacement fuse that was provided with the Accessory Kit. Customers can refer to the Product Information Letter dated April 20, 2010 for additional labeling to assist with the correct fuse replacement. The Customer Reply form should be completed and the letter should be kept with the Operator's Manual. Questions should be directed to Customer Support at 1-877-4ABBOTT. The firm, Abbott Laboratories, sent a "Product Correction" letter dated March 27, 2009 to its customers. The letter included a Customer Reply form and CELL-DYN accessory kits to inform new customers of the necessity of replacing fuses correctly. The letter described the product, problem and actions to be taken. The customers were instructed to check that the ampere rating of the fuse matches the voltage of operation that is currently being utilized with their CELL-DYN 3700 System and follow the instructions provided, and complete and return the Customer Reply form via fax to 1-800-777-0051 or e-mail QAGCO@abbott.com, even if they no longer have the instrument. The firm intends to service items in the field. In April 2010, additional labeling was sent to customers providing additional information when replacing fuses. .
Z-1738-2011 | Mar-21-2011 | 2011 | TEG 5000 - TEG 5000 Thrombelastograph Hemostasis Analyzer System; a multipurpose system for in vitro coagulation studies; Haemonetics; Haemoscope Division; Niles; IL 60714; list number 07-022a multipurpose system for in vitro coagulation studies | Haemoscope Division of Haemonetics Corp | 2 | Haemonetics has received a report of smoke from the unit with a power supply failure. | Haemoscope sent letters to all of their U.S. accounts on January 24, 2011, via certified mail. The letters listed guidelines regarding proper use of the TEG 5000 System Power Supply to eliminate any potential risks such as electrical shorting or power supply damage. The information, which was printed in the letter, will also be available in an updated TEG User Manual soon to be released. A Caution label with installation instructions was included with the letter, to be applied to the back of the TEG unit. The accounts were requested to complete the enclosed acknowledgement form and fax it back to Haemoscope at 847-588-0455. Any questions were directed to 1-800-438-2834.
Z-2023-2011 | Apr-19-2011 | 2011 | Siemens brand 550 TxT Patient Treatment Table - Siemens brand 550 TxT Patient Treatment Table; part number 7346534; Product is manufactured and distributed by Siemens AG; Medical Solutions; Roentgenstrasse 19-21; Kemnath; Germany 95478.The 550 TxT Treatment Table is to be used as an integral part of the Siemens linear accelerator systems to position the patient accurately according to the treatment plan for various radiation therapy treatment techniques. | Siemens Medical Solutions USA; Inc | 2 | Potential for causing overheating of the power supply during treatment; and potential for unexpected table descent. | Siemens issued a Customer Information letter dated March 21, 2010. There were two safety issues. There were two occurrences of smoke caused by an overheating of the power supply during patient treatment. Also, there had been a report of an unexpected lowering of the 550TxT table in patient rescue mode. With the released kit that accompanies this letter, Siemens has implemented a minor modification on the 550TxT to include an additional fuse within the power supply circuit to address the first issue described above. They also included an additional modification to address the lowering of the table by the integration of additional relays. As soon as the "move enable" button is released, the table brake will be enabled which is controlled by the implemented relays. No further action is required for the continued operations of the 550TxT table. Further questions regarding this recall please call (925) 246-8407.
Z-2068-2011 | Apr-26-2011 | 2011 | Advanced Instrument Model 2020 Multi-Sample Osmometer - Advanced Instrument Model 2020 Multi-Sample OsmometerIntended to measure the osmotic pressure of body fluids. | Advanced Instruments; Inc. | 2 | A poorly soldered power cord may pose a potential electrical shock hazard. | Advanced Instruments Product Service Department contacted each customer on 1/19/11 via an Urgent Product Recall Notice. The letter identified the reason for recall along with the affected products. The letter provided instructions to safely examine the date code of the AC Power Inlet assembly on the affected product. If customers were unable to perform the inspection or had questions concerning the instructions, they were to contact Advanced Instruments Product Service Manager at 781-471-2147 or via email jorges@aicompanies.com. The attached response forms were to be completed and returned as instructed.
Z-3043-2011 | Aug-29-2011 | 2011 | Medtronic SynchroMed II - Medtronic SynchroMed II; Model 8637; (The SynchroMed¿ II Pump is supplied in 20 ml or 40 ml reservoir size.) Sterilized using ethylene oxide. The pump is part of an infusion system that stores and delivers a prescribed drug to a specific site. The implantable components of the SynchroMed II Infusion System include the pump; catheter; and catheter accessories. | Medtronic; Inc. - Neuromodulation | 1 | Medtronic is updating information regarding the potential for reduced battery performance that can lead to sudden loss of therapy in a small percentage of Medtronic Model 8647 SynchroMed II pumps that was communicated with Healthcare providers in July 2009. The purpose of the current communication is to provide updated information regarding the scope and occurrence of this issue and to emphasize | Medtronic Neuromodulation notified physicians with an Urgent Medical Device Correction letter beginning July 05, 2011. A press was issued July 08, 2011. The letter described the issue, severity, recommendations if they notice issue, and patient management recommendations. UPDATE: in November 2011, MDT began to exchange unused pumps in inventory with pumps that contained the new battery. A letter dated November 2011 was left behind at locations in which the MDT representative swapped out the devices.
Z-1224-2011 | Feb-11-2011 | 2011 | Pump in Style Advanced Breast Pump - Pump in Style Advanced (PNSA)Breast Pump; a daily use; double pump; electric powered breast pump; Manufactured exclusively by and for Medela; Inc.; P.O. Box 660; 1101 Corporate Drive; McHenry; IL 60051-0660; the affected breast pumps were sold in the following configurations: a) article 57027 - Kit PNSA Shoulder Bag; b) article 57027T - Kit PNSA Shoulder Bag Tagged; c) article 57036T - Kit PNSA The Metro Bag Tagged; d) article 57043 - PNSA without kit; e) article 57062 - Kit PNSA Backpack; f) article 57062-03 - Kit PNSA Backpack CA/3; g) article 57062T - Kit PNSA Backpack Tagged;h) 57102 - Motor Replacement PNSA Backpack 57062; i) 57103 - Motor Replacement PNSA Shoulder Bag 57027Intended use: A powered breast pump for the removal; collection and storage of mother's breast milk. | Medela Inc | 3 | The performance of the pumps may be compromised due to the potentiometer used on the circuit board. Affected pumps may not turn on; may turn off on its own; may reduce or lose suction during the pumping process; or may make a clicking noise. | This firm, Medela, sent an "Important Device Recall" letter dated January 11, 2011 to all of direct accounts via UPS or US Mail. The accounts were informed that a component on the control board, after a period of usage, may cause the motor to slow down or stop, resulting one or more of the following symptoms: the pump will not turn on; the pump turns off on its own; the pump reduces or loses suction during the pumping process or the pump is making a clicking noise. If the users have a pump that exhibits any of these symptoms, they were requested to visit Medela's website at www.medelareturns.com to expedite shipment of a replacement unit. Questions about the recall process were directed to Medela Customer Service at 1-800-435-8316 or e-mail at medelareturns@medela.com. Distributor/retailers were requested to post the recall notice within their establishment visible to their customer base, and to notify their customers of the recall. Information on the recall was posted on Medela's website www.medelareturns.com on January 11, 2011.
Z-1094-2011 | Feb-10-2011 | 2011 | CELL-DYN Emerald - Abbott brand CELL-DYN Emerald; CELL-DYN Emerald CPU Board Hematology System; List Number: 09H39-01; Field Replaceable Unit (FRU) Numbers: 8701686901; 8701686902;Product is distributed byAbbott Diagnostics Division; Santa Clara. CA and manufactured in France for Abbott Diagnostics Division of Abbott Laboratories; Abbott Park; ILIntended use: The CELL-DYN Emerald is an automated hematology analyzer designed for in-vitro diagnostic use in clinical laboratories. | Abbott Laboratories | 2 | There is the potential for the instrument to not power up due to the flash memory issue. This will result in an error message displayed is "No Memory Available" and has the potential for delay in generating patient results. | The firm, Abbott Laboratories, sent a "Product Correction" letter dated October 15, 2010 to all customers. The letter described the product, problem and action to be taken by the customer. The customers were instructed to install a new printer driver in the software location using the provided instructions and to complete and return the Customer Reply Form via fax 1-800-777-0051 or email to QAGCO@abbott.com even if they no longer have the instrument. If you have any questions regarding this information, U.S. customers should call Customer Support at 1-877-4ABBOTT; customers outside the U.S., please contact your local hematology customer support representative.
Z-0772-2011 | Jan-04-2011 | 2011 | HiSpeed FX/I - HiSpeed FX/I (2200997 with Gantry mo. #2201000) | GE Healthcare; LLC | 2 | A defect with the hardware and software displaying the incorrect dose data. | Important Electronic Product Radiation Warning letter was sent to customers identifying the affected product, explaining the defect and related hazards, giving instructions, and the product correction. Customers are to decrease the CTDI value by a factor of 30% when using 5, 7, 10 mm slice thickness, and by a factor of 20% when using 3 mm slice thickness for more accurate estimate of the CTDI. GE Healthcare will remedy the defect free of charge and provide an updated Operator Manual and configuration update to correct the CTDI display to all affected systems. Questions should be directed to the contact numbers provided in the letter.
Z-2120-2011 | Apr-29-2011 | 2011 | KAPPA 600/700/650/800/900/RELIA/SIGMA/MEDTRONIC 350/ENPULSE/ENRHYTHM/AT500/ADAPTA/VERSA/SENIA IPGS - Medtronic; Kappa 600KDR600;KDR601;KDR603;KDR606;KDR651;KDR653The devices are implantable pulse generators (IPG) used to detect bradycardia events and provide pacing support. Within each product family; there are varying features; rate responsive pacing; and diagnostic data. Medtronic;Kappa 700KD700;KD701;KD703;KD706;KDR700;KDR700V;KDR701;KDR701PTO;KDR701V;KDR701VIT;KDR703;KDR703V;KDR706;KDR706V;KDR720;KDR721;KDR730;KDR731;KDR733;KVDD700;KVDD701Medtronic;Kappa 800KDR801;KDR803;KDR806Medtronic;Kappa 900KD901;KD903;KD906;KDR901;KDR903;KDR906;KDR921;KDR931;KDR933;KVDD901Medtronic; Enpulse DRE1DR01;E1DR06;E1DR21;E2D01;E2D03;E2DR01;E2DR03;E2DR06;E2DR06J;E2DR21;E2DR21PJ;E2DR31;E2DR31J;E2DR33;E2DR33J;E2VDD01Medtronic; Adapta/Versa/Sensia DRADD01;ADDR01;ADDR03;ADDR06;ADDRL1;ADDRS1;ADVDD01;SED01;SED01N;SEDR01;SEDR01N;SEDRL1;VEDR01Medtronic; ReliaRED01;REDR01;REVDD01(approved in the US but has not been sold in US)Medtronic; Vitatron Extension G and E series DRE50A1;E60A1;G70A1(not sold in the US) | Medtronic Inc. Cardiac Rhythm Disease Management | 2 | A rare measurement lock-up condition that may (at a rate of approximately 1 in 18;000 devices) inappropriately trigger ERI/RRT in the Medtronic Dual Chamber Pacemakers: Kappa 600; 700; 800; 900; EnPulse; Adapta; Versa; Sensia. This is caused by a random lock-up of the measurement system hardware that may result in an incorrect battery voltage reading of zero. This issue does NOT impact battery lon | The firm, Medtronic, sent an "Important: Medical Device Correction" letter dated April 2011 to its customers. The letter described the product, problem, and actions taken. Medtronic informed the customers that this issue does NOT impact battery longevity and does NOT require device explant. Currently, the device can be reset to normal operation by a Medtronic representative. Reset devices are no more likely to experience a recurrence of this issue. Medtronic is planning to release a programmer software update that will allow the clinician to reset the device. In addition, Medtronic has consolidated their product performance data with the vision of creating a primary reference tool for performance updates. If you have any questions, please contact your local Medtronic Representative or Medtronic Technical Services at 800-505-4636.
Z-1644-2011 | Mar-14-2011 | 2011 | Plum A+ SIngle Channel Infusion Pumps - Plum A+ Single Channel Infusion Pumps; Hospira; Inc.; Lake Forest; IL 60045; the pumps were sold under the following configurations:a) list 11971 b) list 11973 - software version 10.3c) list 12391 - software version 11.3d) list 20679 - Hospira MedNet Softwaree) list 20792 - DriverThe Plum A+ dual-line volumetric infusion system designed to meet the growing demand for hospital-wide; as well as alternate site and home healthcare; standardization. With their primary line; secondary line; and piggyback fluid delivery capability; the Plum A+ is suited for a wide range of medical/surgical and critical care applications. Full compatibility with LifeCare Plum Series administration sets and accessories and the LifeShield¿ needleless protection systems makes the Plum A+ a convenient and cost-effective infusion system. | Hospira Inc. | 2 | Hospira has received customer reports of the Plum A+ infusion pumps with no audible alarm conditions; which have been associated with failure of the piezoelectric assembly due to component quality issues. | The firm, Hospira, sent "URGENT DEVICE RECALL" letters dated February 14, 2011 to their customers on the same date. The letter described the product, problem and action to be taken by the customers. The Customer/Clinicians were instructed to weigh the risk associated with continued use of the device versus removal from service. If they elect to continue the devices, they were provided with instructions to perform alarm tests prior to each clinical use (during their pump cleaning process). If the alarm is not audible, they were instructed to discontinue use of the pump and contact Hospira Global Product Safety and Complaints at 1-800-441-4100. The customers were requested to contact Hospira Advanced Knowledge Center at 1-800-241-4002 for technical assistance. The accounts were also requested to complete the attached Reply Form and return it to Hospira via fax at 1-888-216-7330.
Z-3016-2011 | Aug-12-2011 | 2011 | AirLife Nebulizer Heater - AirLife Nebulizer Heater; Catalog Code 2M8021The AirLife Nebulizer Heater is part of the heated nebulizer system intended to provide warm bland aerosol therapy to respiratory patients. The system is used in conjunction with a standard oxygen flowmeter; which by providing pressurized oxygen powers the nebulizer creating ultra-fine particles in the therapeutic aerosol. The heater warms the solution before it reaches the aerosol jet in the nebulizer and heats the dilution air before it is entrained into the aerosol stream. An aluminum heat exchanger on the nebulizer cap transfers the thermal energy to a nebulizer solution bottle warming the solution prior to aerosolizing. | Carefusion 211 Inc | 2 | The recall was initiated because CareFusion's Investigation has confirmed the Nebulizer Heater can; over time; develop a failure mode that results in an internal electrical short. The short can result in smoke and sparks being emitted from the device. In the presence of oxygen enriched environments and a fuel source; this failure represents a potentialrisk of flame. CareFusion is voluntarily rec | Carefusion sent a URGENT PRODUCT RECALL letter dated July 15, 2011, to all affected customers. The letter identified the product, the problem, and the action needed to be taken by the customer. Customers were instructed to examine their inventory for the affected product as indicated on the sample shown in the enclosed Attachment A-1 and remove any affected heaters present in their facility. Confirm the total number of affected units in their possession and enter the quantity and respective serial number on the enclosed Customer Response Card (Attachment A-2). Customers were instructed to contact CareFusion's Technical Support Department at 800-554-8933, (6:30 am to 5:00 pm PDST), to obtain an RMA (returned materials authorization) number for the return of the affected product. Upon receipt of their response, CareFusion will ship to their customers new replacement Nebulizer Heaters in exchange for the affected units. Customers were instructed to acknowledge receipt of this communication by completing the enclosed Customer Response Card and faxing or sending a copy by return mail to the address indicated on the Customer Response Card, (Attachment A-2).
Z-2885-2011 | Jul-26-2011 | 2011 | LIFEPAK 500 AED - The LIFEPAK 500 Automated External Defibrillator is a semi-automatic defibrillator that uses a patented Shock Advisory System. This software algorithm analyzes the patient's electrocardiographic (ECG) rhythm and indicates whether or not it detects a shockable rhythm. The LIFEPAK 500 AED requires operator interaction to defibrillate the patient.The LIFEPAK 500 AED is intended for use in the pre-hospital and workplace and community environments. It has been tested to RTCA/DO-160D; "Environmental Conditions and Test Procedures for Airborne Equipment" (refer to Specifications on page 5-15 of the Operating Instructions 3005338-001). The LIFEPAK 500 AED is intended for use by personnel who are authorized by a physician/medical director and have; at a minimum; the following skills and training: -CPR training -AED training equivalent to that recommended by the American Heart Association -Training in the use of the LIFEPAK 500 AED. | Physio Control; Inc. | 2 | The potential reliability issue is the result of a design characteristic of a SRAM (Static Random AccessMemory) electronic component on the main printed circuit board assembly. | Physio-Control, Inc. sent an "URGENT- ACTION REQUIRED" letter dated June 2011 to all affected customers. The letter describes the product, problem, and the action to be taken by the customer. The letter recommends that users who store their device(s) in hard-shelled carrying cases remove them from the case(s)immediately and discontinue the use of the case(s). Additionally, the letter recommends that customers perform regular maintenance in accordance with the Operating Instructions. A Confirmation Sheet was enclosed with the letter for customers to complete and return via fax at 1-866-448-9567. Customers can call the Technical Support Team at 1-800-442-1142, Option 5.
Z-2888-2011 | Jul-27-2011 | 2011 | WarmAir Model 135 Convective Warming Device - WarmAir Model 135 Convective Warming Device; Cincinnati Sub-Zero Products; Inc.Used to prevent hypothermia and/or reduce cold discomfort before; during; and after surgical procedures through convective heat transfer from the controller to a warm-air heated blanket | Cincinnati Sub-Zero Products Inc | 2 | Due to its design the circuit breaker on the device could trip causing it to shut down and not provide heated air to the blanket. | Cincinnati Sub-Zero Products, Inc. sent an 'URGENT Medical Device Recall Field Action Notice" dated July 12, 2011 to all affected customers. The letter includes: description of product, problem, and information to request a repair tubing kit free of charge. It also asks customers to complete and return a "Recall Response Form". For additional information e-mail WA135-2011@cszinc.com or call (800) 989-7373 or (513) 772-8810.
Z-2134-2011 | May-03-2011 | 2011 | Ultraview SL Patient Monitor; Model 91387 - Ultraview SL Patient Monitor; Model 91387. All modules may be affected except Models 90470; 90478; 90496; & 91496.Affected model number is 91387 shipped after February 10; 2010 with PCBA 670-1298-01 (serial numbers beginning with 1387-1xxxxx).The data sheet is labeled in part: "***91387-27 SL2700 Patient Monitor***91387-28 SL2800 Advanced Patient Monitor***91387-38 SL3800 Central Monitor***Spacelabs Healthcare ***Ultraview SL***The Ultraview SL 2700 and SL2800 bedside monitors include interactive networking; graphic and tabular trends; remote view and alarm watch support. They accommodate two single-high parameter modules; use external flat panel displays; and include a DC power supply that provides power for the monitor and up to three Flexoria¿ system interfaces. The SL3800 central monitor includes interactive networking; graphic and tabular trends; remote view; alarm watch; and central alarm manager. The central monitor accommodates a dual-high recorder module; uses an external flat panel display; and includes a DC power supply. In addition to central monitoring; the SL3800 provides bedside monitoring operation through Biomed-level control***".Spacelabs Healthcare patient monitors; functioning as either bedside or central monitors; passively display data generated by Spacelabs Healthcare parameter modules; Flexport interfaces; and other SDLC based products in the form of waveform and numeric displays; trends and alarms. Key monitored parameters available on the models 91367; 91369; 91370 and 91387; when employing the Spacelabs Command Module; consist of ECG; respiration; invasive and noninvasive blood pressure; Sp02; temperature and cardiac output. Additional parameters and interfaces to other systems are also available depending on the parameter modules employed. Spacelabs Healthcare patient monitors are intended to alert the user to alarm conditions that are reported by Spacelabs Healthcare parameter modules and/or other physiologic monitors via Flexport interfaces. These devices determine: a) when an alarm condition is violated; b) the alarm priority (i.e. high; medium or low); c) alarm limits; and d) when to initiate and terminate alarm notifications. The patient monitors are also capable of displaying alarm conditions on other monitors that are on the network through the Alarm Watch feature. Spacelabs Healthcare patient monitors may also function as a generic display or computer terminal. As a generic display or terminal; the patient monitors allow net worked based applications to open Windows and display information on other net worked monitors. Spacelabs Healthcare patient monitors are also designed to communicate with a variety of external devices such as displays; network devices; serial devices; user input devices; audio systems; and local/remote recorders. Spacelabs Healthcare patient monitors are intended for use under the direct supervision of a licensed healthcare practitioner; or by personnel trained in proper use of the equipment in a hospital environment. | Spacelabs Healthcare; Llc | 2 | The circuit that supplies backup power to the modules in the Medical Patient Monitor; Model 91387; is not working. If there is a power interruption; the three-minute backup may not be operational and customers may lose any changes they have made to the alarm settings. | SPACELABS Healthcare sent an Urgent Medical Device Correction letter dated April 1, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to weigh the benefits versus the risks when deciding whether or not to continue to use their monitors until they can be upgraded. Customers were recommended to brief their staff about the possibility of losing changes to the alarm settings and to restore them after a power interruption. Field service engineers will contact customers to schedule a mutually agreeable time to come to their hospital and upgrade their affected patient monitors at no cost to the hospital. US Consignees can call 1-800-522-7025, Select 2 for Technical Support. Consignees outside of the US can call 1-425-657-7200 x5089 for any questions about this recall.
Z-1001-2011 | Jan-31-2011 | 2011 | Varian brand C Series Clinic¿s - Varian brand C Series Clinic¿s (Includes Trilogy; and Novalis Tx); Model Numbers: H14; H26; H27; H29; Distributed by and/or Manufactured by: Varian Medical Systems Inc.; Palo Alto; CA The system consists of two major components; a photon; electron; and diagnostic kV X-ray radiation beam-producing component that is installed in a radiation-shielded vault and a control console area located outside the treatment room. Intended use: The Trilogy Mx" System is indicated for stereotactic radiosurgery and precision radiotherapy for lesions; tumors and conditions anywhere in the body when radiation treatment is indicated. | Varian Medical Systems Oncology Systems | 2 | A circuit design anomaly whereby a faulty balance potentiometer can lead to an undetected electron beam asymmetry. | The firm, Varian Medical Systems, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated December 1, 2009 via Federal Express to consignees/customers. The letter described the product, problem and actions to be taken by the customers and the firm. The customers were instructed to perform a verification of field symmetry on electron beam energies to determine the threshold and report any field symmetry issues to their local Varian service office for attention. The firm, Varian, is developing new functionality to detect the symptoms of this component failure condition under all conditions of electron beam operations and asserting a beam interlock if it is detected. Note: Varian indicated until the solution can be developed, they will conduct an audit of installed Clinacs to identify those the suspect "blue" electron beam balance potentiometers installed, and will execute a program of prophylactic replacement of these potentiometers. If you require further clarification, please contact your local Varian Customer Support District or Regional Manager at phone: USA and Canada-1-888-VARIANS (888-827-4265); Europe- +41 41 749-8844 or email support-emea@varian.com; North America: support-americas@varian.com; Australia/New Zealand: support-anz@varian.com; South East Asia: seasia.apps.helpdes@varian.com; China/Asia: china.apps.helpdesk@varian.com; Japan: Japan.apps.helpdesk@varian.com; Latin America: soporte.al@varian.com; Internet: Oncology Systems - www.myvarian.com; Varian Medical Systems -www.varian.com.
Z-0294-2012 | Nov-28-2011 | 2011 | Health o Meter Professional - Health o Meter Professional Remote Display Digital Physician Scale; Model 498KL; digital patient weighing scale; Pelstar LLC; 11800 South Austin Avenue; Unit B; Alsip; IL 60803; U.S.A.; Made n China; Model 498KLModel 498KL scale with remote display head uses sophisticated microprocessor technology. Each instrument is designed to provide accurate; reliable and repeatable weight measurements and features that make the weighing process simple; fast and convenient. The scale is set up to determine the weight of a stationary patient. The weight can be displayed in pounds or kilograms. The unit can be operated by an AC adaptor (optional) or by 6-AA cell batteries (included). | Pelstar; Llc | 2 | There is a potential for an electrical short when installing or replacing the batteries in the remote display head of the scales. This short may cause the batteries to smoke; burn or burst. | Pelstar LLC. sent a Urgent Device Recall letter dated October 7, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter informed them of the potential of an electrical short caused when installing or replacing the batteries which may result in minor injuries to the person installing the batteries. The customers were requested to immediately examine their inventory and quarantine all product subject to the recall. In addition, if they distributed this product to end users, they were requested to identify all customers and notify them of this product recall at once in writing, using the enclosed recall notification letter. The customers were also requested to complete and return the enclosed mandatory response form as soon as possible. Any questions were directed to Health O Meter Professional at 800-815-6615 or 708-377-0600.
Z-1654-2011 | Mar-16-2011 | 2011 | Allura Xper Systems with Velara Generator - Philips Allura Xper Systems with Velara Generator; x-ray system. FCO72200184. | Philips Medical Systems North America Co. Phillips | 2 | A wire in the connector of the Generator Grid Switch Supervisor box might short-circuit between 12V power and ground; and it will cause a System breakdown. | On 10/15/2010, Philips began sending out the URGENT-Device Correction Notice Failure Causing the Generator to Stop Working notification letter to their consignees. The consignees were informed that a short circuit may occur due to a faulty wire connection in the Supervisor Box Connector in the Generator, and it will cause the generator stops working. When the system fails during a critical interventional case, this may cause risk to patients (no Fluoro available). The systems involved in this recall are Allura Xper with the Velara Generator, Allura Integris with OMCP Generator, MultiDiagnost Eleva, and OmniDiagnost Eleva Imaging system. Consignees can contact Philips Healthcare Call Center at 800-722-9377, #5 and Reference FCO 70800114, 70800115, 72200184, and 72200185.
Z-1363-2011 | Feb-17-2011 | 2011 | Electrosurgical generator. - ConMed System 2450 Electrosurgical Generator; REF 60-2450-120; 120 V; ConMed Electrosurgery; Centennial; CO.A general-purpose electrosurgical generator used in conjunction with an electrosurgical accessory handpiece for delivery of high (HF) electrosurgical current through an accessory electrode for cutting and coagulation at the operative site. | ConMed Electrosurgery | 2 | Electrosurgical generator may allow possibility of electrical shock and/or burns to user or patient. | ConMed Electrosurgery sent an Urgent Device Recall letter dted December 27, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to notify anyone who may have received the affected product.. They were to review their inventory and return a Business Reply Form via fax to 303-699-9854 and to contact ConMed Electrosurgery Technical Services at 800-552-0138 to schedule a time to have their equipment upgraded.
Z-1983-2011 | Apr-14-2011 | 2011 | Maxi Sky 1000 and BHM Medical V10 Bariatric Ceiling Lift - Maxi Sky 1000/BHM Medical V10 Bariatric Ceiling Lift System; a patient transfer lift; ArjoHuntleigh; Addison; IL 60101; Model numbers LF20010; LF20011; LF20013; LF20014; LF20015; LF21015; LF21019; 9120001; 9120004; 9120005; 9120008; 9120009 and 9121009The Maxi Sky 1000 is designed for lifting patients in a homecare setting; at nursing homes and other assisted living centers; as well as in institutional or hospital settings. Patient transferring is performed under the supervision of an appropriately trained caregiver staff; in accordance with the instructions for use. All other uses must be avoided. | Arjo; Inc. dba ArjoHuntleigh | 2 | The battery-powered control board in the ceiling lift motor may overheat; resulting in smoke and possibly flames and the lift becoming inoperable. | ArjoHuntleigh sent an Urgent Field Safety Notice dated March 4, 2011, to all affected customers via UPS Ground on March 21, 2011. The notice informed the accounts of the potential for smoke and flames coming out of the Maxi Sky 1000/V10 bariatric ceiling lifts. The accounts were requested to immediately cease use of the affected lifts, powering them down by activating the emergency stop, and tag the lifts to prevent further use. If there is a clinical need to continue use of the lift before it is retrofitted or replaced, specific instructions were given to reduce the risk of a device malfunction. The accounts were requested to complete the Customer Response Form online, using the enclosed card's instructions. Questions were directed to the ArjoHuntleigh at 1-800-323-1245, ext 57985.
Z-1566-2011 | Mar-07-2011 | 2011 | Terumo Advanced Perfusion System 1 - Terumo Advanced Perfusion System 1; 100/120V System 1 Base; REF 801763; Terumo Cardiovascular Systems Corp.; Ann Arbor; MI.The Terumo Advanced Perfusion System 1 is indicated for use in extracorporeal circulation of blood for arterial perfusion; regional perfusion; and cardiopulmonary bypass procedures only when used by a qualified perfusionist who is experienced in the operation of Sarns or similar equipment. | Terumo Cardiovascular Systems Corporation | 2 | Terumo Cardiovascular Systems (CVS) has received 19 reports of a circuit breaker for the Terumo Advanced Perfusion System 1 tripping and causing the system to switch to battery back-up power. In two cases; the incident was caused when the circuit breaker came into incidental contact with the base of an IV pole; in others there was no known cause. Terumo CVS' investigation has determined that | Terumo sent an URGENT MEDICAL DEVICE CORRECTION letter dated February 14, 2011, addressed to Chief of Perfusion, Department of Cardiovascular Surgery or Director of Operationd Room Services. The letter explains the problem with the circuit breakers and hazards of the problem. The letter states that a service technician will visit the facility to install a guard. The service will be performed at the next scheduled preventative maintenance visit for those consignees with a service contract. For customers without a service contract, Terumo CVS will contact the consignees to schedule an appointment for the correction. Customers were instructed to review the Medical Device Correction notice. Assure that all users are aware of the notice. Confirm receipt of the communication by faxing the attached Customer Response Form to the fax number indicated on the form. For questions or concerns customers were instructed to contact Terumo CVS Customer Service at 1-800-521-2818.
Z-0071-2012 | Oct-19-2011 | 2011 | TAPSCOPE 550 - ***REF #200***TAPSCOPE 550. Each pouch contains one (1) 18 French esophageal stethoscope with bipolar pacing and recording electrodes with a YSI 400 series male connector thermistor. LOT xxxxxxx 400095-0200-C 100902. www.cardiocommand.com. ***REF #200-0010***TAPSCOPE 550 CONTAINS: Ten (10) 18 French esophageal stethoscopes with bipolar pacing and recording electrodes with YSI 400 series male connector thermistors. NON-STERILE. SINGLE USE ONLY. DO NOT REUSE. CAUTION: Federal (U.S.A.) law restricts this device to sale; distribution and use by or on the order of a licensed physician. Manufacturer: CARDIOCOMMAND; INC. Tampa; FL USA. (800) 231-6370 (813) 289-5555. www.cardiocommand.com. LOT xxxxxxx 400155-B 091101.***REF #200-0025***TAPSCOPE 550 Contains: Twenty five (25) 18 French esophageal stethoscopes with bipolar pacing and recording electrodes with YSI 400 series male connector thermistors. ***LOT xxxxxxx***FOR SINGLE USE ONLY. DO NOT REUSE. NON-STERILE. CAUTION: Federal (U.S.A.) law restricts this device to sale; distribution and use by or on the order of a licensed physician. Read Model 2A operator's manual prior to initiating pacing procedures. Manufacturer: Tampa; FL USA 33607. (800) 231-6370 (813) 289-5555. www.cardiocommand.com. 400215-A 100902. | Cardiocommand Inc. | 2 | CardioCommand; Inc. is recalling their TAPSCOPE 550 esophageal pacing/recording catheters because there maybe electrical discontinuity; which can result in the inability to reliably pace and/or record ECG. | On October 25, 2010, an Urgent Medical Device Recall Notice was sent by certified mail to all facilities advising them of the recall. The letter identified the affected product and reason for the recall. Users were asked to discontinue use and return the affected product to the firm. If all product was already used, then this should be indicated on the form provided. The firm will send replacement product in quantities equal to the returned product. The attached form should be completed and returned via fax. Customers were instructed to report any suspected incidents related to this issue to CardioCommand Customer Service at 1-800-231-6370. Also questions or comments should be directed to CardioCommand Customer Service.
Z-0208-2012 | Nov-17-2011 | 2011 | ACT Cellular Phone Monitor - ACT Cellular Phone Monitor; a component of the ACT Ambulatory Cardiac Telemetry; an automatically activated system with a high performance EKG; office hookup and auto send/auto detect that requires no patient intervention to either capture or transmit an arrhythmia when it occurs; Manufactured by Card Guard; 2 Pekeris St; Rehovot; 76100; Israel; there are two models used with either the ACT 1 sensor or the ACT III sensor:a) Model Ozone-HTCb) Model BlackJack2 SGH-i617LifeStar ACT provides critical ECG information by capturing the onset and termination/escape of an arrhythmia to assist in correctly identifying and treating the patient. Patients are monitored for up to 30 days and all of the data is available at the request of the physician during the service period and for 7 days following the end of service. | Lifewatch Services Inc | 2 | LifeWatch received complaints of minor shocks and burns during use of the LifeStar ACT systems with shorts in the leads. | The firm, LifeWatch, sent an " Urgent Device Recall" letter dated September 27, 2011 to the affected customers. The letter describes the product, problem and action to be taken. The customers were instructed to immediately refrain from providing these devices to their customers; stop using the affected sensors, return their entire kit in the pre-paid mailer, and use the new device kit included with the recall letter. In addition, the letter informed the users that the replacement device has upgraded software and firmware that allows for detection of shorts in the cable, and if detected, cuts off current flow to the patient to prevent minor shocks and burns. If you have any questions, please call (847) 813-4258.
Z-0882-2011 | Jan-07-2011 | 2011 | Terumo Advanced Perfusion System 1 - Air bubble detect module; catalog 802110; Terumo Cardiovascular System Corp; Ann Arbor; MI.The air bubble detect module is used with an air sensor to detect gross air bubbles in the perfusion circuit. | Terumo Cardiovascular Systems Corporation | 2 | The firm identified that a component within the System 1 ABD and Occluder Modules that may malfunction. The Air Bubble Detector Module (ABD) is used with an air sensor to detect gross air bubbles in the perfusion circuit. The Occluder module is used with the Occluder head to provide a computer controlled tube clamping mechanism to regulate the flow in the venous line.The firm identified a fai | Terumo Cardiovascular Systems Corporation issued an URGENT MEDICAL DEVICE CORRECTION letter dated 12/2/2010 to the Director of Cardiac Operating Room Services or Operating Room Nurse Manager. The letter identified the product, the problem, and the action to be taken. The letter stated the firm will replace all modules in the affected population, and that a representative would contact the consignee regarding the replacement schedule. Customers were to review the Medical Device Correction. Assure that all user were aware of the notice. Confirm receipt of the communication by faxing the attached Customer Response Form to the fax number indicated on the form. For questions regarding this recall call 1-800-521-2818.
Z-3080-2011 | Aug-29-2011 | 2011 | Model 754/754M Disposable MRI Ventilator Circuit - Model 754/754M Disposable MRI Ventilator Circuit Impact Instrumentation Inc.; 23 Fairfield Place; West Caldwell; NJ 07006.Accessory for 700 Series 754/754m Ventilators during MRI examinations | Impact Instrumentation; Inc. | 2 | Model 754/754 M Disposable MRI Ventilator Circuits are being distributed without market clearence (510k). | Impact Instrumentation Inc. sent an "Urgent Product Recall" dated May 27, 2011 to all affected customers. The letter informed customers the product is not a cleared product and to return the product back to Impact Inst. for credit. A notification form is included in the letter to be return to the firm. For information on this recall contact the firms Customer Service Department at (800) 969-0750, 8:30 am to 5:00 pm EST.
Z-2800-2011 | Jul-14-2011 | 2011 | Model 134; WarmAir Convective Warming Device - Model 134; WarmAir Convective Warming Device; 115V; Part #86180; Device Listing #D111865; The WarmAir 134 is used in conjunction with a single-patient use blanket to warm a patient's body temperature. The device draws in room air through a filter by means of an electric blower. It heats the air to the desired temperature by controlling the air temperature entering a flexible connecting hose; and conveys it to the disposable blanket (i.e. FilteredFlow Blankets; Warming Tube) in contact with the patient; in order to warm the patient. This device provides a continuous flow of heated air as a means of warming patients suffering the common; but significant problems of hypothermia. The system can be used in Post Anesthesia Care Units (PACUs); Intensive Care Units (ICUs); Surgical Intensive Care Units (SICUs); Emergency Rooms (ERs); Operating Rooms (ORs); or any other departments or hospital facilities requiring patient temperature management. | Cincinnati Sub-Zero Products Inc | 2 | A component defect. The recalling firm determined that one of the connectors attached to the firm's WarmAir 134 device's Printed Circuit Board (PCB) can overheat. Should this event occur; the excess heat build up within the device may result in damage to the device's external case and/or melting of some of the internal components of the unit. | CSZ Cincinnati Sub-Zero sent an "URGENT MEDICAL DEVICE RECALL FIELD ACTION" notice to all affected customers. The letter identifies the product, problem, and actions to be taken. CSZ instructs customers to complete and return an enclosed recall response form and to state whether or not their facility is in possession of the device. The completed form should be returned via fax at 513-772-9119 or via e-mail to WA134Recall2011@cszinc.com. Additionally, customers are instructed to discontinue use of the device and return to their Biomedical Engineering and/or Service Department for device exchange and return. Questions regarding this notice are directed to the WarmAir 134 field action (recall) administrator at 1-800-989-7373.
Z-0142-2012 | Oct-31-2011 | 2011 | CDI 500 Monitor - CDI Blood Parameter Monitoring System 500 including all Catalog Numbers to include: Catalog # 500AVHCT with Arterial and Venous blood parameter modules and Hematocrit/Saturation probe. Catalog Number 500AHCT; CDI 500 Monitor with Arterial blood parameter module and Hematocrit/Saturation probe. Catalog Number 500AV CDI 500 Monitor with Arterial and Venous blood parameter modules; and Catalog 500A; CDI 500 Monitor with Arterial blood parameter module.Intended for use during cardiopulmonary bypass procedures when continuous monitoring of arterial and /or venous pH; PCO2; K+; or oxygen saturation; hematocrit; hemoglobin; and temperature are desired. | Terumo Cardiovascular Systems Corporation | 2 | Terumo CVS identified a defective circuit board; which could potentially cause the CDI monitor to not function after a power loss. Once power is lost; the monitor will not power up again; even if AC or battery power are restored. | Terumo Cardiovascular Systems Corporation sent an "URGENT MEDICAL DEVICE CORRECTION" notice dated September 23, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. The letter provides information on what to do in the event of a failure of the device. Additionally, a Customer Response Form was attached to the letter for customers to complete and return. Terumo CVS will schedule the return of affected products for inspection and replacement of defective parts. Contact Terumo CVS Customer Service at 1-800-521-2818 for questions regarding this notice.
Z-3017-2011 | Aug-12-2011 | 2011 | Animas 2020 Insulin Infusion Pump - Animas 2020 Insulin Infusion Pump.Animas Corp.An insulin infusion. | Animas Corporation | 2 | Incorrect component (1nF capacitor) was used to produce 191 printed circuit boards used in six pump lots of which 380 pumps were distributed. This issue may result in the pump displaying a "Replace Battery" alarm not related to the condition of the battery. | Animas Corporation sent an "Urgent: Insulin Pump Recall" letter dated July, 29 20011 to all affected customers. The letter included affected product and problem. They requested anyone who received the pump to send it back to Animas and a replacement pump will be sent at no charge. For additional information contact the Customer Support Department at (877) 937-7867.
Z-2659-2011 | Jun-24-2011 | 2011 | Precision Ice - Precision Ice One-Patient-Use Control Unit and Wrap System (Model # 4BB02BK).The device is intended to reduce edema and promote healing of soft-tissue injuries and lymphedema through stimulation of the lymphatic system. The device is a therapeutic system which consists of a control unit and fluid circulating pads. The pads are mounted in wraps designed for specific parts of the body such as ankle; knee; elbow; and shoulder. The pads provide cold treatment and pressure which help promote healing.The owner's manual is labeled in part: "PRO TRAINERS' CHOICE COMPANY; INC.***www.precisionice.com***info@precisionice.com***Mailing Address PO Box 27***Kingston; WA 98346 USA***Phone: 360-297-3902***Shipping Address***26129 Calvary Lane NE; Suite 200***Kingston; WA 98346 USA***Fax: 360-297-3903".The sticker which is placed on the back of the control board is labeled in part: "MOD#4BB02BK***SER#100***".The shipping box is labeled in part: "Pro Trainers' Choice' LLC***Model #4BB02BK***Serial #100____***".The knee wrap is labeled in part: "***Pro Trainers' Choice Co. Inc./Precision Ice One Patient Use Knee Wrap***#1BK01051KN***This device is to be used for one patient only; do not re-use!***". | Pro Trainers' Choice Company | 2 | The control panel of the Precision Ice One Patient Use Control Unit and Wrap System has a overheating issue. The control panel printed circuit boards were missing a jumper connection across resistors R28 A; B; C; and D. | Pro Trainers' Choice sent a Mandate-Medical Device Recall on April 18, 2011 and April 20, 2011 to all affected costumers. The letter included a description of the product, affected model and lots. It instructed the costumers to discontinue using the Precision Ice, Model 4BBK02BK until the device is returned and corrected by Pro Trainers' Choice. The device should be returned to PO Box 27, Kingston, WA 98346. The customers were instructed to notify their sub-account customers of this recall if the devices were further distributed. Customers can call Pro Trainers' Choice Company at 360-297-3902 or visit the website www.precisionice.com for any questions about this recall.
Z-0972-2011 | Jan-27-2011 | 2011 | Bio Med Devices - Bio Med Devices Disposable Pediatric Breathing Circuit; One SetCatalog Number: 4040LC5 foot /15mm corrugated EVA hoseIntended use: Breathing circuit is a pediatric disposable tubing assembly for use with Lifecare ventilators | Biomed Devices | 2 | Breathing circuit may disconnect between the adapter adjacent tubing cuff | The firm,Bio Med Devices, sent a "URGENT MEDICAL DEVICE RECALL" letter dated October 18, 2010 via fax to the customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed to immediately examine their stock for the affected lots, if found, discontinue distributing and using the product and promptly return via UPS #025106 to Bio-Med Devices 61 Soundview Rd, Guilford, CT 06437, Attention: 4040LC Recall Return. The customers were also instructed to immediately contact their accounts, if the product was further distributed , advise them of the recall situation and have their outstanding product returned, and to complete and return the RETURN FAX-IMMEDIATE RESPONSE REQUIRED form via fax to 203-458-0440. Note: Bio-Med informed the customers that they will credit the return product; however, they only plan to replace the product in an emergency situation, because they have very few of the 3/8" x 15mm adapters left and are looking to discontinue the 4040LC due to the discontinuance of this product by the manufacturer. If you have any questions, please contact (203) 458-0202 x 217.
Z-2042-2011 | Apr-20-2011 | 2011 | COULTER¿ LH 500 Series Analyzer - COULTER¿ LH 500 Series AnalyzerPart number: 178832; 178833; 178834The COULTER LH 500 Analyzer is a quantitative; automated hematology analyzer and leukocyte differential cell counter For In Vitro Diagnostic Use in clinical laboratories. The LH500 Analyzer also provides a semiautomated reticulocyte analysis. | Beckman Coulter Inc. | 2 | The recall was initiated because Beckman Coulter has identified a problem with Integrated Circuit Chips used in different boards throughout the Beckman Coulter LH500/HmX/HmX Autoloader Hematology analyzer. | The firm, Beckman Coulter, sent an "URGENT PRODUCT CORRECTION" letter dated March 2, 2011, with attached PCA Response Form (via US Postal Service for US Customers) to all customers who purchased the COULTERÂ® LH 500 Series Analyzer, COULTERÂ® HmX Hematology Analyzer, and COULTERÂ® HmX Hematology Analyzer with Autoloader. The letter described the product, problem and actions to be taken. Customers were informed that the suspect Integrated Circuit Chips have been removed from all systems in inventory and manufacturing. The customers were instructed to share this information with their laboratory staff and retain this notification as part of their laboratory Quality System documentation. If customers have forwarded any of the affected product(s) listed on the PCA letter to another laboratory, please provide a copy of the letter to them. The customers were instructed to complete and return the enclosed Response Form within 10 days so Beckman are assured they have received this important communication and are taking appropriate action. The customers with any technical questions regarding the Product Corrective Action were instructed to contact Beckman Coulter Customer Service at 800-526-7694 in the United States and Canada or contact their local Beckman Coulter Representative.
Z-2580-2011 | Jun-17-2011 | 2011 | Mettler Traction Decompression System; Model: ME 4000 - MTD 4000 - Mettler Traction Decompression System; Model: ME 4000Product Usage: Provide traction and mobilization of skeletal structures and skeletal muscles | Mettler Electronics Corp | 2 | The recall was initiated because Mettler has confirmed the possible failure of internal mechanical component and software detection for potential failure regarding Electronics Traction Device - MTD4000. The firm are initiating the recall because component failure might result in patient injury. Use of the device should cease immediately. | Mettler Electronics sent an "URGENT MEDICAL DEVICE RECALL" letter dated April 14, 2011 to all affected consignees. The letter identified the product, the problem and the actions to be taken. Customers were instructed to immediately cease use of the recalled device, examine their inventory and quarantine product subject to recall. In addition, consignees were told if they have further distributed the affected product to please identify their customers and notify them at once of this product recall. The letter asks that that each includes a "Recall Return Response Form" that consignees are required to complete and return. If you have any questions please contact Mettler Electronics Corp 714-533-2221 x331, Monday through Friday, between the hours of 8:30 A.M. and 4:00 P.M. Pacific Time.
Z-0352-2012 | Dec-02-2011 | 2011 | Servo-i ventilator system - Maquet Servo-i ventilator systemDevice Part number 64 87 800Intended for treatment of and monitoring of patients in range of neonates; infants and adults with respiratory failure of respiratory insufficiency; used only by healthcare providers in hospitals or healthcare facilities and for in hospital transport. | Maquet Inc. | 2 | Servo-I and Servo--S ventilator systems are equipped with a specific revision of a printed circuit board (PC1772 revision 22) that includes a component that has generated a number of nuisance alarms due to technical error message 24 (TE 24). | The firm, Maquet Getinge Group, sent an " PRODUCT RECALL URGENT - MEDICAL DEVICE CORRECTIVE ACTION" letter dated September 19, 2011, to its customers. The letter describes the product, problem and actions to be taken. Maquet service team will contact the customers within two weeks of receiving this letter to schedule an on-site service of the affected equipment. If you have any questions, call technical support at our toll free number at 1-888-627-8383.
Z-2060-2011 | Apr-25-2011 | 2011 | Luxtec¿ MLX 300 Watt Xenon Light Source - Luxtec¿ MLX 300 Watt Xenon Light Source;Model OOMLX; Surgical/Medical Light SourceThis light source is designed to supply high-intensity light to a fiber optic cable for illumination of a surgical field during a surgical and/or medical procedure. | Integra LifeSciences Corp. | 2 | Malfunction of the MLX Control Board may cause loss of manual control and Lamp shut down. | Integra notified all customers/distributors via written communication, dated 3/08/2011. On 3/09/11, Integra Sales Representatives and Service Technicians were notified via telephone conference call and provided copies of written communications. Customers were advised to immediately discontinue use of any affected product and to place in a secure area. Note: This defect only applies to MLX units that were repaired between 12/01/2009 and 1/31/2011. This defect does not occur in manufactured units that have not been serviced, nor does it apply to MLX units that were serviced outside of this specific timeframe. Customers with devices already known to require repair and/or servicing will be contacted directly to arrange for the return of their device(s). Additionally, sales specialists were instructed to examine each customer's inventory for those device(s) that may potentially have the jumper-wire and to perform a stand-by test to determine if repair and/or servicing are required.
Z-0342-2012 | Dec-01-2011 | 2011 | Synergy XVI kV Generator - Elekta Synergy XVI X-ray Volume Imaging SystemTo be used as part of radiation therapy treatment process. | Elekta; Inc. | 2 | A small number of sites have experienced situations with kV generator transformers overheating due to a diode failure on the interface board. | The firm, Elekta, issued a "FIELD CHANGE ORDER MANDATORY ACTION" notice dated October 20, 2011 to its customers to add fuse protection for the kV transformers. The notice describes the product, problem, provides instructions and procedures and actions to be taken. The customers were instructed to follow any instructions or recommendations covered in the notice and to file the document in the Important Notice section of the appropriate User Manual. Elekta will contact the customer if the FCO is applicable to their kV generator. If you have any questions, call 770-300-9725.
Z-2585-2011 | Jun-17-2011 | 2011 | Integra CUSA Selector NXT Ultrasonic Surgical Aspirator - Integra CUSA NXT Ultrasonic Tissue Ablation System Service ModuleThe CUSA NXT Ultrasonic Surgical Aspirator System is indicated for use in surgical procedures where fragmentation. emulsification and aspiration of soft and hard (e.g. bone) tissue is desirable. including Neurosurgery. Gastrointestinal and affiliated organ surgery. Urological surgery. Plastic and Reconstructive surgery. General surgery; Orthopedic surgery. Gynecological surgery; Thoracic surgery. laparoscopic surgery and Thoracoscopic surgery. | Integra Limited | 2 | The insulated covering of a wire cable that powers the vacuum pump in the CUSA NXT Service Module can become abraded over time resulting in the exposure of bare wire. If this bare wire were to come in contact with the metal surface of the pump; it could result in a short circuit and a loss of aspiration. | Integra sent a "FIELD ACTION NOTIFICATION" dated May 2, 2011 to all affected customers. The letter identified the product, the problem and the actions to be taken. Customers were informed of a Field Corrective Action that is being initiated. All devices (customer owned, Integra Loaner Pool, and Sales demos) will undergo rework of the vacuum pump assembly. The rework will be scheduled, at customer sites by Integra Field Service Technicians or returned to a designated Integra Service Center for rework. A loaner will be presented to customer accounts upon request. The letter asks that each customer complete and return the enclosed Recall Acknowledgement and Return Form to Integra, maintaining a copy of the letter and form for your records. For any additional information please call 609-936-2485.
Z-1005-2011 | Feb-01-2011 | 2011 | Acuson Antares PE - 10032746 Acuson Antares PE Product Version 5.0; manufactured by Siemens Medical Solutions; Mountain View; CA.Intended for Abdominal; Intraoperative; Small Parts; Transcranial; OB/Gyn; Cardiac; Transesophageal; Pelvic; Neonatal/Adult Cephalic; Vascular; Intravascular; Musculoskeletal; Superficial Musculoskeletal; Great Vessel; and Peripheral Vascular applications. | Siemens Medical Solutions USA; Inc. | 2 | ECG waveform does not accurately align with 2D; M-Mode or Spectral Images. | The firm's Customer Safety Advisory notice was mailed to all affected customers. The letter describes when the malfunction occurs and the potential risk. It also recommended that users collect corresponding evidence with spectral waveforms that would confirm or deny findings based on ECG, in order to avoid potential risks for misdiagnosis. The firm planned to exchange the faulty hardware component in the system. Customers would be contacted to schedule for the hardware exchange. Questions regarding this issue should be directed to the local service support person.
Z-3258-2011 | Sep-23-2011 | 2011 | NANOKNIFE - ANGIODYNAMICS NANOKNIFE IRE GENERATOR; Catalog/REF No. 20300101; ANGIODYNAMICS; 603 Queensbury Ave.; Queensbury; NY 12804Intended to remove tissue and control bleeding by the use of high- frequency electrical current (surgical ablation of soft tissue). | Angiodynamics; Inc. | 2 | Some units of Nanoknife have the potential to deliver a shortened pulse width without indicating an error. The affected units contain a "GOMe" circuit board which may cause the generator to drift from the calibration points set at the time of manufacture. | AngioDynamics, Inc. sent a "NOTIFICATION OF MEDICAL DEVICE CORRECTION" letter dated April 4, 2011 to all affected customers. The letter describes the product, problem, and action to be taken. The letter instructs customers to review the measurement graphs (displayed on generator) after each delivery of energy, and to contact AngioDynamics Customer Service at 1-800-772-6446 if shortened pulses are detected. Contact your local sales representative or call Customer Service for any additional concerns or questions regarding this notice.
Z-2631-2011 | Jun-23-2011 | 2011 | 1500T14 Cardiac Ablation System - Generator - 1500T14 Cardiac Ablation System - GeneratorThe 1500Tl4 Cardiac Ablation System operates in conjunction with a commercially available external Disposable Indifferent (Dispersive) Patch (DIP) electrode. The catheter delivers RF power in a monopolar mode between its distal electrode and the DIP electrode. The Generator incorporates a medical grade power supply. This power supply can be used for various line voltages and frequencies without any modifications. Utilizing microprocessor control; the Generator produces a continuous nonmodulated radio frequency (RF) output at 485 KHz. Its front panel displays the command power output; tissue impedance; and tissue temperature. The Generator measures root-mean-square (RMS) voltage; RMS current; and command power output by taking the average value of the product of voltage and current. This reflects the effective heating power delivered to the tissue from the catheter large electrode (tip electrode or other ablating electrodes). Impedance is calculated as RMS voltage divided by RMS current. The amount and duration of RF power delivery is user- selectable. The Generator has four (4) independent channels for monitoring the tissue temperature simultaneously using thermocouple sensor and i-channel (thermistor) for catheters using a thermistor sensor. The desired temperature is user-selectable when busing a catheter with a temperature sensor. If using a catheter with a temperature sensor; the output power of the Generator will shut off if the measured temperature exceeds 80¿C or if the measured temperature exceeds the user-selected temperature set point by more than 5¿C for greater than 3 seconds. If using a catheter with 4 temperature sensors; the output power of the Generator will shut off when the highest temperature detected from anyone of the temperature sensors is exceeds the preset temperature by more than 5¿C for greater than 3 seconds or exceeding 80¿C.The Generator has built-in safety features which include automatic RF power shutoff if the measured tissue impedance falls below 50 ohms or exceeds 300 ohms or the preset impedance value. | St Jude Medical | 2 | The recall was initiated because St. Jude Medical has confirmed its T14 and T12 RF Generatormay malfunction upon encountering an unexpected electric current (e.g. static electricity).The resulting 'malfunction results in a failure of the generator to respond to key pressesproperly; and may cause unexpected energy delivery or inability to discontinue energydelivery with the start / stop but | The firm, St. Jude Medical (SJM), sent a "Product Correction" letter dated May 9, 2011 to it's affected customers. The letter described the product, problem and actions to be taken. The customers were instructed to discontinue use of the T14/T12 generators and remote controls until the units have been properly serviced; complete and return the attached field action form, and utilize the units in accordance with the Operators Manual. In case of system malfunction, the customers were instructed to shut off the RF power delivery by turning off the generator via the rocker switch, dislodging the power cord, or disconnecting of the patient cable or indifferent electrode connection. An SJM sales representative will contact the customers in the near future to facilitate repair or exchange of the impacted unit. If you have questions regarding this notice, contact the Marketing Manager, AF at (905-286-4015), Marketing Director, AF/CRM, at (905-286-4107); Product Manager (Tokyo) 03-6255-5740, and/or your SJM sales representative.
Z-2232-2011 | May-20-2011 | 2011 | MAKO PN 204588 Camera - MAKO PN 204588 Camera204588 Label With Camera Stand Number ***PN 200294 SN CAM XXX 2011-01.*** Weight 86.2kg.***Class I Equipment. Conforms to IEC 60601-1/A2: 1995; EN 60601-1/A2: 1995; UL 60601-1: 2003; CAN/CSA-C22.2 No. 601.1-M90.*** Manufactured in USA*** MAKO SURGICAL CORP. 2555 Davie Road; Ft. Lauderdale; FL 33317.***204589 Robot Arm Label***PN 203999 SN ROB XXX 2011-01.***V 100/120/230 A 9.6/8.0/4.2 Hz 50/60***Weight 39.4kg***Class I Equipment Conforms to IEC 60601-1/A2: 1995; EN 60601-1/A2: 1995; UL 60601-1:2003 CAN/CSA-C22.2 No. 601-1-M90.***Manufactured in USA***MAKO SURGICAL CORP. 2555 Davie Road; Ft. Lauderdale; FL 33317.***204588 Label With Guidance Module Number***PN 201251 SN GUD XXX 2011-01***Weight 51.7kg***Class I Equipment. Conforms to IEC 60601-1/A2: 1995; En 60601-1/A2: 1995; UL 60601-1: 2003; CAN/CSA-C22.2 No. 601.1-M90***Manufactured in USA***MAKO SURGICAL CORP. 2555 Davie Road; Ft. Lauderdale; FL 33317.***THE RIO uses an infra-red camera for the purpose of spatial navigation. This camera is present on the Camera Stand and communicates with the Guidance Module to provide the locations of the tracked instruments (i.e. robotic-arm; patient trackers; probes; etc). The information is then used to provide the virtual representation of the patient anatomy and the tools on the monitor; and provides necessary information to the robotic-arm to ensure proper bone preparation. | Mako Surgical Corporation | 2 | MAKO Surgical Corporation recalled ONE infra-red camera MAKO PN 204854 which was a component of their RIO Robotic Arm Interactive Orthopedic System with SSN P7-03127; on November 22; 2010. This component can function improperly and cause the accuracy of the system to be outside of the specification listed by the manufacturer. The recall was initiated after the camera manufacture; NDI; Waterloo; On | The firm, MAKO, sent a "RECALL NOTICE" letter dated November 22, 2010, to its customer. The letter described the product, problem and action to be taken. The customer was instructed to complete and return the attached Acknowledgement Form to their MAKOplasty Specialist (MPS) via mail to: MAKO Surgical Corp., Attn: Jim Pomeroy, 2555 Davie Road, Fort Lauderdale, FL 33317. The letter stated that "It is important that the camera be replaced immediately, and MAKO is currently addressing the situation with the camera manufacturer." In addition, MAKO will replace the camera with a camera that is not affected by the recall and verify the customers system to meet the functional specifications. If you have any questions or concerns, please contact your MPS or Sr. Director of Quality Assurance at 954-927-2044 ext 604, and /or Sr. Director, Customer Support at 954-927-2044 ext. 442.
Z-0933-2011 | Jan-21-2011 | 2011 | Baxter FloGard 6201 Volumetric Infusion Pump - Baxter FloGard 6201 Volumetric Infusion Pump. Baxter Healthcare Corporation; Medication Delivery Division. Product code 2M8063For delivery of intravenous solutions; drug solutions; enteral feedings and blood. | Baxter Healthcare Corp. | 2 | The FloGard Infusion Pump was released from service without having the Current Leakage test performed. This could lead to the patient/user to receive an electrical shock. | The firm, Baxter, telephoned the one customer on November 17, 2010, informing them that electrical safety testing was not performed on their FloGard pump serial number 11091330FA during the last service. Since the test was not conducted on the pump, the device may not function as designed and could lead to a hazard of electrical shock. Baxter requested that they locate the pump and remove it from use as soon as possible. Baxter advised the customer that a Baxter Field Service Engineer would be sent to their facility to perform electrical safety testing or alternatively, the customer may choose to send the pump to Baxter for testing. If you have any questions regarding this information, please contact Vice President, Quality, GIS at 847-270-5507.
Z-3008-2011 | Aug-11-2011 | 2011 | Colleague CX Infusion Pump - Baxter Colleague CX Infusion Pump; Single Channel Volumetric Infusion Pumps. Baxter Healthcare Corporation; Deerfield; IL; product code 2M8161 | Baxter Healthcare Corp. | 2 | A review of Baxter service records indicates that the Baxter Colleague CX Infusion Pump has not met certain voltage specifications. This issue could potentially cause a delay in initiation of therapy. | The firm, Baxter, called by telephone on July 22, 2011, and sent a follow-up "Urgent Product Recall" letter dated July 25, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately locate and take the affected pump out of service; call Baxter Global Technical Services at 1-800-THE-PUMP (1-800-843-7867), Option 1, between the hours of 6:00 AM and 7:00 PM Central Time, Monday through Friday, if arrangements have not yet been made to return the pump, and promptly complete and return the CUSTOMER REPLY FORM via fax to 847-270-5457, or scan and email to fca@baxter.com. If you have any questions regarding this communication, please call the Center for One Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm Central Time.
Z-2130-2011 | May-06-2011 | 2011 | Synchron LX20 - Synchron LX20; Part Number: 466200.Designed for the in vitro determination of a variety of general chemistries; therapeutic drugs; and other chemistries. | Beckman Coulter Inc. | 1 | The recall was initiated because Beckman Coulter has received reports of incorrect electrolyte results on Synchron LX systems due to a variety of maintenance-related hardware issues. | The recall communication was initiated on 04/05/2011 with Beckman Coulter forwarding a Product Correction letter with attached customer response form to the affected customers via US Postal Service. The letter informed customers of the products affected, the issue, the impact, and the actions to be taken. Customers were instructed to follow the steps provided in order to address the issues related to microbial contamination; and to complete and return the enclosed response form within 10 days so that receipt of this important information could be confirmed. In addition, the letter requested customers share the information with their laboratory staff and retain the notification as part of your laboratory Quality System documentation. For customers who have any questions regarding this update, the letter indicated to contact Customer Support Center Call Center via the Beckman Coulter website at http://www.beckmancoulter.com/customersupport/support/. They may also contact Beckman Coulter at (800) 854-3633 in the United States and Canada or through your local Beckman Coulter representative. Beckman Coulter will issue two additional recall notification ("UPDATE TO URGENT PRODUCT CORRECTION Synchron LX20, LX20 PRO, LXi 725 Clinical Systems") beginning on 05/20/2011, one to Laboratory Directors and one to Physicians. The notices will be mailed via US Postal Service on Friday, May 20th. Customers who have subscribed to email notifications will receive an email containing the letters on Monday, May 23th. Both letter informs the Laboratory Directors and Physicians that Beckman Coulter has received reports of erroneous/aberrant sodium, potassium, chloride and calcium results that have been released from the laboratory. Electrolyte shifts can be low or high and may occur above, below or within the normal range. Laboratory Directors were instructed to immediately disseminate the Attachment to this letter (via notification/alert procedure appropriate to your .
Z-2217-2011 | May-19-2011 | 2011 | Philips Audio-Video Engineering PAS-210 Remote Active Speakers Kit - Philips Audio-Video Engineering PAS-210 Remote Active Speakers Kit used with the IntelliVue Information CenterThe intended use of the Information Center Software is to display physiologic waves; parameters and trends; formal data for strip recordings and printed reports and provide the secondary annunciation of alarms from other networked medical devices at a centralized location. The Information Center Software provides for the retrospective review of alarms; physiologic waves and parameters from its database. An additional intended use is to provide primary annunciation of alarms and configuration and control access for networked telemetry monitors. | Philips Healthcare Inc. | 2 | PAS-210 Kit Remote Active Speaker Assembly used with the IntelliVue Information CenterSpeaker cable connection may lead to intermittent or loss of audio alarm annunciation | Philips Healthcare issued an Urgent - Medical Device Correction letter dated April 28, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to follow the "Action to be taken by Customer/User" section of the instructions while they await the correction, which will likely come in the form of a replacement audio cable for all affected speaker assemblies. For further information or support concerning this issue, please contact Philips Healthcare Customer Care Service Center at 800-722-9377 # 3, # 1 and reference the letter or your local Philips representative in your area.
Z-2584-2011 | Jun-17-2011 | 2011 | Philips Healthcare - Philips SureSigns VS3 Vital Signs Monitor Models:VS3 863069; 863070; 453564041251; 453564041261; VS3 863071; 863072; 863073; 863074; 453564041271; 453564041281; 453564041291; 453564041301The SureSigns VS3 vital signs monitor and VM4; VM6 and VM8 patient monitors are for use by health care professionals whenever there is a need for monitoring the physiological parameters of patients. The SureSigns VS3 vital signs monitor and VM4; VM6 and VM8 patient monitors are for monitoring; recording; and alarming of multiple physiological parameters of adults; pediatrics; and neonates in healthcare environments. Additionally; the monitor is intended for use in transport situations within a healthcare facility. | Philips Healthcare Inc. | 2 | Speakers on the SureSigns Patient Monitors and Vital Signs; failure may be identified by the display of a Speaker Malfunc message in the alarm message area or an AUDIO FAILED message in the main screen of the device or by the absence of audible sound | Philips Healthcare sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated May 2011 to all affected customers. The letter identifies the products, problem, and actions to be taken. The letter states that customers may continue to use the affected products while awaiting replacement of the speaker, provided that the user routinely checks for the display of the "Speaker Malfunct." INOP. Additionally, the letter states that if the customer experience this INOP or there is no sound from the SureSigns Monitor, the customer is to remove the device from use and contact their local Philips service representative. The letter explains that Philips is completing the design for a replacement speaker assembly, which is expected to be available within 8 weeks. When available, Philips will replace the affected speaker assemblies free of charge. Customers are instructed to contact their local Philips Customer Care Center at 1-800-722-9377 if further information or support is required.
Z-0958-2011 | Jan-26-2011 | 2011 | S5 Display and Control Module - S5 Display and Control Module; Part No.: 28-95-10; Sorin Group Deutschland Gmbh; Lindberghstrasse 25; Munchen; Germany.Used to perform; control and monitor extracorporeal blood circulation during an operation for up to six hours. | Sorin Group USA; Inc. | 2 | Touch screen may become unresponsive; inhibiting user input. | All affected US customers and distributors were notified by certified mail on Nov 12, 2010 via a Field Correction letter. They were told that they would be contacted by Sorin to arrange an appointment to inspect and replace affected touch screens on site. They were told that they could safely continue using their S5 systems in accordance with the Operator's Manual and previously supplied instructions until the inspection and necessary replacement had been completed. A customer response form was also included for completion and return to the firm. Customers can contact the firm if they have any questions.
Z-1385-2011 | Feb-22-2011 | 2011 | AUTOMIX 3+3 - AUTOMIX 3+3 Compounder System; product 2M8286 (115V) and 2M8286K (230V); automated nutrition compounders; Baxter Healthcare Corporation; Clintec Nutrition Division; Deerfield; IL 60015The Automix Compounder Systems are electromechanical positive displacement fluid compounding devices used in the admixture of large volume parenteral nutrition formulations via gravimetric weighing. | Baxter Healthcare Corp. | 2 | The AUTOMIX compounder may not respond correctly to the key pressed by the operator using the keypad. The incorrect keypad response issue can occur both in manual data entry mode (STD Mode) and when entering order numbers to download from host software (MAN ID Mode). It is possible that an incorrect volume and/or an incorrect solution could be admixed into the final bag. | Baxter Healthcare Corp. sent Urgent Device Correction letters dated January 27, 2011, via first class mail to all affected customers. informing personnel that use Automix compounders to exercise care when entering data, and verify the accuracy of the keypad response. The accounts were requested to convey this information to all users of the Automix compounders, emphasizing the need to verify data accuracy each time the keypad is used. The accounts were also requested to complete the attached customer reply form listing the model and serial number of each of the Automix compounders at the facility and confirming their receipt of the letter, and fax it to Baxter at the number provided on the form. Any questions regarding the communication were directed to their local Baxter representative, and technical questions were directed to Baxter's Global Technical Services at 1-800-626-2667. After determining the root cause of the incorrect keypad responses, Baxter issued a follow-up correction letter on 4/27/11 outlining the recommended work practices to reduce fluid penetration into the Automix compounder keypad/control module.
Z-0987-2011 | Jan-27-2011 | 2011 | Physiological Monitoring; Patient Monitor - Physiological Monitoring; Patient MonitorM31250 Information Center Local DatabaseReleases A; B; E.01; G; H; J; K; L; & M of Philips IntelliVue Information Center Software used in conjunction with PAS-210 Kit Remote Active Speakers.The intended use of the Information Center Software is to display physiologic waves; parameters and trends; formal data for strip recordings and printed reports and provide the secondary annunciation of alarms from other networked medical devices at a centralized location. The Information Center Software provides for the retrospective review of alarms; physiologic waves and parameters from its database. An additional intended use is to provide primary annunciation of alarms and configuration and control access for networked telemetry monitors. | Philips Healthcare Inc. | 2 | Some PAS-201 active speaker assemblies in use with the Philips IntelliVue Information Center (PIIC) have experienced intermittent audio or loss of audio. Speaker failures may delay recognition of an alarm condition; if operators do not recognize that there is no alarm sound from the monitor and the monitor is not connected to another audible alarm annunciation system. | Philips notified all affected customers with an Urgent Medical Device Correction notification letter dated December 17, 2010. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to observe their system for any of the described symptoms. If they experience intermittent audio or there is no sound from their PAS-210 speaker while it is not connected to another audible alarm annunciation system, customers were instructed to remove the speaker from use and contact their local Philips service representative. Customers were advised not to rely exclusively on the audible alarm system for patient monitoring, the most reliable method of patient monitoring combines close personal surveillance with correct operation of monitoring equipment. ACTIONS PLANNED BY PHILIPS Philips will replace affected speaker assemblies free of charge. A Philips Healthcare representative would contact the customer when the replacement was available. Field Service Engineers will remove and replace defective speakers to correct the problem. The defective speakers will be returned to vendor (Beach Wire) to be discarded. This correction will be done free of charge for customers.
Z-0113-2012 | Oct-27-2011 | 2011 | Precess MRI Patient Monitoring System - Radio used in the Precess MRI Patient Monitoring System. 3160 (Precess) and HE97B; 94044; 94045 (radio and radio kits).Label located on the back of the WPU. Product is labeled in part: "***Invivo Corporation 12601 Research Parkway; Orlando; FL 32826***Made in U.S.A.***REF 3160***SN***CAUTION: Law in U.S.A. Canada and EU restricts this device to sale by or on the order of of a licensed medical practitioner only.***" Label located on the bottom of the DCU. Product is labeled in part: ***Invivo Corporation 12601 Research Parkway Orlando; FL 32826 Made in U.S.A.***REF 3160DCU***SN***CAUTION: Law in U.S.A. Canada and EU restricts this device to sale by or on the order of of a licensed medical practitioner only.***" Display Controller Unit (DCU) is labeled in part: "***Invivo***PRECESS***". Wireless Processing Unit (WPU) is labeled in part: "***Invivo***". NOTE: WPU AND DCU MAY HAVE THE SAME OR DIFFERENT SERIAL NUMBERS.The radio is responsible for wireless communication between the Wireless Processing Unit and the Display Controller Unit. Precess is intended for use by healthcare professionals to monitor vital signs for patients undergoing MRI procedures. | Invivo Corporation | 2 | Philips Invivo has determined that it is possible for the system's alarm limits and other device settings to spontaneously change without user intervention. This problem is associated with intermittent and frequent loss of communication between the wireless Processing Unit and the Display Controller Unit. | A Field Safety Notice entitled Urgent - Medical Device Correction letter, dated September 2011, with attached Field Safety Notice for customers were sent to distributors. The letter described the reason for the recall and the intention of the Field Safety Notice. Also, the hazards involved and instructions on how to identify units that are affected by the recall were provided . The actions to be taken by the user and the actions planned by the Philips Invivo were explained. Users were to follow the instructions provided on how to safely use the device until the unit can be serviced by Philips Invivo. For further information or support concerning this issue, customers are to contact Philips Invivo Customer Support at (877)-468-4861 or their local Philips Invivo representative.
Z-3047-2011 | Aug-22-2011 | 2011 | Infinity Delta Portable Patient Monitor - Infinity Delta Portable Patient Monitor. Draeger Medical Systems; Inc.Portable monitor used to monitor vital signs of neonatal; pediatric; and adult patients for bedside and transport applications. | Draeger Medical Systems; Inc. | 2 | Monitor keys may spontaneously become inoperative or active. This may cause a membrane switch panel malfunction; which could result in the monitor discharging a patient automatically. | Draeger Medical sent an "Urgent Medical Device Recall Notification" letters dated July 2011 to all affected customers. The letter included product and problem affected. The letter informed the customer a Draeger representative will contact them to upgrade the system. Until the solution is implemented, Draeger recommends the Delta and/or Gamma XXL may be used. If the affected monitors become inoperative or activate without user interaction, Draeger advises removing the monitor from service and contacting a Draeger representative. For questions or for information on this recall please call Draeger Service at (800) 543-5047 (press 1 at the prompt, then 2, then 32349).
Z-2181-2011 | May-10-2011 | 2011 | Varian Medical System TrueBeam - Varian Medical System TrueBeam system for stereotactic radiosurgery and radiotherapy. Model number H19; | Varian Medical Systems; Inc. Oncology Systems | 2 | Position sensor failure mode may result in an inaccurate position calculation. The error manifests itself in one of two possible ways.1. The video returned by one of the two cameras inside the Spectra is all white or all black. The consequence of this failure symptom is that the Spectra stops tracking immediately.2. The video returned by one of the two cameras inside the Spectra is returned b | Varian Urgent Medical Device Correction, Urgent Field Safety Notice letters were sent to all affected facilities by certified mail on April 5, 2011.
Z-0082-2012 | Oct-21-2011 | 2011 | ENDOSCOPIC Camera - Smith & Nephew; HD1200 Autoclavable Camera HeadCatalog Number: 72203360Product Usage: Endoscope & Accessories | Smith & Nephew; Inc. Endoscopy Division | 2 | HD1200 camera heads may cause image issues | Smith Nephew sent an "URGENT  PRODUCT RECALL NOTIFICATION" letter dated sent September 23, 2011 to all affected consignees. The letter identified the product, problem and actions to be taken. Consignees were instructed to complete requested information, return affected product and order replacement product. To return product and order replacement product, consignees were instructed to contact the Smith & Nephew Endoscopy Returns Group at 800-343-5717 (option 3) for a Return Authorization (RA) or send an email to endo.andreturns@smith-nephew.com. For questions call 405-917-8556.
Z-1984-2011 | Apr-14-2011 | 2011 | VARIANT(tm) nbs Sickle Cell Prograin Reorder Pack. - VARIANT nbs Sickle Program Resin Update CD-ROM; Model Number:250-3020; included in 250-3000; Distributed and Manufactured by Bio-Rad Laboratories; Inc. The VARIANT nbs Sickle Cell Program is intended as a qualitative screen for the presence of hemoglobins F; A; S; D; C and E in eluates of neonatal blood collected on filter paper by high-performance liquid chromatography(HPLC). The VARIANT nbs Sickle Cell Program is intended for Professional use only. | Bio-Rad Laboratories Inc | 3 | The affected resin Update CD-ROMs cause all user-defined settings in the Setup/Test/Data Setup sub screen and Set/Test/Pattern Setting sub screen to be overwritten when the Update Kit procedure is performed. | Bio-Rad Laboratories sent out the Urgent Medical Device Correction notice on 4/14/2011, by Federal Express. An Update was sent on May 13th, advising customers of a minor discrepancy in the parameters included on the affected Resin Update CD-ROMs. Customers were instructed to examine your inventory for the affected reorder pack lots. Please complete and return the Medical Device Correction Response Form Questions regarding this recall please call (510) 724-6114.
Z-3283-2011 | Sep-28-2011 | 2011 | Boston Scientific; LATITUDE¿ Patient Management System Communicator - Boston Scientific; LATITUDE¿ Patient Management System Communicator; Model 6476; Distributed by Boston Scientific; 4100 Hamline Avenue North; St Paul; MN 55112-5798Boston Scientific; LATITUDE¿ RF Communicator & Software; Model 6482.Distributed by Boston Scientific; 4100 Hamline Avenue North; St Paul; MN 55112-5798LATITUDE Communicator REF 6468 (only distributed outside the US)The LATITUDE Patient Management system is intended for use to remotely communicate with a compatible Boston Scientific pulse generator and transfer data to a central database. | Boston Scientific CRM Corp | 3 | A limited number of Model 6476; 6482; and 6468 (Not available in the US) LATITUDE Communicators are unable to authenticate; and thus cannot connect to the LATITUDE System. Customers are instead receiving Latitude Temporarily Unavailable (LTU) message. | Boston Scientific CRM Corporation sent a " Dear Patient", letter dated August 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. A new communicator is being shipped separately to you with a corrected security certificate that will allow it to be activated and function normally. 1. Please activate this new equipment at your convenience to begin home monitoring. Activation instructions are enclosed with the communicator. 2. We also ask that you return the original communicator previously delivered to you: " Place the original equipment in the shipping box used for the new equipment. " Pack newspaper or other filler around the equipment to prevent damage during shipping, and seal the box. " Enter your address in the FROM section of the prepaid Federal Express label, and apply the return label to the outside of the box directly over the original label. " Call Federal Express at 1.800.463.3339 to schedule a pick-up. We apologize for this inconvenience. If you have further questions, please call 1.866.484.3268.
Z-0364-2012 | Dec-02-2011 | 2011 | TrueBeam and TrueBeam STx; - TrueBeam and TrueBeam STx; Model Number: H19; Ref/FSCA identifier: CP-06381 are intended to provide stereotactic radiosurgery and precision radiotherapy for lesions; tumors; and conditions anywhere in the body where radiation treatment is indicated. | Varian Medical Systems; Inc. Oncology Systems | 2 | An anomaly has been identified with the TrueBeam and TrueBeam STx systems where; under certain tuning conditions; the electron beam emerging from the bend magnet may have an elongated spot shape. | On 11/04/2011, an Urgent Medical Device Correction Letter was distributed to all affected users, with a description of the problem and user corrective action steps. Varian is developing a corrective action for this issue. A customer service representative will contact you when the correction is available to schedule installation on your system. Questions regarding the letter should be directed to the Varian Oncology Help Desk Contact Information at USA and Canada: 1-888-827- 4265 or Europe +41-41-749-8844.
Z-1374-2011 | Feb-18-2011 | 2011 | AMO brand WaveScan WaveFront System - AMO brand WaveScan WaveFront System; OPM software versions 3.67 and 3.67;Part Number: 0070-1531; Catalog Numbers: 0070-1664;0070-1649;0070-1654: Version 3.670070-1721: Version 3.671 0070-1662: Version 3.67; US version;0070-1668; IVersion 3.67 International versionProduct is manufactured and distributed by AMO Manufacturing USA; LLC; Milpitas; CAThe WaveScan WaveFront System is a class III accessory device used in conjunction with the VISX¿ STAR S4 and STAR S4 IRTM Excimer Laser Systems; which are indicated for wavefront-guided laser assisted in situ keratomileusis (LASIK) eye surgery in certain patient populations. The WaveScan WaveFront System is a diagnostic instrument intended for the automated measurement; analysis; and recording of refractive errors of the eye. The WaveScan System software uses the diagnostic data of the eye's refractive errors and wavefront aberrations to calculate the target treatment shape and then generates the commands for the VISX STAR laser to create the target shape on the cornea. | AMO Manufacturing USA; LLC | 2 | The failure to comply with Minimum System Requirements; increased system errors could result; and a miscreated treatment table (incorrect LASIK refractive treatment) could result. | Abbott Medical Optics, Inc. sent an Urgent Medical Device Field Correction letter to all affected customers starting July 15, 2010. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to take the following actions: Surgeons were instructed to complete the facsimile attachment to acknowledge receipt of the notification letter and fax to 714-247-4530. For those using the affected product a 3.68 OPM disc must be requested by e-mailing a request to OPM@pacful.com. For any questions please contact AMO Technical Service Representatiave at 1-877-266-4543.
Z-2052-2011 | Apr-26-2011 | 2011 | LATITUDE Patient Management System - Boston Scientific's LATITUDE Patient Management System:LATITUDE 6.1 Patient Management System WebsiteServer software; Model 6488 for United StatesServer software; Model 6465 for European UnionLATITUDE 6.2 Patient Management System WebsiteServer software; Model 6441 for CanadaServer software; Model 6442 for Australia / New ZealandIntended for use to remotely communicate with a compatible Boston Scientific pulse generator and transfer data to a central database. | Boston Scientific CRM Corp | 2 | Boston Scientific has determined that data within certain printed reports obtained from the LATITUDE Patient Management System may not be accurate; in that; the Most Recent Presenting Electrogram (MRP-EGM) for some of the selected patients may be either missing or incorrect. | Consignees were sent a Boston Scientific "LATITUDEÂ® Medical Records Information" letter dated January 2011. The letter was addressed to "Dear Doctor". The letter described the problem and the product involved in the recall. The letter also included the Rate of Occurrence, Action and Further Information sections.
Z-0185-2012 | Nov-16-2011 | 2011 | LIFEPAK 20 defibrillator - The LIFEPAK 20 defibrillator/monitor is an AC powered external defibrillator with battery backup used by healthcare providers in hospital and clinic settings.Affected Product Part Numbers: 3202487-000 thru -012; 3202487-014 thru -027; 3203487-029 thru -039; 3202488-000 thru -019; U3202487-000; U3202487-015; U3202488-000. | Physio Control; Inc. | 2 | When users do not disconnect the test plug and reconnect the QUIK-COMBO electrodes to the therapy cable; the test plug signal can be misinterpreted as a patient waveform. | Physio Control sent a " Urgent Medial Device Safety Alert" letter dated January 2009, to all affected customers. The letter identified the product, the problem and the action needed to be taken by the customer. The letter instructed the customers to follow the Operating Instructions for AED and Manual Defibrillation mode (Section 4-Therapy, pages 58 and 61) that instructs users to "Connect therapy electrodes to the therapy cable and confirm cable connection to the defibrillator" to prevent users from mistakenly leaving the test plug and reconnect the QUIK-COMBO electrodes to the therapy cable, the test plug signal can be misinterpreted as a patient waveform. Recommendations to users who use a Therapy Cable: -Disconnect the QUIK-COMBO test plug prior to patient use -Confirm therapy electrodes - to - therapy cable connection -Confirm therapy cable - to - defibrillator connection -Confirm therapy electrodes - to - patient connection Adding 2005 American Heart Association software that includes audible and visual prompts to LIFEPAK 20 defibrillators manufactured prior to January 2007 is optional and not required. If the customers would like to have this optional kit installed in your defibrillators, please contact Physio Control, Inc. Customers can call Technical Support at 1-800-442-1142 for any questions about this recall or or visit our website at www.physio-control notices.com/LP20testplug.
Z-1020-2011 | Feb-02-2011 | 2011 | i-STAT Level 2 Control - i-STAT Level 2 Control Value Assignment SheetsHuman Consumption | Abbott Point of Care Inc. | 2 | Three i-STAT Level 2 Control Value Assignment Sheets and the i-STAT CLEW A20 Software System Manual Update CD contain the incorrect range for pCO2; BUN and Glucose for the i-STAT EC8+; 6+; EC4+; E3 and G cartridge types. | Abbott Point of Care Inc., sent an URGENT RECALL NOTICE to all affected customers on 8/3/2010. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to replace all copies of the value sheets with the ones provided in the letter. ASLO, if they have a disc with the Value sheets on it, it is to be replaced with the updated disc and the original discarded. If the control materials, discs or value sheets have been forwarded to another facility, a copy of the letter was to be sent to them. The Business Reply card is to be completed and returned to Abbott. Questions or comments are to be directed to Abbott Point of Care Technical Support at 800-366-8020, Option 1.
Z-3075-2011 | Aug-25-2011 | 2011 | GE Healthcare; DASH 3000; 4000; 5000 Patient Monitors - GE Healthcare; DASH 3000; 4000; 5000 Patient MonitorsThe Dash 3000/4000/5000 patient monitor is intended for use under the direct supervision of a licensed healthcare practitioner. The intended use of the system is to monitor physiologic parameter data on adult; pediatric and neonatal patients. The Dash 3000/4000/5000 patient monitor is designed as a bedside; portable; and intra-hospital transport monitor that can operate in all professional medical facilities including but not limited to: emergency department; operating room; post anesthesia recovery; critical care; surgical intensive care; respiratory intensive care; coronary care; medical intensive care; pediatric intensive care; or neonatal intensive care areas located in hospitals; outpatient clinics; free-standing surgical centers; and other alternate care facilities. Physiologic data includes but is not restricted to: electrocardiogram; invasive blood pressure; noninvasive blood pressure; heart rate; temperature; cardiac output; respiration; pulse oximetry; carbon dioxide; bi-spectral index; impedance cardiography; oxygen; and anesthetic agents as summarized in the operator's manual. The Dash 3000/4000/5000 patient monitor is also intended to provide physiologic data over the UNITY NETWORK TM indirectly to clinical information systems (via our Enterprise Gateway) and allow the user to access hospital data at the point-of-care. The information can be displayed; trended; stored; and printed. The Dash 3000/4000/5000 patient monitor was developed to interface with nonproprietary third party peripheral devices that support serial data outputs. | GE Healthcare; LLC | 2 | GE Healthcare has recently become aware of an issue with the Dash 3000/4000/5000 patient monitors configured with the 802.11b wireless LAN option that may impact patient safety. The Dash monitor may exhibit reboots on some networks when the 802.11b wireless LAN option is used to connect to the hospital wireless LAN infrastructure. This reboot results in a loss of patient monitoring for approxima | GE Healthcare sent an URGENT MEDICAL DEVICE CORRECTION letter dated June 15, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. GE recommended actions that should be taken by a qualified service representative. For any questions regarding the medical device correction or identification of affected items contact Technical Support at 1-800-558-7044 or customers should contact their local Service Representative. UPDATE: A second letter dated 12/13/11 was sent on 1/20/12 to consignees. The revised letter provides an update to the safety instructions and product correction.
Z-2573-2011 | Jun-20-2011 | 2011 | Ultraview SL Capnography Module - The Ultraview SL 91517 Capnography module is a mainstream and sidestream gas analyzer designed to measure the concentration of carbon dioxide in a gas mixture; to aid in determining the patient's ventilatory; circulatory; and metabolic status.Mainstream monitoring uses a self-calibrating CO2 sensor to ensure continuous monitoring without interruption of connection. Reusable; lightweight (<14 grams); or single-use airway adapters are available for both adult and neonatal (low dead space) patients. Respiration can be monitored via an airway adapter and endotracheal or tracheostomy tubes using the mainstream function of the module.The intended use statement that apply to the 91517 are continuous monitoring of carbon dioxide and oxygen. The Model 95000-A Bedside/Transport Monitor is used at the patient's bedside to actively acquire; analyze; monitor and display a variety of clinical parameters (in clinically relevant formats) for adults; pediatric or neonatal patients. The Model 95000-A Bedside/Transport Monitor is designed for use by a healthcare professional in a hospital; healthcare facility (i.e.; clinic); or intra-hospital transport. This product is not designed for home use or for use in transport vehicles. | Spacelabs Healthcare; Llc | 2 | Capnography Module; Model 91517; resets 20 seconds after power failure; instead of after three minutes as specified. This reset clears the RAM and the module returns to default settings. | Spacelabs Medical sent a "URGENT-MEDICAL DEVICE CORRECTION" letter dated May 12, 2011 to all affected customers. The letter identifies the product, problem, and actions to be taken. Customers were recommended to set all their Capnography Module Configuration Manager alarm settings to "ON". The letter also recommended that customers brief their staff about the possibility of losing changes to the default settings and to restore them after a power interruption. A field service engineer will be contacting customers to set up a date and time when affected modules can be upgraded at no cost to the customer. For questions about this corrective action program, call Spacelabs Medical at 1-800-522-7025, Select 2 for Technical Support. For outside the US, customers can call 1-425-657-7200, x5089.
Z-2073-2011 | Apr-26-2011 | 2011 | Unicel¿ DXC and SYNCHRON¿ LX Clinical Systems - UniCel¿ DxC860iThe UniCel DxC Integrated Workstation combines a UniCel DxC 600 or UniCel DxC 800 analyzer and a UniCel DxI 600 or UniCel DxI 800 analyzer into a single instrument presentation. Samples are loaded from a single point of entry through a UCTA (UniCel Closed Tube Aliquotter). The UCTA functions as the sample processing manager by aliquotting and routing samples to the DxC and the DxI analyzers according to programming requirements. The UniCel DxC 600 or DxC 800 Synchron Clinical System is a fully automated; computer-controlled clinical chemistry analyzer intended for the in vitro determination of a variety of general chemistries; therapeutic drugs; and other chemistries of clinical interest in biological fluids such as serum; plasma; urine; or cerebrospinal fluid (sample type is chemistry dependent). The UniCel DxI 600 or DxI 800 Access Immunoassay System is an in vitro diagnostic device used for the quantitative; semi-quantitative; and qualitative determination of various analyte concentrations found in human bodily fluids. | Beckman Coulter Inc. | 2 | This recall was initiated because Beckman Coulter has confirmed that for User-Defined Chemistries (UDRs); the polychrome correction is not automatically updated when wavelengths are modified. | The firm, Beckman Coulter, sent an "URGENT: PRODUCT CORRECTIVE ACTION (PCA)" letter dated October 20, 2010 with attached PCA Response Form (via US Postal Service for US Customers) to all customers who purchased Synchron LX and UniCelÂ®Â® DxC Clinical Systems Operating Software (all versions). The letter describes the product, problem and actions to be taken. The customers were instructed to do the following: -After modifying any UDR wavelength, they MUST reboot to apply the update of the polychrome correction. -If the modified UDR requires calibration, recalibrate after rebooting. - Complete and return the enclosed RESPONSE FORM within 10 days via fax to (714) 961-4234 or mail to: Beckman Coulter, Inc., Attn: Regulatory Affairs, 250 S. Kraemer Blvd, PO Box 8000, Brea, California, 92282-8000; Mail Code E2.SE.08. In addition, the customers were instructed to share this information with their laboratory staff and retain this notification as part of their laboratory Quality System documentation. If customers forwarded any of the affected product listed above to another laboratory, please provide a copy of this letter to them. If customers have any questions regarding this Product Corrective Action letter, please call Beckman's Customer Support Center (Hotline) at (800) 854-3633 in the United States and Canada or www.beckmancoulter.com/customersupport/supportrequests.
Z-2165-2011 | May-10-2011 | 2011 | Instrumentation Laboratory GEM Premier 4000 - Instrumentation Laboratory GEM Premier 4000 PN 00025000000The GEM Premier 4000 is a portable critical care system for use by health care professionals to rapidly analyze whole blood samples at the point of health care delivery in a clinical setting and in a central laboratory. The instrument provides quantitative measurements of pH; pC02; p02; sodium; potassium; chloride; ionized calcium; glucose; lactate; hematocrit; total bilirubin and CO-OXimetry (tHb; 02Hb; COHb; MetHb; HHb) parameters. Total bilirubin can also be quantitated from heparinized plasma samples when analyzed in the tBili/CO-Ox mode. These parameters; along with derived parameters; aid in the diagnosis of a patient's acidlbase status; electrolyte and metabolite balance and oxygen delivery capacity. Total bilirubin measurements are used in the diagnosis and management of biliary tract obstructions; liver disease and various hemolytic diseases and disorders involving the metabolism of bilirubin. In neonates; the level of total bilirubin is used to aid in assessing the risk of kernicterus | Instrumentation Laboratory Co. | 2 | Measurement (amperometric) spike can occur early in cartridge life on the glucose and lactate sensors during patient blood analysis; leading to erroneously highresults | Instrumentation Laboratory notified accounts by a Field Safety NotifIcation Letter dated Feb 2011. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to use whichever of the following mitigation options works best for their institution. Both options will have the same effect to eliminate the potential impact of a rare spike on the glucose and lactate sensors. Customers were instructed to complete the enclosed Response Tracking Form and return immediately fax to 781-861-4207 or e-mail to ra-usa@ilww.com. Customers were advised to contact their local representative with any questions.
Z-1730-2011 | Mar-18-2011 | 2011 | Coulter HmX Hematology Analyzer and HmX Hematology Analyzer with Autoloader - Coulter HmX Hematology Analyzer and HmX Hematology Analyzer with Autoloader.Hematology Analyzers; Part Number: 6605522; 6605523; 6605524; 6605525; 6605526; 6605527; A85566; and A85564.Quantitative; automated hematology analyzer and leukocyte differential cellcounter for In Vitro Diagnostic Use in clinical laboratories. | Beckman Coulter Inc. | 2 | The recall was initiated because Beckman Coulter has confirmed the Coulter MAXM; MAXM AL; HmX and HmX AL Analyzers omit the tilde (~) character when thetilde is used as part of the Sample 10 within a barcode label scanned by the primary mode barcode reader. There is a potential for Specimen or Patient misidentification to occur. | The recall communication was initiated on 12/14/2010 with Beckman Coulter forwarding an Urgent: Product Corrective Action (PCA) letter with attached PCA Response Form (via US Postal Service for US Customers) to all customers who purchased the COULTER HmX Hematology Analyzer, COULTER HmX Hematology Analyzer with Autoloader, COULTER MAXM Analyzer, COULTER MAXM Hematology Analyzer with Autoloader. The letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were instructed not to use the tilde (~) character in Specimen or Patient identifiers. In addition, customers were instructed to complete and return the enclosed PCA Response Form within 10 days. Customers with any technical questions regarding this Product Corrective Action were instructed to call Beckman Coulter Customer Service at 800-526-7694 in the United States or Canada, or contact their local Beckman Coulter representative.
Z-0937-2011 | Jan-24-2011 | 2011 | Welch Allyn PIC 50 Portable Intensive Care System - Welch Allyn PIC 50 Portable Intensive Care System. Model 971081 series.Designed for use in and out of the hospital. | Welch Allyn Protocol; Inc | 2 | Potential likelihood that the Portable Intensive Care (PIC) 50 Defibrillator's screen display may fail. The loss of the ability to view the patient's Electrocardiogram (ECG) on the screen could cause the user to delay or fail to administer defibrillating shocks needed to resuscitate a patient. | On 9/17/10, Welch Allyn began sending out the Important Product Alert letter to all their direct consignees. On 9/24/2010, the firm sent out the letters to their distributors. The Product Alert letter informs the consignees of a low potential likelihood that the Portable Intensive Care (PIC) 50 Defibrillator's screen display may fail. The consignees are advised to follow the instruction in the PIC 50 user's manual requiring routine shift equipment checks daily or at each shift change to minimize the likelihood of the malfunction from occurring. The checks ensure that the PIC 50 and its component (including the screen display) are fully operation and undamaged. The letter further states "our analysis indicates that routinely doing such checks will likely lead to discovery of this malfunction as it occurs, and thus timely maintenance and repair". The distributors are advised to communicate this Product Alert to their customers if the devices were further distributed. For customers in US and Canada, distributors are advised to provide the electronic Excel file with customer's name, address, serial number and part number to WA2221@Stericycle.com. For customers in other countries, they are advised to provide information to PIC50Product Alert@welchallyn.com Consignees can call Welch Allyn at 800-289-2500 for any questions during normal business hours.
Z-2185-2011 | May-12-2011 | 2011 | CDI 500 Blood Parameter Monitoring System - CDI Blood Parameter Monitoring System 500 with Arterial and Venous blood parameter modules and Hematocrit/Saturation probe; Catalog 500AVHCT; Terumo Cardiovascular Systems Corporations; Ann Arbor; MI. The CDI 500 is indicated for use to display patient blood gas values during extracorporeal procedures. | Terumo Cardiovascular Systems Corporation | 2 | The firm has received 234 complaints of inaccuracy of displayed values on the CDI 101 and 500 monitors. In some cases the customer reported inaccurate potassium values. Other inaccuracies include CO2; hematocrit; O2 saturation; pH; hemoglobin and other miscellaneous inaccuracies. The CDI 101 and CDI 500 system are stand alone systems that are used during extracorporeal procedures to continuou | The firm, TERUMO, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated October 27, 2010 to its customers. The letter described the product, problem and actionst to be taken. The customers were instructed to reveiw this notice; assure that customers are ware of this notice; confirm receipt of this notification by faxing the attached Customer Response Form, and place the attached stickers in a visible place in the Operators' Manuals, such as the inside of the front cover, if additional stickers are needed, the customers should contact Terumo CVS Customer Service at 1-800-521-2818. The letter states that when a displayed value is significantly different from expectations based on the clinical situatioin, the user should verify the accuracy by independent means before treating. The letter included "Caution" stickers that were to be applied to the CDI system operatros manual. The letter also states "Please note that we are not recommending that you stop using your CDI monitor system." If you have any questions or concerns, please contact Terumo CVS Customer Service at 1-800-521-2818. Customer Service hours are from Monday-Friday, 8AM - 6PM EST.
Z-2188-2011 | May-12-2011 | 2011 | CDI 500 Blood Parameter Monitoring System - CDI Blood Parameter Monitoring System 500 with Arterial Blood parameter module Catalog number 500A ; Terumo Cardiovascular Systems Corporations; Ann Arbor; MI. The CDI 500 is indicated for use to display patient blood gas values during extracorporeal procedures. The CDI 500 system is a stand alone systems that is used during extracorporeal procedures to continuously monitor the blood in the extracorporeal circuit; providing ongoing information about the blood parameters. | Terumo Cardiovascular Systems Corporation | 2 | The CDI 500 monitors may display inaccurate values. These errors include inaccurate potassium; CO2; hematocrit; O2 saturation; pH; hemoglobin and other miscellaneous inaccurate values. | The firm, TERUMO, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated October 27, 2010 to its customers. The letter described the product, problem and actionst to be taken. The customers were instructed to reveiw this notice; assure that customers are ware of this notice; confirm receipt of this notification by faxing the attached Customer Response Form, and place the attached stickers in a visible place in the Operators' Manuals, such as the inside of the front cover, if additional stickers are needed, the customers should contact Terumo CVS Customer Service at 1-800-521-2818. The letter states that when a displayed value is significantly different from expectations based on the clinical situatioin, the user should verify the accuracy by independent means before treating. The letter included "Caution" stickers that were to be applied to the CDI system operatros manual. The letter also states "Please note that we are not recommending that you stop using your CDI monitor system." If you have any questions or concerns, please contact Terumo CVS Customer Service at 1-800-521-2818. Customer Service hours are from Monday-Friday, 8AM - 6PM EST.
Z-1522-2011 | Mar-03-2011 | 2011 | ApexPro Telemetry System - GE Healthcare; ApexPro Telemetry System The ApexPro Telemetry System is intended for use under the direct supervision of a licensed healthcare practitioner. The system is designed to acquire and monitor physiological data for ambulating patients within a defined coverage area. The system processes this physiological data to detect various ECG arrhythmia events and select physiological parameter limit violations. The ApexPro Telemetry System is intended to be installed in the hospital or clinical environment in order to provide clinicians with patient physiological data; while allowing for patient mobility. These systems are typically deployed in sub acute care areas in hospitals or clinical sites where patient mobility can enhance the effectiveness of the medical procedures administered. The physiological parameters monitored include ECG; noninvasive blood pressure; non-invasive temperature and SpO2. The ApexPro Telemetry System is intended to provide ECG data via Ethernet to the computer platform for processing. The ApexPro Telemetry System is also intended to provide physiologic data over the Unity Network to clinical information systems for display. | GE Healthcare; LLC | 2 | GE Healthcare has become aware of a potential for loss of patient monitoring when ApexPro Telemetry System is used in conjunction with a mirrored view CARESCAPE CIC Pro in combination monitoring mode (combo mode). The safety issue includes the potential for a delay in annunciation of audible alarms when used in combo mode; and may result in thedelay of treatment of a cardiac event. | GE sent an "Urgent Medical Device Correction" letter dated December 29, 2010. The letter was addressed to Healthcare Administrator / Risk Manger, Chief of Nursing and Director of Biomedical Engineering. The letter described the Safety Issue, Affected Product Details, Product Correction, Safety Instruction and Contact Information. Customers were instructed to contact Technical Support or their local Service Representative with any questions regarding this medical device. For any questions regarding this recall call 262-573-4122.
Z-2916-2011 | Aug-02-2011 | 2011 | Vital Sign elance Monitor; Model 93300 - Vital Sign elance Monitor; Model 93300Spacelabs Healthcare; Llc.A vital signs patient monitor with arrhythmia detection. | Spacelabs Healthcare; Llc | 2 | Model 93300 Vital Signs Monitor may stop operating while in use; and when this occurs; the monitor may also fail to power ON. The resulting situation may cause a loss of monitoring for the patient. | Spacelabs Healthcare, Llc sent an "Urgent-Medical Device Correction" letter dated 18 July 2011 to all affected customers. The letter notified the customers of the problem and recommended customers make their staff aware. It also informed them a Field Service Engineer will contact them to check, and if necessary make adjustments. Customers can contact Spacelabs Healthcare at 800-522-7025 and select 2 for Technical Support or outside the U.S. at +86512-8717-1888, extension 304.
Z-0077-2012 | Oct-20-2011 | 2011 | Medtronic Pressure Display Box - Medtronic Pressure Display Box; Models 66000; 64000; 63000; 60000. It is capable of measuring pressure between -100 mmHg and +500 mmHg with an accuracy of +/-5 mmHg from -100 mmHg to +300 mmHg and +/-10 mmHg from +301 to +500 mmHg. The Pressure Display Box has an upper and lower alarm setting. The alarm will be activated when the pressure is equal to or beyond the set value. Manufacturer: Medtronic; Inc. 710 Medtronic Parkway; Minneapolis; MN 55431-5604.Product Usage: This product is intended for use in displaying line pressures during cardiopulmonary bypass surgery. | Medtronic Perfusion Systems | 2 | The DFU for the Medtronic Pressure Display Box; model 66000; states that; for display readings of -100 to +300 mmHg there will be an accuracy of +/-5 mmHg; and for display readings of +301 to +500 mmHg there will be an accuracy of +/-10 mmHg. Through reports from customers; Medtronic has become aware that the Pressure Display Box may not always accurately display readings. During use in the -40 | Medtronic Perfusion Systems, sent an "Urgent Medical Device Notification letter dated September 9, 2011, via UPS 2-day delivery to all affected customers. The letter identified the affected devices, described the issues, and the actions needed to be taken. The letter indicates when approved, the updated Directions for Use for the Pressure Display Box and Pressure Display Sets will be sent to customers. Customers were asked to complete the Customer Confirmation Certificate and return to Medtronic per the instructions on the form. The letter states that if the recalled products have been further distributed, forward this letter immediately to all customers who received the affected products, and gather proof of notification. For further questions concerning this notification, contact Medtronic Cardiovascular Technical Support at 877-526-7890 or Medtronic Sales Representative.
Z-1473-2011 | Mar-02-2011 | 2011 | AdvantageSim MD - GE ADVANTAGE SIM MD versions 5.x; 6.x; 7.x up to and including version 7.7.0Advantage Sim MD is used to prepare geometric and anatomical data relating to a proposed external beam radiotherapy treatment prior to dosimetry planning. Anatomical volumes can be defined automatically or manually in three dimensions using a set of CT images acquired with the patient in the proposed treatment position. Definition of the anatomical volumes may be assisted by additional CT; MR; PET or SPECT studies that have been co-registered with the planning CT scan. Additionally; CT & PET data from a respiratory tracked examination may be used to allow the user define the target or treatment volume over a defined range of the respiratory cycle. The geometric parameters of a proposed treatment field are selected to allow non dosimetric; interactive optimization of field coverage. Defined anatomical structures and geometric treatments fields are displayed on transverse images; on reformatted sagittal; coronal or oblique images; on 3 D views created from the images; or on a beam eye's view display with or without the display of defined structures with or without the display of digitally reconstructed radiograph | GE Healthcare; LLC | 2 | GE Healthcare has become aware that inconsistent DRR (Digitally Reconstructed Radiographic) visualization may occur when the blended DRR feature or the Isocenter Move mode feature is used in combination with specified other functions of the AdvantageSim application on Advantage Workstation. These inconsistencies may affect patient safety. Other AdvantageSim functions are not affected. | GE Healthcare sent an "Urgent Medical Device Correction" letter dated November 23, 2010. The letter was addressed to Hospital Administrators / Risk mangers, Oncology Department Managers / Radiation Oncologists, & Radiation Physicists / Dosimetrists / Therapists. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information. For any questions or concerns regarding this recall, please call one of the following phone numbers: United States: 800 437 1171 Japan: 0120 - 055 - 919 For other countries, please contact your local GE Healthcare Service Representative.
Z-2593-2011 | Jun-20-2011 | 2011 | XiO Radiation Treatment Planning System - XiO Radiation Treatment Planning System; XiO Release 2.1.0 and aboveUsed to create treatment plans for any cancer patient for who external beam radiation therapy or brachytherapy has been prescribed. | Computerized Medical Systems Inc | 2 | XiO Software: The Patient Scan position is being confused with the treatment orientation of the patient. The Patient Scan position is determined from the data entered in PFM. Subsequently; when a New Teletherapy Plan is created for a patient in Teletherapy; the user is prompted with Head/Foot Gantry on the text page. This prompt is asking for the orientation of the patient when they are treated. | Computerized Medical Systems Inc sent an undated Safety Notice to all affected customers. The notice indicated the problem, the clinical impact, and the workaround. The notice stated the sites will be notified when a solution is available.
Z-2408-2011 | Jun-03-2011 | 2011 | XiO Radiation Treatment Planning System - XiO Radiation Treatment Planning System; XiO Release 1.0.2 and above.Used to create treatment plans for any cancer patient for who external beam radiation therapy or brachytherapy has been prescribed. | Computerized Medical Systems Inc | 2 | When RT Plans are exported from XiO and RT Structure Sets and Images are exported from Monaco; the DICOM iso-center coordinates may not be accurate. | Elekta / Computerized Medical Systems Inc began mailing an Important Safety Notice to all affected customers in September 2010. The notice identified the product, the problem, and the action to be taken by the customer. Customers were advised not to mix treatment plan information from multiple planning systems. Customers were asked to return the postcard as proof of receipt and their understanding of the Safety Notice. For questions regarding this recall call 408-830-8000.
Z-1901-2011 | Apr-06-2011 | 2011 | Prismaflex - Prismaflex System; Catalog Number 107493; Indicated for continuous solute and/or fluid removal in patients with acute renal failure or fluid overload. | Gambro Renal Products; Incorporated | 2 | Prismaflex touch screen may become unresponsive during Continuous Renal Replacement Therapy procedure. | The firm, GAMBRO Renal Products, sent an "URGENT - MEDICAL DEVICE RECALL" letter dated March 14, 2011 to all customers. The letter described the product, problem and actions to be taken. The customers were provided with response procedures to use if the frozen screen problem occurs while using the product. The customers were instructed to closely monitor the return line for potential air and/or clots in the circuit; review the instructions provided in the attached copy of Chapter 5: Troubleshooting, "Softkeys won't work." of the Prismaflex Operators Manual in conjunction with the attached Troubleshooting Checklist; be aware if performing Manual Termination with Blood Return, air detection is not provided, and complete and return the Customer Reply Form via fax to: Regulatory Affairs at 1-866-447-4471 or 303-222-6916, or by email to: GRPRegulatory.Affairs@us.gambro.com. If you have any questions or concerns, please contact for Clinical questions: ICON 800-554-1312 (available 24/7) or your local Gambro Representative; for Technical questions: Gambro Customer Support 800-525-2623, option 2 (Available 6:00am to 5:30pm MT, Monday thru Friday).
Z-1526-2011 | Mar-04-2011 | 2011 | Fresenius 2008T Hemodialysis Machine - Fresenius 2008T Hemodialysis MachinePart Number 190573; 2008T Hemodialysis System; OLCIDIASAFE PLUSPart Number 190713; 2008T Hemodialysis System with CDXPart Number 190766; 2008T Hemodialysis System with biBagThe Fresenius 2008T is indicated for acute and chronic dialysis therapy. | Fresenius Medical Care Holdings; Inc. | 2 | Cursor on the dialysis screen may disappear and not respond to user input | Fresenius Medical Care contacted customers by mailing a Customer Notification letter "Urgent correction FMNCA Model 2008T Hemodialysis Machine" via Certified-Mail, Return Receipt requested on January 19, 2011. The Technical Services department began contacting consignees by telephone on January 13, 2011, to schedule an appointment for installing the electronic board on the 2008T machines at the customers' facilities. For questions call Technical services 1-800-227-2572.
Z-1944-2011 | Apr-11-2011 | 2011 | CDI 100 - Terumo Cardiovascular Systems Corporations; CDI 100 Hematocrit/ Oxygen Saturation Monitoring System. In-line blood-gas sensor used in cardiopulmonary bypass. | Terumo Cardiovascular Systems Corporation | 2 | Inaccurate result. Possibility for displayed values for potassium; CO2; hematocrit; O2 saturation; pH; hemoglobin and other measured blood parameters to be inaccurate. | On 11/1/2010, Terumo sent "URGENT MEDICAL DEVICE CORRECTION" notices to Chief Perfusionist, Department of Cardiovascular Surgery or Director of Operating Room Services. The letter informed the consignees about the incidents of the CDI 100 and 101 displaying inaccurate values. The letter states that when a displayed value is significantly different from expectations based on the clinical situation, the user should verify the accuracy by independent means before treating. The letter included "Caution" stickers that were to be applied to the CDI system operators manual.
Z-1733-2011 | Mar-18-2011 | 2011 | CELL-DYN Emerald Hematology Analyzer System - Abbott brand CELL-DYN Emerald Hematology Analyzer System; List Numbers: 09H39-01; The product is manufactured in France for Abbott Diagnostics Division of Abbott Laboratories; Abbott Park; IL The CELL-DYN Emerald is an automated hematology analyzer designed for in-vitro diagnostic use in clinical laboratories. | Abbott Laboratories | 2 | Abbott has determined that gain settings for the non-standard specimen types may not match the gain settings for the standard specimen type. There is the potential for impact to product performance. | The firm, Abbott, sent two "Product Correction" letters one dated March 23,2010 and a follow-up letter dated July 1, 2010, via express mail to all its consignees/customers. The letters described the product, problem and actions to be taken. The customers were instructed to perform the following steps, noted in the letter, to print and verify their gain setting. Abbott informed the customers that they would receive a mandatory software update and were instructed to install the CELL-DYN Emerald according to the attached instructions, and to complete and return the Customer Reply form via fax to 1-800-777-0051 or email to QAGCO@abbott.com. Should you have any questions or would like to discuss in more detail, contact Director Quality Assurance and Regulatory Affairs Abbott Hematology at 408-567-3403.
Z-2965-2011 | Aug-09-2011 | 2011 | STAR S4 IR Excimer Laser System; Software Version 5.00 - STAR S4 IR Excimer Laser System; Software Version 5.00Mfg by : AMO Manufacturing USA; LLC510 Cottonwood DriveMilpitas; CA 95035Iris registration is designed to compensate for pupil centroid shift and cycloration; both of Which may occur because of patient characteristics and the different environments at the time of measurement and at the time of treatment. | AMO Manufacturing USA; LLC | 2 | Reports of lower success rates with Iris Registration (IR) image capture with the STAR S4 IR Excimer Laser System v 5.3 that with previous version. | Abbott Medical Optics (AMO) sent an "ADVISORY NOTICE" dated July 18, 2011 to all customers using the AMO WAVESCAN Wavefront System. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to create a separate patient file for each eye on the WAVESCAN System (STAR S4 IR Excimer Laser Systems,software versions 5.00-5.30) according to the directions in attachment A. Additionally, customers are to complete and return an attached form to AMO via fax at the regional fax number. Contact an AMO representative for questions regarding this notice.
Z-2610-2011 | Jun-22-2011 | 2011 | CDITM 101 Monitor - CDI TM 101 Monitor; CDI TM 101 Hematocrit/Oxygen Saturation Monitoring System; The CDI TM 101 Hematocrit/Oxygen Saturation Monitoring System is intended for use during cardiopulmonary bypass procedures where you need continuous monitoring of the blood's oxygen saturation; hematocrit and hemoglobin levels. | Terumo Cardiovascular Systems Corporation | 2 | Terumo discovered through its own testing that the presence of electromagnetic interference (EMI) in the operating room may affect the accuracy of the valuesdisplayed by the CDI"101 Hematocrit/Oxygen Saturation Monitoring System: hematocrit (HCT); hemoglobin (Hgb) and Oxygen Saturation. | Terumo Cardiovascular Systems, Corp issued an "URGENT MEDICAL DEVICE CORRECTION" letter dated April 28, 2011, to all affected customers. The letter identifed the product, the problem and the action needed to be taken by the customer. After a solution has been verified, Terumo CVS will implement a correction for all affected units. Customers were instructed to review this Medical Device Correction notice, Assure that all users are aware of this notice. Confirm receipt of this and to fax back the Customer Response form. For further question please contact Terumo CVS Customer Service at 1-800-521-2818, Monday - Friday, 8 AM - 6 PM EST.
Z-3012-2011 | Aug-12-2011 | 2011 | GE Healthcare; Seno Advantage 2.1 and 2.2 Workstations - GE Healthcare; Seno Advantage 2.1 and 2.2 Workstations with RIS CCOW integration; GE Healthcare 3000 North Grandview; Waukesha; WI 53188Seno Advantage is a medical image review station that allows easy selection; processing; filming and media interchange of multi-modality images from a variety of diagnosis imaging systems. When interpreted by a trained physician; mammographic images displayed on the high-resolution monitors may be used as an element for diagnosis. Furthermore filmed images from all modalities may also be used as an element for diagnosis. | GE Healthcare; LLC | 2 | GE Healthcare has become aware that patient mismatch could occur on some Seno Advantage 2.1 and 2.2 Workstations with RIS CCOW integration. This issue may impact patient safety.When Seno Advantage workstation is synchronized with RIS using CCOW (an HL7 standard protocol designed to enable disparate applications to synchronize in real-time); the following issues could occur: 1. Two different | GE Healthcare sent a URGENT MEDICAL DEVICE CORRECTION letter to all affected customers dated May 31, 2011. The letter identified the product , the problem and the action needed to be taken by the customer. GE Healthcare will upgrade all affected workstations at no cost. A GE Healthcare service representative will contact you to arrange for this correction. If you have any questions or concerns regarding this notification, please call one of the following phone numbers: United States: 800 437 1171 Japan: 0120 - 055 - 919 For other countries, please contact your local GE Healthcare Service Representative.
Z-1223-2011 | Feb-14-2011 | 2011 | Leksell GammaPlan - Leksell GammaPlan; Model 5.34.Intended to be used for planning the dosimetry of treatments in stereotactic radiation therapy. | Elekta; Inc. | 2 | Investigation found that if a user accidentally selects the wrong image/tube position images can be displayed flipped. | Product bulletin Leksell GammaPlan - "Leksell GammaPlan version 5.34" (dated October 2005) was sent to affected Customers as notification of the problem. Field Change Order FCO-DCO-05-0006 (dated November 22, 2005) and Technical Notes Leksell GammaPlan Leksell SurgiPlan Important Notice - "Potential issue using flipped angiographic images" (dated October 2005) was sent to Elekta Service Engineers as notification of the problem. The letters described the issue and the FCO asked the Service Engineers to confirm their receipt of the documentation. The Technical Notes also discussed image visualization requirements and listed recommendations.
Z-1252-2011 | Feb-15-2011 | 2011 | Medical linear accelerator (with patient imaging) - Elekta Synergy XVIIntended use: Radiation therapy treatment of malignant neoplastic diseases; as determined by a licensed medical practitioner. | Elekta; Inc. | 2 | Flexmap calibration did not cover the full 360 degrees of gantry rotation (0 to 355 degrees) and that this caused an isocenter displacement with 2D PlanarView images required at 180 to -180. | The firm, Elekta Limited, sent an "Important Notice A344" dated January 5, 2011 to all consignees/customers. The notice described the product, problem and action to be taken by the customers. The customers were instructed to follow any instructions or recommendations covered in the notice. The notice provides suggestions on how to check and prevent this situation from occurring; it instructs the User to file the document in the Important Notice section of the appropriate User Manual; to complete a FlexMap Check each time they acquire a new FlexMap and amend their local procedures to include the FlexMap Check and complete and return the Important Notice Confirmation of Receipt via Fax to: Field Support Administration, Elekta Limited, Linac House, Fleming Way Crawley, West Sussex, RH10 9RR UK +44 (0)1293 654401. If you have any questions, contact Elekta Limited at +44 (0)1293 654200 or www.elekta.com.
Z-1827-2011 | Apr-06-2011 | 2011 | Phoenix x/ray v/tome/x series - Phoenix x/ray v/tome/x series.Phoenix x¦ray v¦tome¦x series cabinet x-ray systems are used in non-destructive testing; failure analysis; process control; and metrology. | GE Inspection Technologies; LP | 2 | Some of the systems do not have a safety interlock on the service access panel. Although the systems have two independent means for indicating when x-rays are being generated; some systems may have one indicator that is not discernible. Also; some systems lack proper placement of a caution label and others do not have this label affixed. Two of the systems also lack the sign indicating the meaning | The firm has completed all of the following necessary actions and no further actions have been indicated: A Field Service Memo was issued to customers. It described the affected product and described the issue. GE Sensing and Inspection provided and installed retrofit kits, and the systems in question were tested by qualified Field Service Engineers at no charge to customers. GE asked customers to post the notification letter on the affected equipment until retrofit was completed. Questions were to be directed to the GE Service Center at 800-766-6062.
Z-3105-2011 | Aug-31-2011 | 2011 | Varian Treatment - Varian Treatment version 6.6.5052.Varian Medical Systems;Palo Alto; Ca Intended to provide accurate treatment set-ups for each patient by monitoring linear accelerator set-up parameters and by preventing the radiation therapy device form commencing irradiation while any linear accelerator parameter is out of conformance with the treatment plan. | Varian Medical Systems; Inc. Oncology Systems | 2 | Potential for Varian Treatment application failing to load the physical wedge filter for a treatment field on a Siemens Mevatron linear accelerator. Delivery of a treatment field without the planned wedge would deliver more dose than intended and would change the shape of the dose distribution for that field. | Varian Medical Systems sent out "Urgent Medical Device Correction/Urgent Field Safety Notice" letters dated July 20, 2011 to all affected customers. The letter identified product, problem and provided recommended user actions when setting up and running the system. It also asked customers to notify appropriate personal provided instructions. For information or questions on this recall please call Varian at (888) VARIAN5 (888- 827-4265).
Z-1884-2011 | Apr-04-2011 | 2011 | TEE Transducer V5M - TEE Transducer V5M; when used with the Acuson S2000 Ultrasound System; software versions: VA15; VA15A or VA15B; Model Numbers: 10041461 (Acuson S2000) and 8263703 (V5M Transducer);Product is manufactured and distributed by Siemens Medical Solutions USA; Inc.; Mountain View; CAProduct is a medical device; a multi-purpose mobile; software controlled diagnostic ultrasound system with and on-screen display for thermal and mechanical indices related to potential bio-effect mechanisms. tis function is to acquire primary or secondary harmonic ultrasound echo data and display it in B-mode; MMode; Pulsed(PW) Doppler mode; Continuous (CW) Doppler Mode; Color Doppler Mode; Amplitude Doppler Mode; a combination of modes; or harmonic imaging and 3d/4d imaging on a flat panel display. The device system is intended for the following applications: Fetal; Abdominal; Intraoperative; Pediatric; Small Parts; Transcranial; OB/GYN; Cardiac; Pelvic; Neonatal/Adult Cephalic; Vascular; Musculoskeletal; Superficial Musculoskeletal; and Peripheral Vascular applications. The system also provides ability to measure anatomical structures and calculation packages that provide information to the clinician that may be used adjunctively with other medical data obtained by a physician for clinical diagnosis purposes. The device is a multi frequency; multi plane; phased sector array transducer. In the TEE (transesophageal echocardiography process where the device fails; the transducer is introduced into the esophagus of the patient to obtain images of the heart structure. | Siemens Medical Solutions USA; Inc. | 2 | Device function failure--When performing a TEE exam with the transducer; CW (Continuous Wave) mode is not functional. Due to work flow; this may not be apparent until the transducer is inserted into the patient's esophagus. | Siemens sent a Customer Safety Advisory Notification letter on April 7, 2009, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to pass the notice on to all those within their organization who need to be aware of the issue until the corrective action is completed. For questions regarding this recall call 650-969-9112.
Z-2983-2011 | Aug-10-2011 | 2011 | IMPAC / Elekta Medicals Sequencer - LANTIS Oncology Information System (OIS) Manufactured by IMPAC/Elekta medical as Multi -Access OIS for Siemens.An oncology digital imaging information management system. Lantis OIS in combination with Sequencer and a non-Siemens linac interface.System components include: LANTIS Commander; LANTIS system; LANTIS 6.1 Commander; LANTIS 8.3 Commander.Subcomponents include: Linac Interface; other vendor; LANTIS Varian Clinac; LANTIS GE Saturn; LANTIS Philips SLS;LANTIS Philips 75 ASU; LANTIS Elekta (Philips) and LANTIS Varian EX I/F. | Siemens Medical Solutions USA; Inc | 2 | A potential safety issue exists for data transfer between LANTIS OIS system (containing PRIMEVIEW or PRIMEVIEW Data Gateway) and the Sequencer V&R system via a third party interface is used. | Siemens sent an "Update Instructions/Customer Safety Advisory Notice" to all affected users, beginning June 30, 2011. The letter includes description of the product and problem and provides updated instructions to ensure data is stored in the system. If you have questions contact Siemens at (925) 602-8157.
Z-0305-2012 | Nov-29-2011 | 2011 | Medical linear acelerator (with patient imaging) - Elekta Synergy XVI 4.5Product Usage: Radiation therapy treatment of malignant neoplastic diseases; as determined by a licensed medical practitioner. | Elekta; Inc. | 2 | If "Confirm" or "Unconfirm" settings are clicked during the "Terminated Checking" phase of treatment delivery; Desktop Pro cancels the loaded beam and sets the Beam MUs to zero. Desktop Pro then sends the delivered MUs to MOSAIQ a a zero value. | Elekta sent an IMPORTANT NOTICE letter dated March 9, 2011 to affected customers. The notice identified the product, the problem, and the corrective action to be taken by the customer. The notice advise all customers to follow any instructions or recommendations covered in the Notice. For questions call +44(0)1293 654200.
Z-2149-2011 | May-05-2011 | 2011 | 4Z1c Transducer - 4Z1c Matrix Array Transducer; model number 10033682; used with ACUSON SC2000 ultrasound systems; model number 10433816; manufactured by Siemens Medical Solutions; Ultrasound Business Unit; Mountain View; CA.The 4Z1c transducer is a special purpose ultrasound transducer designed to help clinicians perform adult echocardiography and pre-operative and intra-operative surgical assessments. Secondary applications include pediatric echocardiography; primarily benefiting older and larger pediatric populations; general abdominal and intra-operative pediatric abdominal imaging and peripheral vessel imaging. | Siemens Medical Solutions USA; Inc. | 2 | Burn risk-- due to coolant leaks; patients may suffer burns or blisters due to increased temperature. | A customer safety advisory notification was sent out to customers on 7/10/09, and a mandatory software update will also be conducted.
Z-2218-2011 | May-19-2011 | 2011 | AVOXimeter 1000; AVOXimeter 4000 - AVOXimeter 1000; AVOXimeter 40001000- portable bedside whole blood oximeter that directly measures oxyhemoglobin fraction; the total hemoglobin concentration; and oxygen content. It has enhanced software that optimizes its use int he cardiac catheterization laboratory through calculations of physiologic equations. 4000-a portable bedside whole blood CO-oximeter that directly measures total hemoglobin concentration and the relative concentrations of oxyhemoglobin carboxyhemoglobin concentration methemoglobin. In addition; oxygen content oxygen capacity and oxygen saturation indices are calculated. | International Technidyne Corp. | 2 | An incorrect calibration code that affects the total hemoglobin measurement was applied to affected Avoximeter instruments. | International Technidyne Corp. sent a letter dated March 23, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their inventory to determine if they have any of the affected product. If they have any of the affected product, customers were instructed to stop using it and remove it from their inventory. Customers were also instructed to complete the attached Customer Account Tracking Form and return it by fax to (732-635-0144) or e-mail (techsupport@itcmed.com) or mail to: ITC Technical Support 20 Corporate Place South Piscataway, New Jersey 08854 For any questions call the Distributor or ITC Nexus Dx Technical Support at 800-631-5945 (US) or 732-548-5700 (International), ext 4707.
Z-0052-2012 | Oct-17-2011 | 2011 | SEQUENCER - Impac Medical Systems SequencerProduct Usage: The intended us of Sequencer is to assists in the process of patient treatment on a radiotherapy treatment machine. Sequencer provides the capability to notify the user of actions that need to take place prior to treatment; to display reference images for set up purposed and to auto setup the machine to predefined settings. The actual settings are read from the treatment machine; via the machines communication interface; and compared to predefined values. if a mismatch occurs between the planned values and the actual machine settings; treatment is prohibited. At appropriate points during the treatment; the actual delivered values are recorded to provide treatment tracking. | Impac Medical Systems Inc | 2 | Varian reported a problem related to certain collimator rotation values which are incorrect on TrueBeam compared to planned values in Mosaiq. | " Impac Medical Systems Inc., sent an USER NOTICE letter to all affected consignees on September 20, 2011. The letters identified the source characterization error, clinical impact, and final resolution. The letter states that Elekta Support will follow up with each consignee and remotely correct the source characterization. Customers were instructed to sign, date and return the attached Confirmation of Receipt Form by one of the three ways indicated. The letter instructs consignees to distribute this notice to any and all users of IMPAC software that may potentially be affected by the issue. For questions, concerns, and requests for upgrade, email support@impac.com or call. The letter provided consignees with a complete list of contact information..
Z-1572-2011 | Mar-07-2011 | 2011 | MOSAIQ Oncology Information System - MOSAIQ Oncology Information SystemIntended Use : Electronic Patient Records Management | Siemens Medical Solutions USA; Inc | 2 | The potential for injury to a patient should the treatment data be incorrectly uploaded to patient records. | Sent an Urgent Medical Device Correction Safety Advisory Notice to customers on 1/11/11. Letters will be delivered by Siemens Customer Service Group or sent by certified mail. Letter identified the affected product and stated there is a potential risk due to a problem identified by the manufacturer. Customers are asked to refer to the attached Safety Notice from Elekta IMPAC Software. A Siemens Service Engineer will contact customers site to arrange for a patient matching criteria audit to determine if you have a problem. Customers should include the Safety Advisory Notice including the Safety Notice from Elekta IMPAC Software in their product related documentation until further notice. If customers have any questions, they should contact their responsible SIEMENS Service Engineer or Application Specialist.
Z-2241-2011 | May-23-2011 | 2011 | NeuViz 16 Multi-Slice CT Scanner System - NeuViz 16 Multi-Slice CT Scanner System is a computed Tomography X-ray system; and has a Part number 989605858501.The device is labeled in part: "***Neusoft***NeuViz16 MULTI-SLICE CT SCANNER SYSTEM***MANUFACTURED FOR NEUSOFT MEDICAL SYSTEMS CO.; LTD***MANUFACTURED BY PHILIPS AND NEUSOFT MEDICAL SYSTEMS CO.; LTD***ADDRESS: NEUSOFT PARK; HUNNAN INDUSTRIAL AREA; SHENYANG***CHINA***". | Philips And Neusoft Medical Systems Co.; Ltd. | 2 | Dimension measurement displays on a combined image were found inaccurate after the clinician merged multiple image series into one image. | Neusoft Medical Systems sent an Urgent Device Correction Notification letter dated February 9, 2011, via certified mail to their distributor NEUISYS. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to discontinue use of combined images. Instead, use MPR Batch for diagnostic reviews until the planned upgrade is provided to the product. Neusoft Medical System and Philips Medical System released the mandatory Field Change Order (FCO) by March 31, 2011. All nine affected systems will be updated with aforementioned kits in six months after the release. Consignees were instructed to contact Service Support Department by e-mail to nms-service@neusoft.com.
Z-2919-2011 | Aug-03-2011 | 2011 | XiO Radiation Treatment Planning System - XiO Radiation Treatment Planning System; XiO versions equal to or greater than 2.4.0 and less than 4.63.0The XiO RTP System is used to create treatment plans for any cancer patient for whom external beam radiation therapy or brachytherapy has been prescribed. | Computerized Medical Systems Inc | 2 | The XiO Software currently prompts for entry of CT to ED data for a studyset. No indication is given that the Electron Density data should only be used for Photon and Electron calculations and either Relative Stopping Power or Mass Density should be used for Proton calculations (depending on the algorithm). | The firm, Eletka, Inc., issued an "Important Safety Notice" dated May 18, 2011 via mail to all affected sites. The notice explains the issue and identifies a work around to the problem until the new release is available. The customers were instructed to complete and return the enclosed "Confirmation of Receipt" form via FAX to 314-993-1175, attn: QA; e-mail or mail: CMS Software, Attn: QA, 13723 Riverport Drive, Suite 100, Maryland Heights, MO 63043. The notice states that "this problem has existed since XiO Release 2.4.0 and has been resolved in XiO Release 4.63.0 by adding warning messages and explanations to the XiO On-Line Help. Your site will be notified when this release is available." If you have any question, call 314-993-0003, 800-878-4267 or 408-830-8023.
Z-2066-2011 | Apr-26-2011 | 2011 | OB-Link release 2.0.7 - OB-Link release 2.0.7. distributed by Digisonics; Inc.; 3701 Kirby Drive; Houston TX 77098OBLink provides automated transmission of patient data and measurements from ultrasound systems directly to Digisonics OB-View/OB-Windows systems. | Digisonics; Inc | 2 | Data/measurements from a OB report sent from ultrasound systems populated the wrong OB-View report | Digisonics visited consignee in April 2008 as follow-up to complaint.
Z-1676-2011 | Mar-16-2011 | 2011 | CRS Remote - IMPAX Remote Cardiology Review Station; a cardiovascular information system. Model Number CV7.8 | AGFA Corp. | 2 | A discrepancy in the validation testing which resulted in the product not performing as intended. | AGFA sent an "URGENT SAFETY NOTICE" letter dated 2/10/11 via FED-EX to the customer. The letter included an acknowledgment which was to be FAX-Back indicating that the information was received and understood. Agfa has discussed with the customer how the de-installation of their CRS Remote product - Software Version 7.8_HL would be performed.
Z-0263-2012 | Nov-22-2011 | 2011 | Infinity(R) Acute Care System (IACS) Monitoring Solution (M540) - Infinity(R) Acute Care System (IACS) Monitoring Solution (M540) - Physiological Patient Monitoring SolutionNetworked solution/system used to monitor patients' vital signs and therapy; control alarms; review Web-based diagnostic images; access patient records; and scale up functionality based on patient acuity level. | Draeger Medical; Inc. | 1 | During product training on the Infinity(R) Acute Care System (IACS) Monitoring Solution (M540) it was discovered that the weight-based drug dosage calculation may indicate incorrect recommended values. In addition; a misalignment of ECG and ART waveforms delivered from an Infinity(RO Acute Care System (IACS) Monitoring Solution (M540) was observed on the Infinity Central Station. | The firm initiated their recall of this device on October 17, 2011 by issuing a letter to the consignee instructing customers to follow the Instructions for Use and use the Infinity Central Station only for remote assessment of a patient's status. They recommend that you do not consider the waveforms on Infinity Central Station for primary diagnosis purposes. In case of a high priority alarm situation, "Check the patient and treat if necessary". If the misalignment issue is observed, it can be cleared by undocking and redocking the M540 or by changing the order of waveforms displayed on the Cockpit screen (or any other change to the auto display screen layout).
Z-0033-2012 | Oct-11-2011 | 2011 | Sodium Bicarbonate Mixing and Delivery System Software v 7.1. Model # MD-380 - Sodium Bicarbonate Mixing and Delivery System Software v 7.1. Model # MD-380; Serial Numbers: 5103031102; 5101101101; 5102161007; 5103221007; 5104221007; 5105101008; 5109131010; 5111231011; 5114301012; 5102131101 & 5113091012. Manufactured by: Isopure Corp; Simpsonville; KYSystem is used to mix and distribute bicarbonate as a part of a clinical dialysis system. | Isopure Corp | 3 | Software problem causes transfer line between mixing and storage tank to not be rinsed after disinfection; leaving residual chlorine in the line. | Isopure sent a letter dated May 19, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Isopure devised an immediate solution which must be implemented at their faclity in conjunction with the disinfection procedure. Customers were asked to inform all appropriate personnel, clinical and technical, who may be involved in operating the system during the normal disinfection schedule. On 5/19/2011, the firm sent letters to their customers via US mail and email. They also telephoned their customers to notify them of the problem.
Z-1523-2011 | Mar-04-2011 | 2011 | Siemens brand ARTISTE System Digital Linear Accelerator with the syngo RT Therapist version 4.1 Conn - Siemens brand ARTISTE System Digital Linear Accelerator with the syngo RT Therapist version 4.1 Connected to MOSAIQ; Part No. 8168754; Distributed by Siemens Medical Solutions Oncology Care System; Concord CA Medical charged-particle radiation therapy system. The intended use of the SIEMENS branded ARTISTE; ONCOR and PRIMUS family of linear accelerator systems is to deliver X-Ray photon and electron radiation for the therapeutic treatment of cancer. | Siemens Medical Solutions USA; Inc | 2 | A software problem that has the potential for data loss during the transfer of treatment records to the OIS; which may not be recorded in MOSAIQ; and subsequently may lead to mistreatment. | Siemens sent Urgent Medical Device Correction, Safety Advisory Notices to all affected customers on June 29, 2010. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to include the Safety Advisory Notice in their System Owner Manual, Chapter, "Safety Advisory Letter. For questions customers were instructed to contact their Siemens Service Engineer or Application Specialist. For questions regarding this recall call 925-602-8157.
Z-3098-2011 | Aug-31-2011 | 2011 | 4D INTEGRATED TREATMENT CONSOLE - Varian brand 4D Integrated Treatment Console (4DITC) Radiation Therapy; version 8.1 SP2 to 10.0 SP1; Model Number: H51; Reference/PSCA Identifier: CP-03940; Product is manufactured and distributed by Varian Medical Systems Inc.; Palo Alto; CAProduct Usage: 4D Integrated treatment console is designed to assist the operator of a radiation device in providing accurate treatment set-ups for each patient by monitoring set up parameters and preventing the radiation therapy device from commencing irradiation while any parameter is out of conformance with the treatment plan. | Varian Medical Systems; Inc. Oncology Systems | 2 | Varian has scheduled for an upgrade of the device from current version to 10.2. There are circumstances in which the 4DITC may be involved in a failure to correctly record treatment history information. | Varian Medical Systems, Inc. sent an "Urgent Medical Device Correction/Urgent Field Safety Notice" letter dated June 23, 2011 to all affected customers. The letter identified the product, problem and recommended user actions to follow. Customers were instructed to notify all appropriate personnel of the correction notice. The letter states that the recalling firm will issue a Customer Technical bulletin to inform users of the mandatory upgrade and related information. For future reference, this document will be posted to the Varian customer support website: http://www.MyVarian.com. Any questions regarding this Product Notification letter, contact Varian Medical Systems Help desk at 1-888-827-4265 for USA and Canada and +41417498844 for Europe.
Z-0886-2011 | Jan-08-2011 | 2011 | Varian brand C-series Clinac; Trilogy and Novalis Tx; Software Versions 6.X and 7.X; - Varian brand C-series Clinac; Trilogy and Novalis Tx; Software Versions 6.X and 7.X; Model Numbers: H14; H27; H29; HCX; Product is manufactured and distributed by Varian Medical Systems Inc.; Palo Alto; CA The Trilogy Radiotherapy Delivery System is a radiation therapy accelerator intended to deliver megavoltage x-ray treatments for conventional radiotherapy (three dimensional conformal radiotherapy and intensity modulated radiotherapy) and stereotactic radiosurgery and radiotherapy. Stereotactic treatments are intended for therapy of lesions; e.g.; arteriovenous malformations; primary tumors and metastases. Stereotactic treatments may be intracranial or extra cranial and consist of single-session or fractionated delivery. | Varian Medical Systems Oncology Systems | 2 | The Auto goto or Auto setup functions in C-Series version 7 software ignore the couch angle and exceed motion zone (unless a tolerance is defined by the user); it has the potential for collision with the patient on the couch. | Varian sent a Customer Technical Bulletin letter to all affected users on May 27, 2010, and was also posted to Varian's customer support site: http://www.MyVarian.com. The letter identified the product, the problem, and provided instructions on the recall. Customers were instructed to return the proof of notification (via completed "Proof of Notification Form or appropriate Postal system certified receipt of delivery to Varian. For questions regarding this recall call 650-483-3153.
Z-1652-2011 | Mar-14-2011 | 2011 | Factor II (prothrombin) G20210A kit - Factor II (prothrombin) G20210A kit; catalog number 03610195001; Roche Diagnostics; Indianapolis; IN The Factor II (prothrombin) G20210A kit allows the detection and genotyping of a single point mutation (G to A at position 20210) of the human factor II gene; from DNA isolated from human whole peripheral blood. The test is performed on the LightCycler 2.0 instrument utilizing polymerase chain reaction (PCR) for the amplification of Factor II DNA recovered from clinical samples snd fluorigenic target-specific hybridization for the detection and genotyping of the amplified Factor II DNA. The Factor II (prothrombin) G20210A test is an in vitro diagnostic test for the detection and genotyping of the Factor II (Prothrombin) G20210A mutation as an aide to dignosis in the evaluation of patients with suspected thrombophilia. The test is intended to be used on the LightCycler 2.0 instrument using the LightCycler Software 4.05 or 4.1. | Roche Diagnostics Operations; Inc. | 2 | If the bar code is printed from the Batch Results Screen; a wrong batch ID may be printed; while the correct batch ID is still assigned to the batch in the data repository of the MPLC 2.0 software. This may result in sample mis identification. If printed from the ordering screens; the bar code is printed with the correct batch ID. | The firm, Roche, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated November 04, 2010 to its customers. The letter described the product, problem and action to be taken. The customers were instructed to print batch ID barcode labels from the ordering or stage set up screens only; print Cooling Block barcode labels from the Post Elution Edit screen only; complete and return the attached UMDC Faxback Form for MagNA Pure LC 2.0 Instrument [UMDC 10-228] via fax to 1-877-551-7152 and file this correction notice for future reference. The firm plans to issue a software update in first quarter 2011. If you have questions about the information contained in this UMDC, for Clinical customers, please contact Roche Molecular Diagnostics Technical Service at 1-800-526-1247; for Research customers, please contact Roche Applied Sciences Technical Service at 1-800-262-4911.
Z-1836-2011 | Mar-28-2011 | 2011 | ACUSON and SONOLINE Antares and Antares PE ultrasound systems - ACUSON and SONOLINE Antares and Antares PE ultrasound systems at product versions 4.0 and 5.0 with cardiovascular options installed;ACUSON Antares 4.0 & 5.0: 10032747; 10037592; 10040728;ACUSON Antares PE 4.0 & 5.0: 10032746; 10037591; 10038202; 10040729;SONOLINE Antares 4.0 & 5.0: 05936518; 08653771; 10037593;Antares Cardio Options: 10036613; 10036588; 10036585; 10436250;Product is manufactured and distributed by Siemens Medical Solutions USA; Inc.; Business Unit Ultrasound; Mountain View; CAThe ACUSON Antare is a general purpose Diagnostic Ultrasound system. Its function is to acquire primary or secondary harmonic ultrasound echo data and display it in: B-Mode; M-Mode; Pulsed (PW) Doppler Mode; Continuous (CW) Doppler Mode; Color Doppler Mode; Amplitude Doppler Mode; a combination of modes; Harmonic Imaging; or 3D/4D imaging on a FPD display. The Antares ultrasound imaging system is intended for the following applications: Abdominal; Intraoperative; Small Parts; Transcranial; OB/GYN; Cardiac; Transesophageal; Pelvic; Neonatal/Adult Cephalic; Vascular; Intravascular; Musculoskeletal; Superficial Musculoskeletal; Great Vessel; and Peripheral Vascular applications. The system also provides for the measurement of anatomical structures that provide information that is used for clinical diagnosis purposes. | Siemens Medical Solutions USA; Inc.; Mountain View; CA | 2 | There is an error in the estimation of mean pressure gradient in the Cardiac measurement package. The currently implemented calculation will result in an underestimation of the mean pressure gradient (PGmean) when estimated from a trace of the Doppler waveform of the stenotic flow. | Siemens Medical Solutions USA, Inc., sent a Customer Safety Advisory Notification letter on August 17, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to pass the notice on to all those within their organization who need to aware of the issue until the corrective action was completed. For questions customers were to contact their local service support person. For questions regarding this recall call 650-694-5993.
Z-1881-2011 | Apr-04-2011 | 2011 | iSite PACS with Advanced Visualization Tools (AVT) - iSite PACS with Advanced Visualization Tools (AVT) R7.4V6L1; Versions: 3.6.87; 3.6.92; 4.1.87; and 4.1.92; Product is manufactured and distributed by Philips Medical Systems; Foster City; CAiSite PACS: iSite PACS is an image management system intended to be used by trained professionals; including but not limited to physicians; nurses and medical technicians. The system is a software package used with general purpose computing hardware to acquire; store; distribute; process and display images and associated data throughout a clinical environment. The software performs digital image processing; measurement; communication and storage. iSite supports receiving; sending; printing; storing and displaying studies received from the following modality types via DICOM: CT; MR; NM; US; XA; PET; DX; DR; RF; RT; MG; SC; VL; as well as hospital/radiology information systems. | Philips Medical Systems | 2 | A software anomaly that occurs when clinicians use Volume Vision to view images in 3D and uses the measurement tool to measure regions of interest. This could lead to an incorrect interpretation of the measurements which could cause an incorrect diagnosis and incorrect treatment plan. | Philips sent an Urgent Field Safety Notice letter dated February 23, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to assure that all colleagues were aware of the problem. When using the Volume Vision software, check the correctness of measurements (graphics and numerical values) prior to saving a captured image containing measurements. Save captured images containing measurements immediately after creating them and before altering the measurements in the original image view. For any subsequent measurements, create a new capture and save it immediately. For further information or support concerning this issues, customers should contact their local Philips representative at +1-877-328-2808.
Z-1900-2011 | Apr-06-2011 | 2011 | ACUSON S2000 Ultrasound System - Siemens brand ACUSON S2000 Ultrasound System with software versions:VA10; VA10A; VA10B; VA10C; VA15; Model Number: 10041461;Product is manufactured and distributed by Siemens Medical Solutions USA; Inc. Ultrasound Business Unit; Mountain View; CAThe ACUSON SC2000 ultrasound imaging system is intended for the following applications: Cardiac; Neo-natal and Fetal cardiac; Pediatric; Transesophageal; Adult Cephalic; Peripheral Vessel; Intraoperative Neurological; Musculo-skeletal Conventional; and Musculo-skeletal Superficial applications. | Siemens Medical Solutions USA; Inc. | 2 | The product has a software problem in which previous patient measurement data gets associated with another patient's image. | The firm, SIEMENS, issued a "CUSTOMER SAFETY ADVISORY NOTIFICATION" letter to all customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure that the first image in Exam Review is from the current patient; ensure that the Heart Cycle measurement tool is not selected on the Cale menu when entering Exam Review; be aware of this issue as described in the letter and perform four (4) or less measurements of the same measurement label. Should you have any questions, please call (650) 694-5398.
Z-0009-2012 | Oct-04-2011 | 2011 | Plum A+ Single Channel Infusion Pumps - Plum A+ Single Channel Infusion Pumps; Hospira; Inc.; Lake Forest; IL 60045; the pumps were sold under the following configurations: a) list 11971 b) list 11973 - software version 10.3 c) list 12391 - software version 11.3 d) list 20679 - Hospira MedNet Software e) list 20792 - Driver Product Usage: The Plum A+ dual-line volumetric infusion system designed to meet the growing demand for hospital-wide; as well as alternate site and home healthcare; standardization. With their primary line; secondary line; and piggyback fluid delivery capability; the Plum A+ is suited for a wide range of medical/surgical and critical care applications. Full compatibility with LifeCare Plum Series administration sets and accessories and the LifeShield¿ needleless protection systems makes the Plum A+ a convenient and cost-effective infusion system. | Hospira Inc. | 2 | Hospira has received reports of continuous recycling and/or rebooting of Plum A+ infusion pumps when the backlight intensity and/or display contrast settings for the LED display has been adjusted from the original default setting. This condition could cause a delay in therapy if the clinician is unable to start the device due to the continuous recycling and/or rebooting. | The firm, Hospira sent "URGENT DEVICE RECALL" letters dated September 16, 2011 to all affected customers. The letters identified the products, problem and action(s) to be taken. The customers were instructed to read the information provided carefully, complete the attached Reply Form and return it to Hospira via fax at 1-888-345-5358. The letter states that a software upgrade to address the issue is currently being developed. Customers were instructed to notify their consignees if they have further distributed the affected product. The customers were requested to contact Hospira Advanced Knowledge Center at 1-800-241-4002, option 4, for technical assistance.
Z-0281-2012 | Nov-23-2011 | 2011 | Draeger Infinity(R) Delta; Delta XL; Gamma XXL and Kappa Monitors - Draeger Infinity(R) Delta; Delta XL; Gamma XXL and Kappa Monitors are used to monitor; physiologically; the vital signs of patients (with Arrhythmia Detection or Alarms). | Draeger Medical; Inc. | 2 | Firm received 2 complaints reporting Ventricular Fibrillation (Vfib) arrhythmias were not recognized by the monitor and did not alarm when the Masimo SET pod is used for SpO2 measurements. | Firm initiated their recall of this device on October 25, 2011 by sending an "Urgent Medical Device Recall" letter to consignees. A Technical Service Bulletin will be generated and distributed globally that instructs the affected International Draeger subsidiaries to inform customers in their countries. The software anomaly has been corrected in software version 7.3. Draeger Service Representatives will schedule time with customers to upgrade the software in affected monitors. The software upgrade will be provided free of charge.
Z-0007-2012 | Oct-03-2011 | 2011 | Outlook ES Safety Infusion System - Outlook ES Safety Infusion System; Model 621-100ES; 621-200ES; 621-300ES; and 621-400ESB. Braun Medical Inc.; 1601 Wallace Drive; Suite 150; Carrollton; TX 75006The Outlook ES pump is intended for use with B. Braun Medical Inc.; Horizon Pump IV Sets to regulate the flow of primary and secondary fluids when positive pressure is required. | B Braun Medical; Inc | 2 | The Outlook ES Dose Guard Care Area can inadvertently be exited when the pump enters the KVO (Keep Vein Open) state. The same condition can occur when a sequence of HOLD-HOLD key strokes are performed on the pump key panel. | The firm, B. Braun Medical Inc., sent an "URGENT: FIELD CORRECTION" letter dated September 8, 2011 to its customers. The letter described the product, problem and actions to be taken. The firm informed the customers that a B. Braun representative will be contact ing them to schedule a field software upgrade of the affected pumps. The customers were instructed to complete and return the FIELD CORRECTION INFORMATION FORM via mail using the enclosed envelop and contact B. Braun Customer Service for shipping instructions at 1-800-627-7867 if product is to be returned. Should you require any additional information, please do not hesitate to contact B. Braun Medical Inc., Technical Product Support at 1-800-627-7867 (800-627-PUMP).
Z-2788-2011 | Jul-11-2011 | 2011 | Eclipse Treatment Planning System - Eclipse Treatment Planning System;Varian Medical Systems; Palo Alto; CARadiotherapy treatment for patients with malignant and benign diseases. | Varian Medical Systems; Inc. Oncology Systems | 2 | A software anomaly has been identified with the Eclipse Monte Carlo (eMC) algorithm where not-equidistant data points in the open beam profile used during configuration may lead to inaccurate results. | Varian sent an URGENT MEDICAL DEVICE CORRECTION letter via certified mail dated May 20, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to remove the product from use and advise users to manually initiate the configuration process for resampling to be applied when a previous calculation model is copied. Customers were asked to sign and return the notice to confirm receipt. For questions call USA and Canada 1-888-827-4265 For Europe call +41 41 749 8844.
Z-1953-2011 | Apr-12-2011 | 2011 | COULTER¿ LH750 Analyzer - Coulter LH750 AnalyzerThe COULTER LH 750 Hematology Analyzer is a quantitative; automated hematology analyzer and leukocyte differential counter For In Vitro Diagnostic Use in clinical laboratories. The COULTER LH 750 Hematology Analyzer provides automated Reticulocyte analysis and enumeration of nucleated red blood cells (NRBCs) as well as an automated method for enumeration of RBCs and WBCs in body fluids. | Beckman Coulter Inc. | 2 | This recall was initiated because Beckman Coulter identified a problem with some date fields that are transmitted to the Laboratory Information System (LIS). Starting Jan 1; 2011; the date fields will transmit as a blank field to the LIS if the year is greaterthan 2010. The following dates are affected: Run date; Control expiration date; Date of birth(For individuals born in 2011 or later); a | The firm, Beckman Coulter, sent an "URGENT: PRODUCTION CORRECTION" letter dated January 31, 2011 with attached PCA Response Form (via US Postal Service for US Customers) to all customers who purchased the Coulter LH 750 Analyzer, Coulter LH 780 Analyzer, Coulter LH 500 Series System, and Coulter GEN S System. The letter describes the product, problem and actions to be taken. The customers were instructed to work with their US vendor to assess how this affects their laboratory. The actions they choose are dependent on the configuration of their LIS. ACTION FOR ISSUE 1: Use Sample Drawn Date and Date of Birth information available at your LIS in lieu of the analyzer transmitted fields BIRTH and UDATE. Use the current date when the results are received at the LIS in lieu of the transmitted Run Date [DATE]. Print results from the Workstation to keep a copy of the analysis date on the analyzer, if necessary. Use the control expiration date from the analyzer display or printout. ACTION FOR ISSUE 2: If you use the LIS 1G1 transmission format, use the General Information group field names in place of the Retic Parameter group field names to identify each reticulocyte run (see Table in Issue 2 section). Check with your LIS vendor to determine if it is possible to switch to the LH 750 Workstation transmission format. Format requirements are LIS driver specific. For patients with Dates of Birth in 2011 for both Issue 1 and Issue 2, results should be reviewed to determine if decision rules or age-related limits sets were triggered inappropriately. RESOLUTION These issues will be corrected in a future software release for the LH Series products. The GENÂ®Â®S product has been discontinued; no software revisions will be forthcoming for the GENÂ®Â®S product line. Customers were also instructed to complete and return the PRODUCT CORRECTION RESPONSE FORM via fax to 786-639-7500/7501/7502/7503/7504; mail to Beckman Coulter, Inc., Attn: Regulatory Affairs Mail Code 310B06, 11.
Z-3277-2011 | Sep-23-2011 | 2011 | CareSuite" - Critical Care Manager; PACU Manager and Anesthesia Manager - CareSuite" - Critical Care Manager; PACU Manager and Anesthesia Manager Software; transmission & storage; patient data.This product is a non-alarming software patient information system that compiles an electronic medical record utilizing commonly available hardware. A patient record is populated with information from various sources; such as health care professionals; medical devices connected to the system; and data that arrives via hospital and laboratory information systems. The application stores this information in a database and displays it to healthcare providers in different formats for informational purposes. The product is intended for use by healthcare professionals. | Picis Inc. | 2 | Electronic health record software applications malfunctioned in a manner which failed to display the discontinued status for medication orders; as entered by the attending physician | Picis Inc. sent a customer notification letter dated August 11, 2009 to all affected customers. The letter describes the product, problem, and action to be taken. Customers informed of the potential patient safety error were advised to make the appropriate configuration change, within current software versions 7.1.4 or later. Contact your local Picis support representative for questions regarding this notice.
Z-0057-2012 | Oct-18-2011 | 2011 | MicroMaxx ultrasound system and M-Turbo ultrasound system - MicroMaxx ultrasound system and M-Turbo ultrasound system is Medical Diagnostic Ultrasound.General purpose ultrasound system intended for use by a qualified physician for evaluation by ultrasound imaging or fluid flow analysis of the human body. | Sonosite; Inc. | 2 | This is a software anomaly. In some circumstances; if the Systolic/Diastolic (S/D) ratio value is interpreted directly from the display; the clinician could potentially use an incorrect value of the S/D ratio in his/her assessment. | The firm SonoSite, Inc., sent an "Urgent - Medical Device Correction" letter dated September 29, 2011 to its foreign and domestic customers. The letter described the product, problem, and actions to be taken. The customers were instructed to follow the three (3) steps provided in the letter to verify if the extended OB/GYN Calculations (EOBC) is enabled on their system. Additionally, SonoSite will perform a software upgrade to M-Turbo and/or MicroMax ultrasound system in order resolve the issue. The customers were instructed to contact Sonosite Technical Support via website: www.sonosite.com/sdratio; email: regulatory@sonosite.com (include contact information and address), or call in U.S. and Canada: 1-877-590-4982. If you have any questions or concerns regarding this matter, do not hesitate to contact in US and Canada at 1-877-657-8118, Monday-Friday, 6 am-5 pm PST. Customer can also send an email to Technical Support at service@sonosite.com.
Z-0911-2011 | Jan-18-2011 | 2011 | Varian Eclipse - Varian Medical Systems; Eclipse Client software versions with build numbers from 7.3.10 through 8.6.15 (i.e. excluding 8.6.17); Model Number: H48; Radiation Treatment Planning System | Varian Medical Systems; Palo Alto; CA | 2 | Possible incorrect dose calculation: An anomaly in certain Eclipse versions where; under certain specific sequences of user actions; the dose distribution may fail to be re-calculated even though changes have been made which affect dose. | Varian Medical Systems Urgent Medical Device Correction letters were sent to all affected users on August 6, 2010. Letters were also sent to Varian Sales, Marketing and Service organizations to inform them of the recall.
Z-1931-2011 | Apr-08-2011 | 2011 | CARESCAPE Monitor B850 - GE Healthcare; CARESCAPE Monitor B850The CARESCAPE Monitor B650 is a multi-parameter patient monitor intended for use in multiple areas and intrahospital transport within a professional healthcare facility. The CARESCAPE Monitor B650 is intended for use on adult; pediatric; and neonatal patients and on one patient at a time. The CARESCAPE Monitor B650 is indicated for monitoring and recording of; and to generate alarms for; hemodynamic (including EGG; ST segment; arrhythmia detection; EGG diagnostic analysis and measurement; invasive pressure; non-invasive blood pressure; pulse oximetry; cardiac output; temperature and mixed venous oxygen saturation); impedance respiration; airway gases (C02; 02; N20 and anesthetic agents); spirometry; gas exchange; and neurophysiological (including electroencephalography; Entropy; Bispectral Index (BIS) and neuromuscular transmission) status. The CARESCAPE Monitor B650 can be a stand-alone monitor or interfaced to other devices. It can also be connected to other monitors for remote viewing and to data management software devices via a network. The CARESCAPE Monitor B650 is intended for use under the direct supervision of a licensed healthcare practitioner; or by personnel trained in proper use of the equipment in a professional healthcare facility. The CARESCAPE Monitor B650 is not intended for use during MRI. | GE Healthcare; LLC | 2 | GE Healthcare has become aware of a potential safety issue associated with the CARESCAPE Monitor B650 & B850 Cardiac Output measurement using Patient Data Module (PDM); and activation of ECG Leads Off alarm in specific cases.1. When the CARESCAPE Monitor is used to measure Cardiac Output values with the Patient Data Module (PDM) or TRAM Module; Cardiac Output values may be calculated incorrectl | The firm, GE Healthcare, sent "Urgent Medical Device Correction" letters dated March 23, 2011 to its consignees/customers. The letters were addressed to Healthcare Administrator / Risk Manager, Chief of Nursing, and Director of Biomedical Engineering. The letters described the Safety Issues, Safety Instructions, Affected Product Details, Product Correction, and Contact Information. The customers were instructed to follow the safety instructions documented in the letter. GE Healthcare is working on a software correction that will be provided to their customers once it is available and will contact the customer to schedule the software update. If you have any questions regarding this medical device correction or the identification of affected items, please contact Technical Support at 1-800-558-7044, or your local Service Representative.
Z-2126-2011 | May-02-2011 | 2011 | Cirrus HD-OCT Model 400 instruments - Cirrus HD-OCT Model 400 instruments; new and refurbished; with or upgraded to; version 4.6 or 5.1.0 software.The Cirrus HD-OCT is a non-contact; high resolution tomographic and biomicroscopic imaging device. | Carl Zeiss Meditec Inc | 2 | A software defect exists in versions 5.0.0 and 5.1.0 where when following a specific workflow; exam data may be saved under the name of the first patient | Carl Zeiss Meditec, Inc. sent an Field Correction Action letter dated March 9, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The letters requested acknowledgement of the letter and a response and the fact that a follow up packet with revised software will be sent to all customers with V 5.0.0 or 5.1.0 software. For questions, please call 1-855-853-9461.
Z-2040-2011 | Apr-22-2011 | 2011 | Orthoralix 8500 DDE Panoramic Dental X-Ray System - Orthoralix 8500 DDE Panoramic Dental X-Ray System; Extraoral x-ray unit with timer; Gendex Dental Systems; Des Plaines; IL 60018; part numbers 110-0199G2; 110-0199G4; 110-0234G1 and 110-0234G3. The intended use of the device is to be used as an extraoral source of x-rays for imaging of the dento-maxillo-facial area. | Gendex Dental Systems | 2 | There is an anomaly in the Orthoralix firmware which can cause a message corruption which will distort the patient image being transferred from the X-Ray system to the PC. This image distortion/corruption (distorted; fragmented; flipped; negative; reversed or had lines across them) is most likely to occur during periods of high network traffic and is significant enough that it is clear and obviou | Gendex sent a recall letter to their accounts on February 4, 2011 via first class mail, stating that an anomaly in the Orthoralix firmware can cause a message corruption which will distort the patient image being transferred from the X-Ray system to the PC. This image distortion/corruption (distorted, fragmented, flipped, negative, reversed or had lines across them) is most likely to occur during periods of high network traffic and is significant enough that it is clear and obvious these images cannot be used for diagnostic purposes. The letter stated that resending the image from the X-Ray system to the PC will correct the issue present in certain images sent during periods of high network traffic. Instructions were provided to download an Appendix to Orthoralix DDE User Manual that explains how to resend the images. If that does not resolve the issue, users were requested to contact Gendex Technical Support at 1-800-323-8029. Customers were asked to complete and return the enclosed Acknowledgement Form by fax to 215-997-5665 or e-mail: regulatoryaffairs@gendex.com. For questions regarding this recall call 800-323-8029.
Z-1250-2011 | Feb-15-2011 | 2011 | Xeleris Processing and Review Workstation 1.0; 1.1; 2.0; 3.0 - Xeleris Processing and Review Workstation 1.0; 1.1; 2.0; 3.0The display; processing; archiving; and communication of data acquired by Emission Tomography cameras used in diagnostic radiology; including procedures for planar imaging; whole body imaging; tomographic (SPECT) imaging; positron imaging by coincidence; attenuation correction; and anatomical image registration. 3X The system is intended for use by Nuclear Medicine (NM) or Radiology practitioners and referring physicians for display; processing ; archiving; printing; reporting and networking of NM data including planar scans (Static; Whole Body; Dynamic; Multi-Gated) and tomnographic scans (SPECT; Gated SPECT; dedicated PET or Camera-Based-PET) acquired by gammna cameras or PET scanners. The system can run on dedicated workstation or in a server-client configuration. The NM or PET data can be coupled with registered and/or fused CT or MR scans and with physiological signals; in order to: depict; localize; and/or quantify the distribution of radionuclide tracers and anatomical structures in scanned body tissue for clinical diagnostic purposes.. | GE Healthcare; LLC | 2 | GE Healthcare has become aware of safety issues relevant for the Xeleris Workstation 1.0; 1.1; 2.0; 2.1; 3.0:1. Measurements taken on the CT Images on the CT system are different from the Xeleris Workstation2. Inaccurate PET SUV values may be calculated when data is retrieved from PACS after being processed in the Xeleris.3. Wrong marking of Early and Late result series on display in Bra | The firm, GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated January 12, 2011 to all consignees/customers. The letter described the product, problem and actions to be taken by the customers. The customers were instructed to please ensure that all potential users in their facility are made aware of this safety notification and the recommended actions. A GE Healthcare Service representative will perform the required software update on each affected system. If you have any questions or concerns regarding this notification, contact your local GE Healthcare Service Representative. You may also contact GE Healthcare call center in US or Canada at 800-437-1171 and other countries at 262-896-2890.
Z-2220-2011 | May-19-2011 | 2011 | Eclipse Treatment Planning System - Eclipse Treatment Planning System; Client software versions 8.0 through 8.6; Model Number: H48; Product is manufactured and distributed by Varian Medical Systems Inc.; Palo Alto; CA The Eclipse Treatment Planning System (Eclipse TPS) is used to plan radiotherapy treatments for patients with malignant or benign diseases. Eclipse TPS is used to plan external beam irradiation with photon; electron and proton beams; as well as for internal irradiation (brachytherapy) treatments. In addition; the Eclipse Proton Eye algorithm is specifically indicated for planning proton treatment of neoplasms of the eye. | Varian Medical Systems; Inc. Oncology Systems | 2 | An anomaly in Eclipse Treatment Planning System; versions 8.0 to 8.6; whereby Actual SSD (source surface distance) value may not correspond to the Planned SSD value after Planning Approval. | The firm, Varian Medical, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated April 4, 2011, to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to do the following: 1) Set the Planning Approved status in Eclipse (Eclipse displays both the Actual and Planned SSD values during plan approval) and verify both SSD values prior to approving the plan. 2)Verify the SSD values in the Eclipse Treatment Plan Report prior to approving the plan for treatment or proceeding with treatment delivery (this report lists both the Actual and Planned SSD values). 3) Prior to approving the plan in RT Chart, verify both the Actual and Planned SSD values in the Parameters Workspace by selecting Field Parameters>Geometry tab. 4) Prior to treatment, they should verify intended SSD values during independent MU calculation. 5)Re-familiarize themselves with the concepts of Actual SSD and Planned SSD by reviewing the listed information. Please advise the appropriate personnel working in your radiotherapy department of the content of this letter. Varian mitigated the product behavior in version 8.9. If you require further clarification, please feel free to contact the Varian Oncology Help Desk ; phone: USA and Canada - 1888-827-4265; Europe: +41 41 749 8844; email: North America: support-americas@varian.com; Australia/New Zealand: support-anz@varian.com; Europe: support-emea@varian.com; South East Asia: seasia.apps.helpdesk@varian.com; China/Asia: china.apps.helpdesk@varian.com; Japan: Japan.Apps.Helpdesk@varian.com; Latin America: soporte.al@varian.com; Internet: Oncology Systems-www.myvarian.com; Varian Medical Systems -www.varian.com; or your local Varian Customer Support District or Regional Manager.
Z-2702-2011 | Jun-29-2011 | 2011 | Eclipse Radiation Treatment Planning System - Eclipse Radiation Treatment Planning System; model H48;Varian Medical Systems Inc.. | Varian Medical Systems; Inc. Oncology Systems | 2 | In affected Analytical Anisotropy Algorithm versions when modeling the second source; the collimator jaw positions are read from the first control point of the field. This can cause the delivered dose to be higher than planned locally by up to 10%. | Varian Medical Systems, Inc. Oncology Systems sent a "Urgent Medical Device Correction" letters on May 3, 2011 The letter contained recommended actions. For additional information contact Varian Medical Systems, Inc at 888- 827-4265.
Z-1845-2011 | Mar-30-2011 | 2011 | CellQuest Pro SW - CellQuest Pro SW; version 4.0.2. and later;Becton Dickinson; San Jose; CA 95131CellQuest Pro software allows you to acquire and analyze data from your flow cytometer on a Macintosh computer. | BD Biosciences; Systems & Reagents | 2 | Results and gates could be mismatched if you utilize the undo function to undo a gate that was deleted. | Important Product Information letters, dated September 2007, were sent the week of 9/14/2001 via Federal Express to all affected customers. The letters identified the affected product along with the reason for recall. The letter states that "as long as customers do not use the Undo function to undo a gate deletion, the software can be used with full functionality." The letter also states that BD will change the issue in the next software version. If customers believe that they may have analyzed sample files by utilizing the Undo function, then it is recommended that the sample files be re-analyzed. If customers need more information they should contact the Customer Support Center at 877-232-8995.
Z-0932-2011 | Jan-21-2011 | 2011 | 4D Integrated Treatment Console (4DITC) - 4D Integrated Treatment Console (4DITC);All versions up to and including 8.8.17Varian Medial Systems; Palo Alto; CADesigned to assist the operator of a radiation device in providing accurate treatment set-ups for each patient by monitoring set up parameters and preventing the radiation therapy device from commencing irradiation while any parameter is out of conformance with the treatment plan. | Varian Medical Systems; Inc. Oncology Systems | 2 | Multiple isocenter plans will load on the 4DITC and are available for imaging and treatment; but the 4DITC software does not recognize that mulitplex isocenter is present. | Urgent Medical Device Correction/Field Safety notification letters were distributed to all affected users, and to Varian Sales, Marketing and Service Organizations on December 14, 2010 by certified mail. The letters described the issue and recommended user actions along with Varian's actions. Users are to use caution when acquiring actual couch positions for any plan. In addition, each plan should be reviewed for quality assurance purposes in accordance to set departmental or organizational policies. The firm updated the procedure and enclosed a copy with the letter. It was also posted to the customer support site. If further clarification is needed, customers are to contact their local Varian Customer Support District or Regional Manager.
Z-1929-2011 | Apr-08-2011 | 2011 | GE CARESCAPE Monitor B850 - GEThe CARESCAPE Monitor B850 includes both new and existing subsystems interconnected to form a high acuity; multi-parameter patient monitoring system. A typical configuration would be a CARESCAPE Monitor B850 host processing unit running the CARESCAPE Monitoring platform software; a display with integrated keyboard and a frame for the insertion of parameter measurement modules. A variety of options are available to the customer including additional displays; various input devices (keyboard; mouse; bar code reader); additional modules and frames and software options. The new subsystems include a host processing unit (CPU); a 15 inch LCD display with an integrated keypad; 19 inch LCD display with an integrated keypad and touch panel interface; a five slot parameter module frame; a seven slot parameter module frame; a cabled remote control; a cabled remote keypad; and the CARESCAPE Monitoring platform software. Some of these new major subsystems include non-patient contact accessory items (e.g. cables and mounting hardware). The CARESCAPE Monitor B850 interfaces to a variety of existing physiological parameter measurement modules. In addition; the CARESCAPE Monitor B850 interfaces to a variety of existing OEM medical devices via the existing network infrastructure. The CARESCAPE Monitor B850 is a multi-parameter high acuity patient monitor intended for use in multiple areas within a professional healthcare facility. The CARESCAPE Monitor B850 is intended for use on adult; pediatric; and neonatal patients and on one patient at a time. The CARESCAPE Monitor B850 system is indicated for monitoring of Hemodynamic (including ECG; ST Segment; Arrhythmia Detection; ECG Diagnostic Analysis and Measurement; Invasive Pressure; Noninvasive Blood Pressure; Pulse Oximetry; Cardiac Output; Temperature; Impedance Respiration and SpO2 (Mixed Venous Oxygen Saturation)); Airway Gases (Fi/Et C02; 02; N20 and Anesthetic Agent); Spirometry; Gas Exchange (02 Consumption (VO2); C02 production (VCO2); energy expenditure (BE); and respiratory quotient (RQ)) and neurophysiological (including electroencephalography (EEG); Entropy; Bispectral Index (BIS) and Neuromuscular Transmission (NMT) Monitoring) status. The CARESCAPE Monitor B850 provides alarms; trends; snapshots and events; and calculations and can be connected to displays; printers and recording devices. The CARESCAPE Monitor B850 can be a stand-alone monitor or interfaced to other devices. It can also be connected to other monitors for bed to bed viewing and to data management software devices via a network. The CARESCAPE Monitor B850 is intended for use under the direct supervision of a licensed healthcare practitioner; or by personnel trained in proper use of the equipment in a professional healthcare facility. In addition to the healthcare practitioner; the CARESCAPE Monitor B850 is designed to provide configuration and troubleshooting capabilities to qualified service personnel. The CARESCAPE Monitor B850 is not intended for use during MRI. | GE Healthcare; LLC | 2 | GE Healthcare has become aware of potential safety issues associated with the CARESCAPE" Monitor B850 when Patient Data Module (PDM); Tram Acquisition Module (Tram); Patient Side Module (E-PSM or E-PSMP); E- MASIMO or E-(P)RE(S)TN modules or Bed-to-bed alarming feature is used.1. When CARESCAPE Monitor B850 is used with the above listed modules; some SpO2 alarms may be delayed or missed.2. Whe | The firm, GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated February 25, 2011 to its consignees/customers. The letter was addressed to Healthcare Administrator / Risk Manager, Chief of Nursing, and Director of Biomedical Engineering. The letter described the recall product the problem, Safety Issues, Affected Product details, Product Correction, Safety Instructions and Contact Information. The customers were instructed to follow the safety instructions provided in the letter. GE Healthcare is working on a software correction that will be provided to their customers once it is available. If you have any questions regarding this medical device correction or the identification or affected items please contact Technical Support at 1-800-558-7044, or your local Service Representative.
Z-2678-2011 | Jun-28-2011 | 2011 | CDI 500 - Terumo Cardiovascular Systems; CDI blood parameter monitoring system 500 with Arterial and Venous blood parameter modules and Hematocrit/Saturation probe; Catalog AVHCT. | Terumo Cardiovascular Systems Corporation | 2 | Inaccurate readings may result after Methylene Blue; an intravascular dye; is administered to the patient. The presence of Methylene Blue may interfere with the readings from the shunt sensor for the CDI 500 system in a way that would cause inaccurate pH values. Because the pH value is used to calculate other values; these parameters could also be affected by the presence of Methylene Blue: K | Terumo sent URGENT MEDICAL DEVICE CORRECTION letters to Chief of Perfusion, Department of Cardiothoracic Surgery or Director of Operating Room Services, dated 4/21/2011. The letter explained the reason for the correction and the potential hazard. The letter stated that Terumo was updating the Operators Manual for all CDI systems and would provide the manuals to all users when available. Consignees with questions should call 1-800-521-2818. Monday- Friday 8 am- 6 pm.
Z-0891-2011 | Jan-10-2011 | 2011 | 4D Integrated Treatment Console - 4D Integrated Treatment Console; versions: 8.0.12; 8.0.23 SP2; 8.1 SP2; 8.1.1.3; 8.1.1.4; 8.1.2; 8.1.3; 8.1.4; 8.3; 8.6; 8.6.1; 8.6.2; 8.8.0; Model Number: H51; Distributed and manufactured by: Varian Medical Systems Inc.; Palo Alto; CA Intended use: The 4D Integrated Treatment Console function is designed to assist the operator of radiation therapy device in providing accurate treatment setups for each patient by monitoring setup parameters and preventing the radiation therapy device from commencing irradiation when any parameter is out of conformance with the treatment plan. | Varian Medical Systems Oncology Systems | 2 | Software anomaly may result in a design deficiency causing Plan Edit Settings to malfunction under specific circumstances; where the user selects to use a Cone Beam CT set up field. | The firm, Varian Medical Systems, sent a Product Notification Letter/Field Safety Letter entitled "URGENT MEDICAL DEVICE CORRECTION" dated October 16, 2009, to all affected customers. The letter described the product, problem, actions to be taken by the customers and Varian. The customers were instructed to use caution when creating CBCT imaging field and limit use of the 'select all' button when acquiring treatment parameters; review departmental policies and procedures regarding the acquisition of any gantry angle at the 4D ITC, and advise the appropriate personnel of this letter. The firm, Varian, recommended that only vertical, longitudinal and lateral couch parameters are acquired when creating a CBCT field at the 4D ITC; notified all affected customers of this document, and will correct this issue in a future release of the 4D ITC software. If owners of $D ITC, version 8.0.12 and higher, have any questions regarding this product notification letter, please contact Varian Medical Systems Help Desk at 1-888-VARIANS (888-827-4265) for USA and Canada, +41 41 749 8844 or support-emea@varian.com for Europe; Emails for North American: support-americas@varian.com; Australia/New Zealand: support-anz@varian.com; South East Asia: seasia.apps.helpdesk@varian.com; China/Asia: china.apps.helpdesk@varian.com; Japan: Japan.Apps.Helpdesk@varian.com and Latin American: soporte.al@varian.com, and Internets for Oncology Systems- www.myvarian.com and Varian Medical Systems- www.varian.com.
Z-1853-2011 | Mar-30-2011 | 2011 | Acuson Antares Utrasound imaging device - Acuson Antares Ultrasound Imaging System with software versions 4.0 and 5.0. | Siemens Medical Solutions USA; Inc. | 2 | Image captured from the device may be erroneous. | Acuson Customer Safety Advisory notification letters were sent March 1, 2010 by certified mail.
Z-3178-2011 | Sep-12-2011 | 2011 | EP-Shuttle/Stockert J50/Stockert J70 - STOCKERT 70 Radio Frequency Generator; Catalog Number 57001 Indicated for use with a Biosense Webster; Inc. diagnostic/ablation deflectable tip catheter to deliver radio frequency energy during cardiac ablation procedures.Stockert 70 in foreign countries is also known as: EP-Shuttle/Stockert J50/Stockert J70With software versions: 1.035/1.035J/1.035J70 to 1.037/1.037J70 | Biosense Webster; Inc. | 2 | Biosense Webster has initiated a voluntary field notification to all customers that have a Stockert 70 Radio Frequency Generator; to communicate undetected software anomalies in the form of software release note in order to amend the labeling and thereby enhance the safety of the product. | The recall communication was initiated on 7/8/11 and 7/11/11 with Biosense Webster forwarding A Customer Notification Letter, Acknowledgement Form and Software Release Note (Attachments 3, 4 and 1) via FedEx to all U.S. customers who purchased the Stockert 70 Radio Frequency Generator. Two letters and forms were sent to U.S. customers via express mail, addressed to the EP/Cardiology Lab and Risk Management Office, respectively. The Release Note provides tips for the avoidance of potential events when working with the EP-SH UTTLE, STOCKERT 70, STOCKERT J50 or STOCKERT J70 RF generators with SW versions 1.035 / 1.035J /1.035J70 to 1.037/1.037J70.
Z-2403-2011 | Jun-01-2011 | 2011 | Capsule Neuron UMPC - Capsule Neuron UMPC is a component in the DataCaptor" Connectivity System. Catalog Number: DC-NU-MPCCapsule Neuron" UMPC is a bedside device that enables automatic collection of vital signs data that can local cache data and connect to the hospital network through a wiredEthernet connection or via a standard 802.11 b/g/n wireless network; available in a fixed or mobile mounting option; running either DataCaptor continuous or periodic connectivity softwareThe DataCaptor" System is indicated for use in data collection and clinical information management either directly or through networks with independent bedside devices. DataCaptor" is not intended for monitoring purposes; nor is the software intended to control any of the clinical devices (independent bedside devices / information systems) it is connected to | Capsule Tech Inc. | 2 | Software: Data cached on the Neuron due to server unavailability is not being delivered to the server after server availability is restored.If a compact flash failure occurs; the file system may be corrupted and result in system hangs; lock-ups or crashes.Batteries can discharge completely in a Capsule Neuron" UMPC even though it is connected to power source | Capsule Tech initiated three corrective action notifications for the Neuron" UMPC units three failure modes as follows: a) Data cached on the Capsule Neuron" UMPC due to server unavailability is not being delivered to the server after server availability is restored: i) A customer advisory was provided to a subset of consignees, along with DataCaptor software version 6.6.5 on December 1, 2010. ii) A customer advisory was provided to the remaining consignees, along with DataCaptor software version 6.6.7 on January 27, 2011. If a compact flash failure occurs, the file system may be corrupted and result in system hangs, lock-ups or crashes: i) A customer advisory was provided to all consignees on March 23, 2011, via phone call, email and FedEx;.The correction will be provided as a software update. A formal recall notification letter will be provided to all consignees alerting them to the availability of the correction once it is released, via phone call, email and FedEx. c) Batteries can discharge completely in a Capsule Neuron" UMPC even though it is connected to power. Note: the Capsule Neuron" UMPC operates normally when connected to power. i) A customer advisory was provided to all consignees on April 23, 2010, via FedEx. For questions regarding this recall contact Capsule technical support team at (800) 260-9537 or e-mail support@capsuletech.com..
Z-0867-2011 | Jan-07-2011 | 2011 | Stryker Digital Capture Pro 2 - Stryker Digital Capture Pro 2 device (240-050-800; CD version) and (240-050-810; DVD version) designed for digital documentation of medical procedures. Manufactured by Stryker Endoscopy; San Jose; CA. | Stryker Endoscopy | 2 | Device will not able to assign the correct time/date stamp for saved filed after December 31; 2009. On Jan 1; 2010; the device system date for all SDC Pro 2 units will re-start from 1 Oct 2001 and begin incrementing from that point onwards. The system date will be incorrect and users will not be able to correct this issue. Problem will not affect how pictures or videos are captured using SDC Pro | Stryker Endocoscopy sent an URGENT: Device Correction letter dated December 11, 2009, to the OR Supervisor/Risk Manager and all Stryker Sales Representatives. The letter identified the product, the problem, and the action to be taken by the Sales Representatives. Sales Representatives were to identify affected accounts. Go to customer accounts and install the upgrade for each affected unit. Upon completion of the upgrade, the sales rep was to sign the acknowledgment form indicating that all of their affected accounts received the appropriate software fix and mail to Stryker Endoscopy or fax both sides of the postcard to 408-754-8378 or scan in both sides of the postcard and email to sdcpro2recall@stryker.com. For questions regarding this recall call 408-754-2738 or 408-754-2000.
Z-3032-2011 | Aug-17-2011 | 2011 | Spacelabs Medical Patient Monitor - Spacelabs Medical Patient Monitor is a Physiological; Patient monitor With Arrhythmia Detection or Alarms. The Models are 91367; 91369; 91370; 91387; & 91393 are sold with with Perioperative Mode (option D); which includes Start case/End case functionality.Spacelabs Healthcare patient monitors; functioning as either bedside or central monitors; passively display data generated by Spacelabs Healthcare parameter modules; Flexports interfaces; and other SDLC based products in the form of waveform and numeric displays; trends and alarms. Key monitored parameters available on the models 91367; 91369; 91370 and 91387; when employing the Spacelabs Command Module; consist of ECG; respiration; invasive and noninvasive blood pressure; Sp02; temperature and cardiac output. Additional parameters and interfaces to other systems are also available depending on the parameter modules employed.Product Usage: When a clinician enters END CASE mode; there is the expectation the operation is over and the patient is being disconnected from the monitor. Removing the patient cables will trigger multiple alarms so clinicians will enter this END CASE mode to prevent having to listen to false alarms. Spacelabs Healthcare patient monitors are intended to alert the user to alarm conditions that are reported by Spacelabs Healthcare parameter modules and/or other physiologic monitors via Flexport interfaces. These devices determine: a) when an alarm condition is violated; b) the alarm priority (ie. high; medium or low); c) alarm limits; and d) when to initiate and terminate alarm notifications. The patient monitors are also capable of displaying alarm conditions on other monitors that are on the network through the Alarm Watch feature. Spacelabs Healthcare patient monitors may also function as a generic display or computer terminal. As a generic display or terminal; the patient monitors allow network-based applications to open Windows and display information on other net worked monitors. Spacelabs Healthcare patient monitors are also designed to communicate with a variety of external devices such as displays; network devices; serial devices; user input devices; audio systems; and local/remote recorders. Spacelabs Healthcare patient monitors are intended for use under the direct supervision of a licensed healthcare practitioner; or by personnel trained in proper use of the equipment in a hospital environment. | Spacelabs Healthcare; Llc | 2 | The clinician did not realize the monitor was in END CASE mode and was expecting audible alarms. Visual alarms continue to work normally. | Spacelabs sent an "URGENT - MEDICAL DEVICE CORRECTION" letter dated July 25, 2011 along with the Perioperative section of the Operations Manual to all consignees in the US. This letter includes recalled model 91369, 91370, 91387, & 91393 (option D). On July 27, 2011, Spacelabs sent out an updated recall notification letter to include one additional model 91367 to all consignees in the US. The letter identifies the products, problem, and actions to be taken. On August 2, 2011, a recall notification letter (translated as necessary) and the Perioperative section of the Operations Manual was emailed to all international subsidiaries and distributors of record. The notification letter stated that the Operating Room staff may not be aware that alarm tones are off if the device has been left in the END CASE state. If the customers make use of the Perioperative START CASE / END CASE feature, they should advise the clinicians to familiarize themselves with this feature as described in the Perioperative section of the Operations Manual. In particular, they should know how to confirm the state of the alarm tones prior to starting a case. The letter states that the firm is in the process of upgrading the software and will be contacting customers to update monitors at no cost. If customers have any questions about this corrective action program, please contact Spacelabs Healthcare at 1-800-522-7025 and select 2 for Technical Support, or outside the United States 1-425-657-7200 x5089.
Z-1961-2011 | Apr-12-2011 | 2011 | ACUSON Sequoia ultra sound systems with Auxiliary 2 MHz Continuous Wve (CW) transducer - Auxiliary 2 MHz Continuous Wave (CW) transducer; model number 8251046; when used with ACUSON Sequoia ultrasound systems at SW revision 12.X. Manufactured by Siemens Medical Solutions USA; Inc.; Ultrasound Business Unit; Mountain View; CA.The ACUSON Sequoia is a general purpose Diagnostic Ultrasound system. It is intended for the following applications: Abdominal; General Imaging and Cardiology for Fetal; Abdominal; Intraoperative (abdominal and neurological); Pediatrics; Small Organs (breast; testes; thyroid; and penis); Neonatal/Adult Transvaginal; Peripheral Vessels and Musculo-skeletal (superficial and conventional). The system also provides for the measurement of anatomical structures that provide information that is used for clinical diagnosis purposes. The Siemens Aux CW transducer is a non-imaging continuous wave transducer indicated for use in transthoracic/cardiac adult and pediatric applications as well as continuous wave spectral Doppler echocardiography. | Siemens Medical Solutions USA; Inc. | 2 | Overheating. A software bug may cause the transducer to overheat when used with certain transducer models. The transducer can become noticeably hot within a few minutes; which may cause patient discomfort or burns. | The firm, SIEMENS, sent a "Customer Safety Advisory Notification" to its customers on 3/21/2009. The notification described the product, problem and actions to be taken. The customers were instructed to disconnect any transducers that may be in an active transducer port and chose another imaging transducer prior to use of the Aux CW transducer. If they do not have another transducer to select other than those listed in the notification, they should close the study in progress and start a new study before using the Aux CW transducer-this procedure should be followed for each new patient and /or exam type. In addition, the customers were instructed to pass this notice on to all those within their organization who need to be aware of this issue until the corrective action is completed. SIEMENS will release a software update that will correct the issue. The customers local service engineer will install the updates on their system as soon as it becomes available. Should you have any questions regarding this notification, you may contact your local service support person and/or call (650) 969-9112.
Z-3002-2011 | Aug-11-2011 | 2011 | ECLIPSE TREATMENT PLANNING SYSTEM - Eclipse Treatment Planning System;Varian Medical Systems; Palo Alto; CA Product Usage: Eclipse Treatment Planning System is used to plan radiotherapy treatments for patients with malignant or benign diseases. It is used to plan external beam irradiation with photon; electron and proton beams; as well as for internal irradiation (Brachytherapy) treatments. | Varian Medical Systems; Inc. Oncology Systems | 2 | In affected Eclipse versions; the treatment field energy may change from the current selection to a previous selection without use intervention and without invalidating the MU or calculated dose distribution. | Varian sent an "Urgent Medical Device Correction/Urgent Field Safety Notice" letter dated June 30, 2011 to all affected consignees. The letter identifies the affected product, problem, and actions to be taken. The letter states that Varian will contact all affected customers to apply a database configuration script to the database and is preparing a correction for Eclipse Version 10 to address the issue. A Varian representative will contact each site to schedule correction when it is available. The letter also provides RECOMMENDED USER ACTIONS to follow. The Letters go on to reiterate caution and warnings from Eclipse instructions for use. All appropriate personnel should be notified of the correction notice. For future reference, this document will be posted to the Varian customer support website: http://www.MyVarian.com. For further clarification, contact your local Varian Customer Support District or Regional Manager.
Z-1609-2011 | Mar-08-2011 | 2011 | Phiips Medical Systems - Philips HeartStart MRx Defibrillator Monitor (M3535A/M3536A) with software (SW) versions F.01.00 or F.01.01 and with all of the following hardware options EtCO2; IP and Temp.Used for the termination of ventricular tachycardia and ventricular fibrillation. | Philips Healthcare Inc. | 2 | Software: EtCO2 and Ventilation Rate values may be labeled incorrectly. | Philips Healthcare representatives issued an Urgent Medical Device letter on 2/7/11. Customers are asked to follow the Action to be taken by Customer/User section of the Urgent Medical Device Correction Notification. In this section, customers are given the following instructions: Philips software (SW) version F.01.03 (for M3535A and M3536A) and R.01.00 (for M3536A, English only) corrects this problem. If you have affected product a Philips Service Engineer will contact you to install software version F.01.03 or R.01.00. Until your software is updated your HeartStart MRx monitor/defibrillator can remain in service. If you are using the device in one of the use modes listed under problem description, make sure that operators are aware that even though the EtCO2 and ventilation rate values can display interchangeably (unexpectedly or randomly switch) under either the ventilation rate label or the EtCO2 label, the values themselves can be identified by their color (for example if the ventilation rate label is white, the ventilation value is displayed as white). Customers are to contact their local Philips representative at 1-800-722-9377, if they need further information or support regarding this issue.
Z-2065-2011 | Apr-25-2011 | 2011 | Acuson X300 - Acuson X300 systems at software versions 1.00.00 to 1.0.10 and 2.0.00 to 2.0.02; model numbers 10037409; 10132987; and 10038837; manufactured by Siemens Medical Solutions ultrasound business unit; Mountain View; CAMedical device is an ultrasound imaging system for general radiology; fetal; abdominal; intraoperative; pediatric; small parts; neonatal/adult cephalic; cardiac; transesophageal; pelvic; transcranial; ob/gyn; urology; vascular; musculoskeletal; superficial musculoskeletal and peripheral vascular applications. The system also provides for measurements of anatomical structures and for analysis packages that provide information that is used for clinical diagnosis purposes. | Siemens Medical Solutions USA; Inc. | 2 | Software Error with potential for misinterpretation of data-- The previous patient's last capture clip image can appear in the next patient's study. | The firm, Siemens, sent a "CUSTOMER SAFETY ADVISORY NOTIFICATION" letter to its customers on July 30, 2009. The letter described the product, problem and actions to be taken. The customers were instructed they need to update any of the patient registration settings to reset the cine buffer so that no data is transferred between studies, if the system software is not being updated. Siemens informed the customers that a mandatory update software is being installed on their system free of charge. If you have any questions, contact (650) 969-9112.
Z-3101-2011 | Aug-31-2011 | 2011 | On-Board Imager (OBI); Trilogy Mx - On-Board Imager (OBI) 1.3; 1.4; 1.5; and Trilogy Mx; TrueBeam 1.0; Offline Review 1.0 - 2.0;Product is manufactured and distributed by Varian Medical Systems Inc.; Palo Alto; CAThe On-Board Imager device is used for verification of correct patient position in relation to isocenter and verification of the treatment fields relation to anatomical and/or fiducial landmarks. Trilogy Mx is intended to provide stereotactic radiosurgery and precision radiotherapy for lesions; tumors; and conditions anywhere in the body where radiation treatment is indicated. | Varian Medical Systems; Inc. Oncology Systems | 2 | Varian has identified a software anomaly in the 3D Match environment; regarding a slight mismatch between CT image and Structures in On-Board Imager w/s. | Varian Medical Systems sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated June 21, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. The letter recommends several user actions. The firm is developing a correction for the issue and will contact customers when the correction is available. Questions regarding this notice may be addressed by contacting the Varian Oncology Help Desk at 1-888-827-4265.
Z-0114-2012 | Oct-26-2011 | 2011 | GE SENOGRAPHE ESSENTIAL FULL FIELD MAMMOGRAPHY SYSTEM - GE SENOGRAPHE ESSENTIAL FULL FIELD MAMMOGRAPHY SYSTEM (SENO ESSENTIAL).The Senographe Essential system is intended to be used in the same clinical applications as traditional mammographic film/screen systems. It generates digital mammographic images which can be used for screening and diagnosis of breast cancer. | GE Healthcare; LLC | 2 | GE Healthcare has become aware of inaccurate measurement on GE magnified images provided by Senographe DS and Essential Mammography systems when reviewed on IHE compliant workstations. This issue may impact patient safety. | GE Healthcare sent an " URGENT MEDICAL DEVICE CORRECTION" letter dated September 21, 2011 to all affected customers. The letter describes the products, problem, and actions to be taken by the customers. The letter provides safety instructions and advises customers to complete and return a Customer Reply Form via fax to +33(0)1 30 70 39 30. A GE Healthcare service representative will contact customers if the updated software version is required. Contact the Call Center at 800-437-1171.
Z-1819-2011 | Mar-24-2011 | 2011 | Extended Brilliance Workstation (EBW) - Extended Brilliance Workstation (EBW) a component of the Computed Tomography X-Ray System Model Numbers EBW (CT) 728260 & EBW (Portal) 728269.The Extended Brilliance Workplace¿¿ (EBW) is a component of the Computed Tomography X-Ray System that is intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. This device may include signal analysis and display equipment; patient; and equipment supports; components and accessories. The EBW is also provided as a component of PET/CT and SPECT/CT systems; performing the same function as indicated above (also referred to as the NM EBW). | Philips Medical Systems (Cleveland) Inc | 2 | During an FDA inspection on 10/1/2010 the firm was notified by FDA that their field action they initiated in Oct 2009; constituted a recall. The firm determined that the functional analysis feature in the Comprehensive Cardiac Analysis (CCA) was displaying incorrect values on the polar fields. The calculated results for the myocardial wall thickening was incorrect in that the numbers displayed f | Philips sent an URGENT - Field Safety Notice letter dated September 29, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to be aware that the numbers displayed for "Wall Thickening" actually represent a reduction in wall thickness from ES to ED. Philips will be providing a sotware update to resolve the issue. For further information or support customers were instructed to contact their local Philips representative: Customer Care Service Center in the U.S. at 1-800-722-9377, option 5: Diagnostic Imaging option 1: CT or their local Philips representative.
Z-0905-2011 | Mar-04-2011 | 2011 | BuckyDiagnost with Optimus Generator - BuckyDiagnost general purpose x-ray system with Optimus Generator.General Purpose Diagnostic Radiography. | Philips Healthcare Inc. | 2 | Affected systems were determined not to prevent run-on of a selected technique factor if the technique selector and the exposure control were activated simultaneously. | A Customer Information BuckyDiagnost Diagnostic X-Ray System letter, dated January 6, 2011, was sent to customers. The letter identifies the affected product, the reason for the recall, actions that are to be taken by customers, and actions that are planned by the firm. Customers are advised to make sure that the desired exposure values are displayed and that they have released the technique button BEFORE they start the exposure by pressing the exposure switch whenever manually changing the exposure parameters and in order to avoid the problem. Philips will issue a Field Change Order (FCO) to resolve the issue by updating the system software free of charge. Philips will contact customers to implement the FCO. Further information or support concerning this issue can be obtained by contacting your local Philips representative. The Field Service Engineers began software implementation the week of March 21, 2011.
Z-2051-2011 | Apr-22-2011 | 2011 | Acuson S2000 ultrasound system - ACUSON S2000 ultrasound systemManufactured by Siemens Medical Solutions USA; Inc.Business Unit Ultrasound;1230 Shorebird WayIntended for the following applications: Fetal; abdominal; Intraoperative; Pediatric; Small Parts; Transcranial; OB/GYN; Cardiac; Pelvic; Neonatal/Adult Cephalic; Vascular; Musculoskeletal; Superficial Musculoskeletal and Peripheral Vascular Applications. | Siemens Medical Solutions USA; Inc. | 2 | The Firm discovered a software error on the Accuson S2000 ultrasound systems affecting the measurement results in the patient report. | Siemens issued a Customer Safety Advisory Notification letter to all customers on March 15, 2011, regarding a software error on the ACUSON S2000 ultrasound systems affected measurement . They are in the process of releasing a software update that will correct this issue. Siemens recommends reviewing the patient report at the end of each study to confirm that the report contains only measurements relevant to the exam just performed. No patient injury has been reported. Further questions regarding this recall please call (650) 694-5398.
Z-2726-2011 | Jul-14-2011 | 2011 | IQecg - IQecg IQmanager Diagnostic Workstation; Part Numbers: 4-100-1200; Software Version 8.0.1; 8.3; 8.3.1; 8.3.2IQecg is to be used; under the supervision of a Physician; to obtain Electrocardiograms from the human body surface. | Brentwood Medical Technology Corp | 2 | The recall was initiated because when a user of the IQmanager; IQdevices SDK; and IQiC (Centricity) software versions listed above accesses the QT dialogue box to either accept or edit the QT value; the QTc value is recalculated using a Hodges QTc equation; regardless of patient's age. The software does this without notifying the user of the change from use of the Bazett equation; which is utiliz | The firm, Midmark, sent an "URGENT: MEDICAL DEVICE FIELD CORRECTIVE ACTION" letter on May 19, 2010 with an attached Corrective Return Material Authorization (RMA) form via US Postal Service Certified Mail to its consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately examine their IQecg, IQmanager, or lQiC Centricity software to determine the software version currently in use. If it is a listed product, it is subject to this field corrective action. A sample software label was attached to aid the consignees/customers in locating and identifying the software CD. For customers who have already installed the software, the version number can be identified in the "About" page. Consignees/customers were instructed to complete the enclosed Field RMA form and return it with the affected software to: Quality Assurance Department Midmark Diagnostics Group 1125 West 190 S t. Gardena, CA 90248 In addition, if they further distributed this product, identify customers and notify them at once of this product field corrective action. Replacement software was to be shipped to the consignees/customers upon receipt of the returned software CD. If they require a UPS return shipping label or an on-line download of replacement software, contact Midmark Diagnostics Group Support Services at 800-624-8950 extension 5145 or 5108. If you have any questions, call Midmark Diagnostics Group Support Services at 800-624-8950 extension 5145 or QA/RA Manager, Midmark Diagnostics Group 310-516-5116.
Z-2976-2011 | Aug-10-2011 | 2011 | Philips Medical Allura Xper Systems - Allura Xper Systems Models:Allura Xper FD10 (F/C)-Biplane Fixed Allura Xper FD10/10 (incl. OR table)Allura Xper FD20- High resolution fixed-x rayAllura Xper FD20/10Allura Xper FD20/20Allura CV20The Allura Xper F010; FD10/10; FD20; FD20/10 and FD20/20 Use From Labeling): is intended for use in cardiovascular and vascular X-ray imaging applications. including diagnostic;interventional procedures (such as PTCA; stent placement and atherectomies); pacemaker implantations and ElectroPhysiology (EP). The Allura Xper is intended for: Dedicated vascular and neurovascular imaging applications; including diagnostic and interventional procedures. This includes; e.g. peripheral; cerebral. thoracic and abdominal angiography; as well as PTCAs. stent placements; embolisations and thrombolysis. and Cardiac imaging applications including diagnostics; interventional procedures (such as PTCA; stent placing; atherectomies); pacemaker implantations and Electrophysiology (EP). Non-vascular interventions such as drainages; biopsies and vertebroplasties procedures. Use of the Allura Xper FD Series in the OR environment is restricted; the system should only be used for minimally invasive procedures. (Endovascular) | Philips Healthcare Inc. | 2 | System freezes from 20 seconds up to one minute. No fluoroscopy exposure and no geometry movements during this time period. | Philips Health Care North America sent a FIELD SAFETY CORRECTION ACTION letter dated July 4, 2011, to all affected customers. The letter identified the product, the problem and the action needed to be taken by the customer. This FCD will be identified as FCD72200200 and will be free of charge. This FCD consists a CD with the correct firmware and a installation procedure to implement the FCD correctly. Please contact your local Philips representative. For further follow up please call (978 ) 687-1501.
Z-2578-2011 | Jun-16-2011 | 2011 | XiO Radiation treatment Planning System - XiO Radiation treatment Planning System; XiO Release 4.2.0 and aboveUsage: Used to create treatment plans for any cancer patient for who external beam radiation therapy or brachytherapy has been prescribed. | Computerized Medical Systems Inc | 2 | When a dialog box is displayed and the dose status is changed to "Dose not current"; beam edits can be made in the graphics area using the mouse (rotate collimator; reposition collimator; move isocenter; move weight point; etc.). Then if CANCEL is selected on the dialog box; any edits made on the dialog box are undone; the dose is left unchanged; but the beam edits made in the graphics area with t | The firm, Elekta CMS Software, sent an "IMPORTANT SAFETY NOTICE" dated September 6, 2010 to its affected customers. The notice describes the product, problem and actions to be taken (workaround). The customers were instructed to distribute this notice to any and all users of CMS software at their organization who are potentially affected by this issue. The notice also states that this problem has existed since XiO Release 4.2.0 and will be resolved in XiO Release 4.70.00. Elekta will notify the sites when this solution is available. If you have any questions, please call 314-993-0003 or Toll Free: 800-878-4267.
Z-2595-2011 | Jun-20-2011 | 2011 | XiO Radiation Treatment Planning System - XiO Radiation Treatment Planning System; XiO Release 4.0.2 and aboveUsed to create treatment plans for any cancer patient for who external beam radiation therapy or brachytherapy has been prescribed. | Computerized Medical Systems Inc | 2 | XiO Software: when Digitally Reconstructed Radiographs (DRRs) are generated in XiO; and DICOM is used to export the DRRs (File-Print-DRR); the exported images are associated with the Beams; but NOT the Plan. The Plan association had always been done manually by the user. This is functionality that has never been provided on XiO. The problem occurs because XiO is not storing a persistent DICOM | Elekta / Computerized Medical Systems Inc sent an undated Safety Notice to all affected customers. The notice indicated the problem, the clinical impact, and the workaround. The notice stated the sites will be notified when a solution is available. No date was given for the mailing or the solution.
Z-3103-2011 | Aug-31-2011 | 2011 | BrachyVision Software; part of Eclipse TPS - BrachyVision; versions 8.2; 8.5; 8.6 and 8.9. Varian Medical Systems.3100 Hansen Way; Palo Alto; California 94304-1038.Used to plan radiotherapy treatments for patients with malignant or benign diseases. In addition; the module is specifically indicated for planning proton treatment of neoplasms of the eye. | Varian Medical Systems; Inc. | 2 | Medical device for brachytherapy treatment may cause users to enter incorrect syntax and lead to potential misinterpretation of patient treatment plan. | Varian Medical Systems, Inc. sent a "Medical Device Recall" letter dated August 05, 2010 to all affected customers. The letter included description of product affected and problem. Customers were advised to review template files to ensure that there are no characters immediately after the string in the customized dose quality parameters. Device users were requested to return a response form. For information on this recall contact Varian Technical Support at (800) 360-7909.
Z-2579-2011 | Jun-16-2011 | 2011 | XiO Radiation Treatment Planning System - XiO Radiation Treatment Planning System; XiO Release 4.3.0 and aboveUsed to create treatment plans for any cancer patient for who external beam radiation therapy or brachytherapy has been prescribed. | Computerized Medical Systems Inc | 2 | XiO: When the beam template is recalled; XiO prematurely initiates a dose calculation; re-snaps the MLC or Port to the structure of the current patient ; but then fails to re-calculate dose for the new MLC or Port position. If no subsequent change is made to force a dose recalculation; XiO will display the dose from the original MLC/Port shape dose instead of the dose from the re-snapped MLC/Port | The firm, Elekta CMS Software, sent an "IMPORTANT SAFETY NOTICE" dated September 14, 2010 to its affected customers. The notice describes the product, problem and actions to be taken. The customers were instructed to distribute this notice to any all users of CMS software at their organization who are potentially affected by this issue, and complete and return the attached postcard to confirm receipt of the Safety Notice. The notice also states that this problem has existed since the XiO Release 4.3.0 and is resolved in the XiO release 4.26.00. Elekta will notify the customers when this solution is available. If you have any questions, please call 314-993-0003 or toll free 800-878-4267.
Z-0182-2012 | Nov-10-2011 | 2011 | Extended Brilliance Workstation-NM (EBW-NM) - Extended Brilliance Workstation-NM (EBW-NM) with software version 1.0P; 1.1.1A; 1.5H; 1.5.1A and 2.0Q -Field Test PeriodA nuclear medicine image display and processing application suite that provided software applications used to process analyze and display medical images/data. | Philips Medical Systems | 2 | A problem related to a software error has been detected in the Philips Multiple Gated Radionuclide Angiocardiography (MUGA) application. A false negative interpretation due to a falsely elevated ejection fraction or a false positive interpretation due to a false low ejection fraction could occur. | Philips sent all affected sites an Urgent Medical Device Correction Letter on September 27, 2011 for the MUGA issue. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the information contained in the letter with all members of their staff who need to be aware of the contents of the communication. Customers were asked to retain a copy of the communication with the equipment instruction for use. Additionally, a field change order was released on September 26, 2011, customers will be receiving release notes as part of their correction informing them of the changes and corrections to the software. The new software will be installed on affected units by Philips Field Service Engineers. For customers in North America contact Customer Care Solutions Center at 1-800-722-9377. Select option 5. In all other countries contact the local Philips Healthcare office.
Z-1402-2011 | Feb-25-2011 | 2011 | STRYKER NAVIGATION SYSTEM - iNtellect Cranial Software - Stryker Navigation System - iNtellect Cranial Software; Part number 6000-650-000.It is part of the product series of the Stryker Navigation System. The system comprises software for pre-operative surgical planning and supports computer assisted surgery based on a wireless optical tracking localization device for the use in navigated cranial surgery. It includes automatic multi-modality correlation. The system provides an easy-to-use work flow concept for image importation; FESS; Craniotomy; Biopsy; etc. Additional; automatic segmentation of brain surface; tumor; vessels; etc. are supported. | Stryker Instruments Division of Stryker Corporation | 2 | When users imports more than one image set and changes the viewing direction for both images; the orientation of every other image series might be stored incorrectly; as a mirror image. This may lead to insufficient or inaccurate correlation results or incorrectly shown orientation of every second image series during navigation. | Urgent: Medical Device Recall Notification letters were sent on 12/3/2010, and software upgrades were to be performed in the field by the end of 2010. Specific Customer Instructions: 1) Locate the units listed in the notification. 2) Distribute this notification to all departments affected. 3) Return the enclosed business reply form to confirm receipt of this notification and identify how many affected units are currently in your inventory. 4) Once the new software is available, Stryker will dispatch a service tech in order to update the software. Note: Software upgrades for all applications will be available by the end of the year. The upgrades will be scheduled as we receive your business reply form. Note: If you currently have a Time License or Pay Per Use card, a new one will be issued to you at the time of the update. 5) If you have loaned or sold any of the units listed in this notification, please, forward a copy of this notice to the new users and advise Stryker of the new users name and their new location. 6) If you have disposed of any of the affected units and they are no longer in use, please, advise us of their obsolescence by providing us with their serial numbers. Questions regarding this recall should be directed to Stryker Field Action Team at 800-800-4236, ext 3808, 3584, or 4354 between the hours of 8 a.m. - 5 p.m. Monday - Friday.
Z-2568-2011 | Jun-15-2011 | 2011 | XiO Radiation Treatment Planning System - XiO Radiation Treatment Planning System; XiO Release 4.3.0 through 4.50.00.Used to create treatment plans for any cancer patient for who external beam radiation therapy or brachytherapy has been prescribed. | Computerized Medical Systems Inc | 2 | In XiO; when the beams are renumbered; it is possible to get into a state where the dose displayed for a beam does not match the beam placement on the patient. This can occur when dose is calculated; beams are renumbered; and the dose is re-summed only. | Computerized Medical Systems, Inc. sent an CUSTOMER ADVISORY letter dated May 2009, to all affected customers. The letter identified the product, the problem, and the action needed to be taken by the customer. Computerized Medical Systems, Inc have created "patches" to resolve the problem. A User Notice was sent in March 2011, to all users still using the affected software notifying them of the updated software available to correct the problem. This letter was sent by US Mail but no return receipt was included. For further questions, please call (408) 380-8023.
Z-0124-2012 | Oct-27-2011 | 2011 | SYNCHRON CX4 - SYNCHRON CX4 Clinical System Part #: 758300Product Usage: The Beckman SYNCHRON CX4 DELTA; CX7 DELTA; CX4CE; and CX7 systems are fully automated and computer controlled instruments designed for the in vitro diagnostic quantitation of biological fluid components and therapeutic drugs as well as the qualitative determination of drugs of abuse in urine | Beckman Coulter Inc. | 2 | The recall was initiated because Beckman Coulter has confirmed software limitations which affect only manually-loaded reagent cartridges on the CX4/5/7 CE; CX4/5/7 Delta and CX9 ALX Clinical Systems. | Beckman Coulter sent an "URGENT: PRODUCT CORRECTIVE ACTION" letter dated September 20, 2010 with an attached Fax Back Response Form (via US Postal Service for US Customers) to all customers who purchased the SynchronÂ® CX4/5/7 CE, CX4/5/7 Delta and CX9 ALX Clinical Systems. The letter provides the customers wtih an explanation of the problem identified and actions to be taken. Customers were instructed to share this information with laboratory staff and retain this notification as part of their laboratory Quality System documentation. If they forwarded the affected product listed above to another laboratory, they were asked to provide a copy of this letter to them. Questions regarding this Product Corrective Action letter, were directed to Customer Support Center (Call Center) on the Beckman Coulter website at http://www.beckmancoulter.com/customersupport/support or call (800) 854-3633 in the United States and Canada. Outside the United States and Canada, they were directed to their local Beckman Coulter Representative.
Z-0310-2012 | Nov-30-2011 | 2011 | cobas c 311 Analyzer - cobas c 311 Analyzer; Part Number: 04826876001In-vitro diagnostic analyzer capable of performing clinical chemistry; specific protein and electrolyte tests. Analytes are measured photometrically or turbidimetrically; the analyzer also has an optional ISE module for measuring sodium; potassium and chloride. | Roche Diagnostics Operations; Inc. | 2 | The cobas c 311 analyzer software may incorrectly calculate the volume of the ISE Reference Electrolyte solution remaining on the analyzer. In some cases the incorrect results obtained will not be flagged with any data alarms. | October 26, 2011 Urgent Medical Device Corrections were sent to all direct accounts in the USA that have installed the cobas c 311 analyzer and included a faxback form to be completed and faxed to 1-877-766-7452. Consignees were instructed to temporarily replace the ISE Reference Electrolyte bottle based upon the frequency calculated from the maximum number of ISE samples run per day and activate the yellow threshold alarm on the system. Customers with question and concerns are directed to call Roche Diagnostics Technical Support at 1-800-428-2366, available 24-hours, 7 days a week.
Z-2574-2011 | Jun-15-2011 | 2011 | XiO Radiation Treatment Planning System - XiO Radiation Treatment Planning System; XiO Release 1.5.0 and above.Used to create treatment plans for any cancer patient for who external beam radiation therapy or brachytherapy has been prescribed. | Computerized Medical Systems Inc | 2 | In XiO; when Telepathy is exited; re-entered; and the temporary plan re-called; the OLD beam weight will be displayed but the dose distribution will reflect the modified beam weight. | Computerized Medical Systems, Inc. sent an CUSTOMER ADVISORY letter to all affected customers in May 2008. The letter identified the product, the problem and the action needed to be taken by the customer. The issue was resolved in XiO Release 4.50.00 which has been available since June 2009. A User Notice was sent in March 2011, to all the customers reminding them of the availability of the correction. This notice was sent hardcopy via US Mail but there was no return receipt. For further questions, please call ( 408) 380-8023.
Z-0048-2012 | Oct-14-2011 | 2011 | Abbott m2000rt System software - Abbott m2000rt System software; an automated system for performing fluorescence-based PCR that results in quantitative and qualitative detection of nucleic acid sequences; Abbott Molecular Inc.; Des Plaines; IL 60018 USA; catalog #09K15-01The Abbott m2000rt System provides for real-time measurement of the stages of the polymerase chain reaction (PCR). Real-time PCR measures PCR amplification as it occurs; cycle-by-cycle; allowing quantitative measurements to be made during the highly reproducible exponential phase of PCR. | Abbott Molecular | 2 | The Abbott m2000rt System software has software bug which can affect customers who manually enter a Test Order onto the m2000rt and sorts a column on the calibration or control grids of the Test Order screen by double clicking the column header. This causes a mismatch between specific concentration values and the designated calibrators and controls in the final results; which could cause an inval | Abbott Molecular Inc. sent a Field Correction Recall letter dated September 8, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The customer was instructed to follow the instructions identified until an upgrade to your m2000rt software can be installed. Abbott Molecular estimates that this software version will be available the first quarter of 2012. Please complete the enclosed reply sheet and fax it to (847) 775-6728 at your earliest convenience to acknowledge receipt. We appreciate your assistance in this matter.
Z-2626-2011 | Jun-23-2011 | 2011 | BACTEC" FX System Software - BACTEC" FX System Software; catalog #441398 and #441750The BD BACTEC" FX instrument is designed for the rapid detection of bacteria and fungi in clinical specimens. Samples are drawn from patients and injected directly into BACTEC culture vials; which are placed into the instrument for incubation and testing. | Becton Dickinson & Co. | 2 | In vitro diagnostic instrumentation may exhibit random operational errors that could delay effective patient treatment. | BD Diagnostics Systems sent an "URGENT PRODUCT RECALL" letter dated August 2010 to all affected customers. The letter describes the product, problem and action to be taken. The letter described the type of error messages the customers may experience. All affected units will be scheduled for on site software update as well as a disk drive replacement. A BD field service engineer will contact the customers to schedule the on site visit. Customers experiencing issues prior to the scheduled visit or have questions should contact the BD Diagnostics Service Department at 800-638-8663.
Z-1220-2011 | Feb-11-2011 | 2011 | TotalCare AC powered adjustable hospital bed - Hill-Rom TotalCare Bed; P1900; Hill-Rom; Batesville; IN. TheTotalCare Bed System is intended to porvide patient support ideally suited to be used in health care environments. The TotalCare Bed System may be used in a variety of settings including but not limited to; acute care; including critical care; step down/progressive care medical/surgical; high acuity sub-acute care; post anesthesia care unit (PACU) and sections of the emergency department (ED). | Hill-Rom; Inc. | 2 | The firm received one report that a patient was weighed on the bed at 53 Kg when the patient actually weighed 73 Kg. The patient expired the next day and it is inconclusive as to whether the scale was a contributing factor. The firm investigated and found that the software for the scale system did not clear the last patient weight from the history of the scale system. | The firm, Hill-Rom, sent a "URGENT MEDICAL DEVICE CORRECTION" letter dated January 19, 2011, along with a software upgrade (USB stick/flash drive) with instructions, to all affected customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed to requests that facility service personnel immediately upgrade their software; complete and return the Customer Response Form via fax to 812-931-2410 or mail to: Hill-Rom Company, 1069 State Route 46 E, Batesville, IN 47006, Attn: MOD Coordinator J12 and forward a copy of this notice to any other facility personnel. Note: The customer can request that Hill-Rom technicians complete the software upgrade. If you have any questions concerning this Urgent Medical Device Correction, please contact Hill-Rom Technical Support at 800-445-3720.
Z-2294-2011 | May-25-2011 | 2011 | Zoll - Zoll E Series BLS Defibrillator/Pacemaker/MonitorSoftware versions 7.10 and 7.11The product is used by qualified medical personnel for the purpose of converting ventricular fibrillation to sinus rhythm or other cardiac rhythms capable of producing hemodynamically significant heart beats. In addition; the product may be used in synchronized mode to terminate certain atrial and ventricular tachycardias and other arrhythmias resistant to drug therapy. The E Series may be used for cardiac pacing in conscious or unconscious patients for up to a few hours duration as an alternative to endocardial stimulation. The E Series may be used for the monitoring a patients electrocardiogram (ECG). | ZOLL Medical Corporation | 2 | Zoll E Series BLS Defibrillator may fail to auto charge and deliver energy | The firm, Zoll, sent an "Urgent Medical Device Correction" letter dated April 22, 2011 via UPS to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately alert all E Series BLS users to this problem; instruct users to not touch patient and keep the patient still during ECG analysis, Instruct users that if they experience this problem, to press the Analyze button again, and contact ZOLL Technical Service for the software upgrade kit. The letter also states that affected units can remain in service until the software has been updated. If you have any questions, please contact 24/7 technical support at 1 (800) 348-9011 or +1 (978) 421-9460.
Z-0191-2012 | Nov-16-2011 | 2011 | Spectra Optia Apheresis System - Spectra Optia Apheresis System; REF: 61000; Caridian BCT; Lakewood; CO 80215.A blood component separator used for therapeutic plasma exchanges. | Caridian BCT; Incorporated | 2 | Possible higher than expected platelet loss following therapeutic plasma exchange or plasmapheresis. | CaridianBCT sent a "Safety Alert" letter dated September 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The firm will be sending a sending a recall letter to consignees in November 2011, stating they are voluntarily recalling the software and will provide an upgrade once the software has been approved. If you have any further questions please contact your CaridianBCT Representative, the CaridianBCT Support Center at 1-877-3-FYU-BCT ( US Toll free 1-877-299-4228 ) or (303) 231-HELP ) ( 1 (303) 231-4357, or your local CaridianBCT Customer Service office.
Z-2673-2011 | Jun-27-2011 | 2011 | Instrumentation Laboratory GEM Premier 3500 analyzer - Instrumentation Laboratory GEM Premier 3500 Blood Gas AnalyzerPN 00026000000The GEM Premier 3500 blood gas analyzer is a portable system for use by health care professionals to rapidly analyze whole blood samples at the point of health care delivery in a clinical setting. The instrument provides quantitative measurements of whole blood for pH; pC02; p02; Na+; K+; Ca++; glucose; lactate and hematocrit. These parameters; along with derived parameters; aid in the diagnosis of a patient's acid/base status; oxygen delivery capacity; and electrolyte and metabolite balance. | Instrumentation Laboratory Co. | 2 | Software upgrade: misuse of the barcode gun by repeatedly clicking can cause the instrument to attach test results to an incorrect patient ID. | Instrumentation Laboratories sent a "FIELD SAFETY NOTIFICATION" letter dated April 2011 to all affected customers. The letter described the product, problem, and the actions to be taken by the customers. Included with the letter was an upgrade package containing Software V7.2.2 disk along with installation instructions. Also included with the upgrade kit was a Upgrade Tracking Form to be completed and returned via fax at 781-861-4207 or via e-mail at ra-usa@ilww.com. Customers with questions were directed to contact their local representative.
Z-2662-2011 | Jun-24-2011 | 2011 | Elekta Impac Software Sequencer - Elekta Impac Software Sequencer Verify & Record System; for MOSAIQThe intended use of Sequencer (generically referred to as a verification system) is to assist in the process of process of patient treatment on a radiotherapy treatment machine. It gives the capability to notify users of actions that need to take place prior to treatment; to display reference images for set up purposes and to auto setup the machine to predefined settings. | Impac Medical Systems Inc | 2 | There's a problem in the database conversion program that is used when upgrading from the Multi-ACC?S to MOSAIQ or from one version of MOSAIQ to another. If the gantry start angle is incorrect prior to pressing beam-on; the patient will be mistreated. | The firm, Elekta, sent a "USER NOTICE" to its customers with an attached confirmation of receipt form. The notice describes the product, problem and actions to be taken. The customers were instructed to follow the Interim mitigation instructions noted in the notice -verify the Arc start angle of the Simple ARC with MLC treatment field definition prior to delivery and if the customer identifes a field with an errant gantry start angle, to cease using that plan immediately and consult with the original plan for that field and make the correction to the gantry start angle. In addition, the customers were instructed to complete and return the Confirmation of Receipt via fax to: North America and Asia at 702-992-5002 or Europe and Africa +44 (0)1293 654401, Attn: Elekta OIS Su[pport and/or scan and email to: support@impact.com, europe.support@impact.com or by mail; and distribute this notice to any all users of IMPAC software at their organization who are potentially affected by this issue. For questions or requests for upgrade, email support@impac.com (North America/Asia) or europe.support@impac.com (Europe) or call USA 1-800-488-4672.
Z-0304-2012 | Nov-29-2011 | 2011 | Ambulatory Blood Pressure (ABP) Monitor - Ambulatory Blood Pressure (ABP) Monitor; Model 90217A.The Spacelabs Model 90217A Ambulatory Blood Pressure (ABP) Monitor is a small; lightweight unit designed to take blood pressure and heart rate measurements for a 24 hour; 48 hour; or longer period. | Del Mar Reynolds Medical; Ltd. | 2 | If external pressure is applied to the cuff of the Ambulatory Blood Pressure; Model 90217A; during deflation; the monitor will stop working and display an EC75 error code. | On October 21, 2011, a Medial Device Correction customer letter was mailed to all US customers. A return receipt was requested. The letter identified the affected devices and described the problem. If the device displays an EC75 error code, customers can cycle the power off and then on to restore normal functionality. Please advise patients of this mitigation to avoid a delay in obtaining results. The firm is in the process of upgrading their software to avoid the issue and will contact customers to arrange for an update to their monitors. Customers can call Spacelabs Healthcare at 1-800-522-7025, select 2 for Technical support for any question about this recall.
Z-2567-2011 | Jun-15-2011 | 2011 | XiO RTP System - XiO Radiation Treatment Planning System; XiO Release 4.3.1 through 4.40.00.Used to create treatment plans for any cancer patient for who external beam radiation therapy or brachytherapy has been prescribed. | Computerized Medical Systems Inc | 2 | XiO Software: Under certain conditions; when the port shape is altered; dose is not recalculated. The port-change flag is not being set correctly when the port shape is regenerated. The dose distribution will not match the blocking displayed and the mis-match might not be obvious. Monitor Unit calculations will also be based on an incorrect dose distribution. This problem has existed since Re | Computerized Medical Systems, Inc. sent a "XiO CUSTOMER ADVISORY" letter dated December 8, 2008 to all affected customers. The advisory describes the product, problem, and the actions to be taken. The customers were instructed to set the prompt to "Yes" to avoid the problem from occurring according to written instructions provided in the advisory. This problem has existed since Release 4.3.1 and is expected to be resolved with XiO Release 4.50.00. Questions should be directed to the manufacturer's representative at 408-830-8023.
Z-2599-2011 | Jun-20-2011 | 2011 | FOCAL Workstation - Focal Radiation Treatment Planning System; Focal Release 4.1.1 through 4.40.00.Product Usage: for use in radiation treatment planning using generally accepted contouring methods. The Focal software allows users to access XiO and other RTP systems from a remote PC-based system and provides the user with the ability to edit and modify patient data from a Windows-based PC and transfer the information back to the RTP system. | Computerized Medical Systems Inc | 2 | Focal Software: if individual MLC leaf position edits are made; the original MLC aperture is retained. A subsequent aperture edit should not affect the MLC leaf positions however it was found that instead the leaves will then snap (unintentionally) back to the original aperture. Also; if the patient is sent to XiO after the MLC is edited; an isocenter shift made in XiO will also cause the MLCs to | Elekta CMS Software sent an "ADVISORY NOTICE" dated February 6, 2009 to its affected customers. The notice identified the product, the problem, and the actions to be taken. Customers were informed that the workaround method available is to (edit the MLC pattern using the aperture first and the individual MLC leaf edits second). The firm states that this issue will be resolved in Focal version 4.50.00 and above. A "User Notice" will be sent to all users still using the affected software notifying them of the updated software available to correct the problem. If you have any questions, please call 314-993-0003 or toll free 800-878-4267.
Z-3083-2011 | Aug-29-2011 | 2011 | Fresenius - Fresenius 2008K@homeThe Fresenius 2008K@home is indicated for acute and chronic dialysis therapy in an acute or chronic facility. The 2008K@home is also indicated for hemodialysis in the home and must be observed by a trained and qualified person as prescribed by their physician. | Fresenius Medical Care Holdings; Inc. | 2 | Upgrade software to implement new on screen graphics and instructions; the addition of a bracket on the side of the machine to temporarily hold the bloodlines during set up in order to conform to approved 510(k) | Fresenius Medical NA issued a formal field correction notification dated May 4 ,2011, sent by certified mail with return receipt requested. The letter identified the product the problem and the action needed to be taken by the customer. The customer was notified of the retrofit that would be done to their machines. The patients' facilities were made aware of the changes that would be made to each of their patient's 2008K@home machines on May 4, 2011. On May 13, 2011, all facilities with active 2008K@home machines were sent a notification letter of the retrofit. Please return the receipt for confirmation that you have received this notice. Please call (781) 699-2053, to schedule a certified technician to conduct the machine retrofit.
Z-3282-2011 | Sep-27-2011 | 2011 | ADVANTAGE WINDOWS (Workstation) WITH FUNCTOOL OPTION - ADVANTAGE WINDOWS (Workstation) WITH FUNCTOOL OPTIONGE Medical Systems; SCS.A software module that provides supplemental information to those images extracted from CT and MR temporal datasets. | GE Healthcare; LLC | 2 | GE Healthcare has become aware that image orientation of reference series displayed in Functool Diffusion Tensor/FiberTrak protocol on Advantage Workstation may not match with the orientation annotations displayed in the viewport. This issue may impact patient safety. | GE Healthcare sent "Urgent Medical Device Correction" letters dated August 18, 2011 to all affected customers. The letter included name of product affected and described problem. Recommended actions where provided to customers and they were notified a GE Healthcare representative will contact them for a software system upgrade. For question of this recall contact the GE Service Representative at (800) 437-1171.
Z-1019-2011 | Feb-04-2011 | 2011 | PCR Eleva - PCR Eleva. The marketing brochure is labeled in parts: "Philips Medical Systems is part of Royal Philips Electronics... www.medical.philips.com... Philips Medical Systems Nederland B.V".Provides one or more capabilities relating to the acceptance; transfer; display; storage; and digital processing of medical images. | Philips Medical Systems North America Co. Phillips | 2 | There is a risk of selecting the wrong patient; when patient is selected in the patient list and the examination tab pressed with background query are running at the same time. | On 1/30/2009, Philips began sending out URGENT-Device Correction PCR Eleva notifications to their consignees. The consignees were informed of the hazards as when a patient is selected in patient list and examination tab pressed with background query running at the same time, there would be a risk that the wrong patient is selected. The firm issued the Field Corrective Order (FCO) 73200029 to their representative to correct the problem. The corrective action would be a software upgrade that resolve the issue. All consignees can contact Philips Call Center at 800-722-9377, #5, #2 and reference "FCO 73200029" with any questions.
Z-0901-2011 | Jan-14-2011 | 2011 | Life Pulse High Frequency Ventilator - One Cat. Life Pulse High Frequency Ventilator; Catalog Number: 203A; Manufactured by: Bunnell Incorporated Salt Lake City; Utah; Rx Only. | Bunnell; Inc. | 2 | Product software coding error may compromise patient safety. | Customers were notified by an "Urgent: Medical Device Recall or Correction" letter starting 10/22/10. The letter identified the product and described the issues. Customers were given the option of either returning their ventilator(s) for rework or to receive installation/re-test instructions with an upgrade kit to rework themselves. The letter states that all Life Pulse ventilators in clinical use must be upgraded and there is no charge for the software upgrade.
Z-3100-2011 | Aug-30-2011 | 2011 | 4D Integrated Treatment Console (4DITC) - Varian brand 4D Integrated Treatment Console (4DITC) Radiation Therapy; Interpretation of a RapidArc Treatment Plan as a 3D Conformal Arc Treatment Plan; version v8.6 or 8.8 in combination with treatment planning systems; Model Number: H51; Reference/PSCA Identifier: CP-03063; Product is manufactured and distributed by Varian Medical Systems Inc.; Palo Alto; CAProduct Usage: 4D Integrated treatment console is designed to assist the operator of a radiation device in providing accurate treatment set-ups for each patient by monitoring set up parameters and preventing the radiation therapy device from commencing irradiation while any parameter is out of conformance with the treatment plan. | Varian Medical Systems; Inc. Oncology Systems | 2 | The firm has identified that in certain configurations; the 4DITC can interpret a RapidArc or VMAT treatment field as 3D Conformal Arc Treatment Plan. If a RapidArc or VMAT treatment field is incorrectly interpreted as Conformal Arc; the field will be delivered with a constant dose rate and gantry speed for the entire duration of the delivery; which can lead to delivered dose being different than | Varian sent an "Urgent Medical Device Correction/Urgent Field Safety Notice" letter dated June 27, 2011 to all affected customers. The letter identified the product, problem and recommended user actions to follow. Customers were instructed to notify all appropriate personnel of the correction notice. The letter states that Varian Service will contact customers to arrange to perform update to 4D Integrated Treatment Console (4DITC) software. For further clarification, contact your local Varian Customer Support District or Regional Manager. The letter also provides Varian Oncology Help desk contact information 1-888-827-4265 for USA and Canada and +41417498844 for Europe.
Z-2685-2011 | Jun-28-2011 | 2011 | Beckman Coulter UniCel DxI 800 - Beckman Coulter UniCel DxI 800 Access Immunoassay System with Unicel DXI system software version 4.3 | Beckman Coulter Inc. | 2 | Test results may have incorrect ranges. If a default sample type setting on the tests screen is changed while running Unicel DXI system software version 4.3; the new setting is not saved to the system database. The changed sample type setting will reset to the previously saved setting when power is interrupted to the system; or if an operator reboots the system. If the default sample type i | Beckman Coulter forwarded on 7/6/10 a Product Corrective Action letter with attached PCA Response Form (via US Postal Service for US Customers) to all customers who purchased the UniCel Dxl 600 and Dxl 800 Access Immunoassay systems. The letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to: A revised version of UniCel DxI System Software that corrects this issue is currently in development. Until the new software version is available, Beckman Coulter has provided an alternate method for changing the default sample type setting. See the appendix included with this letter for additional information. Refer to the attached Questions and Answers document for additional details and instructions, and to determine if their laboratory is running UniCel DxI Sytem Software version 4.3. Customers may wish to confirm that proper ranges and units of measure are being reported for any Access assays with sample type settings that were changed while running software version 4.3. Customers were asked to share this information with their laboratory staff, and retain the notification as part of their laboratory Quality Ststem documentation. Complete and return the enclosed response form within 10 days so Beckman Coulter can be assured that consignees received the notification. If they need assistance or have any questions regarding this notification, customers were told to contact Technical Support at 1-800-854-3633 in the United States or Canada. Outside the United States and Canada, contact your local Beckman Coulter Representative.
Z-2063-2011 | Apr-25-2011 | 2011 | Optical Guidance Platform - Optical Guidance Platform; version 2.6 and 2.6.1; Reference/FSCA Identifier: CP-03899; Product is manufactured and distributed by Varian Medical Systems Inc.; Palo Alto; CA Optical Guidance Platform is for use with a charged particle accelerator to perform precise positioning of treatment target for stereotactic radiosurgery or radiotherapy treatments on cranial or extracranial lesions. | Varian Medical Systems; Inc. Oncology Systems | 2 | The optical guidance platform may not be properly enforcing a 24 hour time limit between optical camera recalibrations. | The firm, Varian, sent an "URGENT MEDICAL DEVICE CORRECTION-URGENT FIELD SAFETY NOTICE" letter dated March 3, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately check the Calibration timeout on their optical guidance platform. If this exceeds 24 hour, please set the factory default of 16 hours to ensure that the camera will be calibrated daily. The instructions are located in the document. The customers were also instructed to download the most recent version of the Optical Guidance Platform Instructions for Use at www.MyVarian.com or request a copy from their local Varian Customer Support District or Regional Manager. In addition, the customers were instructed to advise the appropriate personnel working in their radiotherapy department of the content of this letter. Varian is preparing a software patch which will be applied to affected units to correct this issue. If you require further clarification, please feel free to contact your local Varian Customer Support District or Regional Manager; USA and Canada -1888-VARIAN5 (888-827-4265) or Europe- +41 41 749 8844.
Z-2176-2011 | May-10-2011 | 2011 | Optical Guidance Platform; version 2.6 and 2.6.1 - Optical Guidance Platform; version 2.6 and 2.6.1; Reference/FSCA Identifier: CP-03976; Model Number: HZ1; Product is manufactured and distributed by Varian Medical Systems Inc.; Palo Alto; CA Optical Guidance Platform is for use with a charged particle accelerator to perform precise positioning of treatment target for stereotactic radiosurgery or radiotherapy treatments on cranial or extracranial lesions. | Varian Medical Systems; Inc. Oncology Systems | 2 | The anomaly that has been identified with the Optical Guidance Platform (OGP) software may not be always displaying the correct transfer date on the patient file. | The firm, Varian Medical Systems, sent an "URGENT MEDICAL DEVICE CORRECTION/URGENT FIELD SAFETY NOTICE" letter dated March 3, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to review the Optical Guidance Platform (OGP) Instructions For Use on Loading a treatment plan; before positioning the patient, compare the plan details loaded into the OGP with the plan details printed from the treatment planning systems; verify that the patient information, isocenter, gantry and couch angles are identical as requested in the Instructions For Use, and advise the appropriate personnel working in their radiotherapy department of the content of this recall letter. Varian has made changes to the OGP Instructions For Use -Varian strongly recommends you download the most recent version at www.MyVarian.com or request a copy from your local Varian Customer Support District or Regional Manager. Varian also is preparing a software correction which will be applied to affected unites to correct this issue. If you require further clarification, please feel free to contact your local Varian Customer Support District or Regional Manager at the follow: USA and Canada - 1-888-827-4265; Europe +41 41 749 8844; E-mail-North America: support-americas@varian.com; Australia/New Zealand: support.anz@varian.com; Europe: support-emca@varian.com; South East Asia: seasia.apps.helpdesk@varian.com; China/Asia: china.apps.helpdesk@varian.com; Japan: Japan.Apps.Helpdesk@varian.com, and Latin America: soporte.al@varian.com; Internet: Oncology Systems customer site - www.myvarian.com and Varian Medical Systems public site-www.varian.com.
Z-2200-2011 | May-16-2011 | 2011 | Optical Guidance Platform; Version 2.6 and 2.6.1 - Optical Guidance Platform; Version 2.6 and 2.6.1; Model Number: HZl; Manufactured and Distributed by: Varian Medical Systems Inc.; Palo Alto; CA.For use with a charged particle accelerator to perform precise positioning of treatment target for stereotactic radiosurgery or radiotherapy treatments on cranial or extracranial lesions. | Varian Medical Systems; Inc. Oncology Systems | 2 | A software anomaly has been identified with the Optical Guidance Platform (OGP) Software v2.6 and v2.6.1 where the transfer of datasets from treatment planning systems other than the FastPlan system result in a lateral offset error. | On 3/30/2011, Varian distributed an Urgent Medical Device Correction Urgent Field Safety Notice to all affected users, with a description of the problem and identification of the affected product. The letter also informed customers of how they can determine if the problem affects them. The letter states that there is no work around for the issue unless they possess a FastPlan system. Varian is preparing a software patch to correct the issue and it will be applied to all OGP v2.6 and v2.6.1 systems. Customers are to advise the appropriate personnel of the content of the letter. If further clarification is needed, customers are to contact their local Varian Customer Support District or Regional Manager.
Z-3036-2011 | Aug-18-2011 | 2011 | Portable patient monitor with arrhythmia detection or alarms. - Product is the Ultraview Spacelabs Compact monitor; models 91367; 91369; and 91370; with touch screen display and wireless option.Product Usage: Spacelabs Medical PCIS patient monitors passively display data generated from Spacelabs Medical patient parameter modules; Flexport interfaces; gas analyzers; and other SDLC based products. Spacelabs Medical PCIS patient monitors may also function as a generic display or computer terminal. As a generic display or terminal; the PCIS patient monitors allow network based applications to open windows and display information on other networked monitors. PCIS patient monitors is intended to sound various alarm tones that are generated by the specific PCIS patient parameter modules. Each patient parameter module determines (a) when an alarm condition is violated; (b) the alarm priority (i.e. high; medium or low); (c) alarm limits; and (d) when to generate and terminate alarms. PCIS patient monitors are also capable of displaying alarm conditions on other monitors that are on the network through the Alarm Watch feature. Spacelabs Medical PCIS patient monitors are designed to communicate with a variety of external devices such as display devices; network devices; serial devices; user input devices; audio system; local/remote recorders and Spacelabs Medical patient parameter modules. In all applications; Spacelabs Medical PCIS patient monitors provide access to that data on the Spacelabs Medical network either via hardwired or wireless transmission.. Through this network; a patients data may be displayed at the bedside monitor; central station or a clinical information system. | Spacelabs Healthcare; Llc | 2 | Monitor may reset repeatedly which can result in a loss of patient monitoring for several minutes. Cause is a software problem. | Spacelabs Medical sent an "URGENT - MEDICAL DEVICE CORRECTION" notification letter to U.S. consignees by mail on September 17, 2010 and to international consignees by e-mail on September 24, 2010. The letter identified the recalled product, warns customers of the potential hazard and identifies some ways to mitigate the risk if they decide to continue use. The letter states to correct this situation, the recalling firm Field Service Engineers will be contact customers to schedule a time to install a free software update. If customers have any questions about this corrective action program, please contact Spacelabs Medical at 1-800-522-7025 and select 2 for Technical Support, or outside the United States 1-425-657-7200 x5089.
Z-2112-2011 | Apr-28-2011 | 2011 | ABL90 FLEX - ABL90 FLEX Portable Blood Gas Analyzer System; Part Number: 393-090; 510 (k) #K092686. All units equipped with software versions: 2.4.1680.25 and all previous versions are subject to recall. The ABL90 FLEX is a portable; automated analyzer that measures pH; blood Gases; electrolytes; glucose; lactate and oximetry in whole blood. The ABL90 FLEX is intended for use by trained technologists; nurses; physicians and therapists. It is intended for use in a laboratory environment; near patient or point-of-care setting. | Radiometer America Inc | 2 | Following the receipt of two (2) reported customer complaints; the firm determined that the current operating software for their ABL 90 Analyzer does not prevent the further use of the sensor cassette when there is the possibility of iron rich plaque build-up in the reference electrode. If the device operator/user ignores the error message and continues to obtain biased measuring results using the | The firm, Radiometer America, sent two Urgent "Medical Device Field Correction" letters issued on February 4, and on February 24th, 2011. The letters described the product, problem and actions to be taken. The customers were instructed to replace the sensor cassette. In order to prevent this situation from occurring, the firm informed the customers in the recall/correction notification that they intend to upgrade all ABL 90 Flex analyzers to a new software version. The new software will perform a check during a normal system FLUSH in order to allow the analyzer to detect the bias and will notify the customer immediately of this condition and will automatically disable the cassette thus preventing further use. The letter goes on to remind the device customers that until such time as the software upgrade is installed, they should make sure to adhere to the procedures stated in the Reference Manual which call for the replacement of the sensor cassette immediately upon receiving message #1310. The letter further instructs the customer to complete and return the attached FAX RETURN FORM via fax to 440-871-0463. The letter also states that the expected release date for the new software should be sometime in the second half of February 2011. If you have any questions regarding this letter, or the upgrade, please contact Technical Support staff at 1-800-736-0600 Opt 4.
Z-2582-2011 | Jun-17-2011 | 2011 | XiO Radiation Treatment Planning System - XiO Radiation Treatment Planning System. XiO Release 4.50.00 and aboveUsed to create treatment plans for any cancer patient for who external beam radiation therapy of brachytherapy has been prescribed. | Computerized Medical Systems Inc | 2 | XiO: When a new electron beam is added to a plan; the calculation Algorithm is set to Pencil Beam; and an Aperture is added to the beam; the user can select Port-Port Properties and edit the Material Thickness values. If the algorithm is then changed to Monte Carlo; the calculation will proceed and the user is not warned the Material/Thickness values just entered are not used in the dose calculati | Elekta CMS Software drafted a User Notice which explains the reason for correction, the clinical impact, and the workaround method. The User Notice was distributed to all affected sites in November 2010. The issue will be resolved in the XiO Release 4.70.00, expected to be available in July 2011.
Z-3221-2011 | Sep-16-2011 | 2011 | TSX-101: Aquilion 32/64 - TSX-101: Aquilion 32/64; System; X-Ray; Tomography; Computed. | Toshiba American Medical Systems Inc | 2 | There is the potential that the current may not be controlled correctly in examinations using the ECG modulation function; which adjusts the x-ray output in synchronization with the ECG waveform. | Toshiba America Medical Systems (TAMS) issued an Urgent: Medical Device Correction notification to consignees via USPS return receipt mail. The letter identified the affected product, the problem, and corrective action. The firm plans to install revised firmware to prevent the occurrence of the problem. A Toshiba Service Representative will contact customers to schedule an appointment for installation. The firm asked customers to stop using the Power Reduction option during ECG modulated scanning until the new firmware has been installed. Customers are to share the information in the letter with all users, reviewing radiologist, and the clinical engineering or Biomedical group at their facility. Customers are also to complete and return the attached reply form via fax or e-mail to TAMS as proof of notification. Questions pertaining to the letter should be directed to 800-421-1968 or customers' local Toshiba Representative at 800-521-1968.
Z-2214-2011 | May-19-2011 | 2011 | GE BRIGHTSPEED DELIGHT CT SCANNER SYSTEM - BrightSpeed Excel Select; BrightSpeed Edge Select ;BrightSpeed Elite Select and BrightSpeed Elite Computed Tomography X-ray systems with software version 09BW35.11 or 09HW30.4.The GE BrightSpeed Delight Series Computed Tomography X-ray system is intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes; for patients of all ages; including Axial; Cine; Helical; Cardiac; and Gated acquisitions. These images may be obtained either with or without contrast. | GE Healthcare; LLC | 2 | GE Healthcare has become aware of a potential set of circumstances that could cause X-ray continuation during an unexpected table stop on certain BrightSpeed systems (BrightSpeed Excel/Edge/Elite Select and BrightSpeed Elite Computed Tomography X-ray Systems ) if this event were to occur on your system; patient safety may be impacted. | The firm, GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION GE Ref: 22932" letter dated January 19, 2011 to its consignees/customers. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction, Contact Information, and Additional Dosimetric Information. A second letter dated March 11, 2011 was sent to the consignee/customer letter consisting of a cover letter and an attachment. "This letter is to notify you of an error in the January 19, 2011, Urgent Medical Device Correction Letter, that you have already received (GE Ref: 22932). In the Safety Issue section of that letter it is written: "If this issue were to occur, it would result in the remainder of the exposure being delivered at the patient position where the table stopped up to the prescribed scan time or 30 seconds, whichever is shorter." The stated 30 seconds limitation is incorrect for your particular model. This information has been corrected on the attached Urgent Medical Device Correction letter, so please discard the original letter if it is still in your possession. " GE Healthcare will provide a software update for all affected systems to address the issue. A GE Healthcare service representative will contact you to arrange for this correction. The customers were also instructed as follows: prior to GE's installation of a software update, continue to monitor patient scanning closely. Should the table stop unexpectedly and X-ray remains on, manually stop the scan, and ensure that all potential users in your facility are made aware of this safety notification and the recommended actions. If you have any questions, please contact GE immediately at the following: United States: 800-437-1171; Japan: 0120-055-919; Korea: 1544-6119; Australia/New Zealand: 800-659-465 and China: 800-810-8188. For other countries please contact your local GE Healthcare Service Representative if you have any questions or concerns regarding this notifi.
Z-2559-2011 | Jun-14-2011 | 2011 | Focal Release 4.34.00 through 4.50.00 software - Focal Radiation Treatment Planning System; Focal Release 4.34.00 through 4.50.00.For use in radiation treatment planning using generally accepted contouring methods. The Focal software allows users to access XiO and other RTP systems from a remote PC-based system and provides the user with the ability to edit and modify patient data from a Windows-based PC and transfer the information back to the RTP system. | Computerized Medical Systems Inc | 2 | Focal Software: When a patient is defined as prone and the shifts are manually entered and the LOCK SHIFT box is checked the X coordinate (left/right) and the Z coordinate (anterior/posterior) are reversed (left to right; anterior to posterior). This could potentially lead to incorrect isocenter definitions from prone patients and thus a mislocation of dose. | Computerized Medical System sent a CUSTOMER ADVISORY dated June 12, 2009, to all affected users. This advisory identified the product, the problem, and the action needed to be taken by the customer. There were no patients mistreated as a result of this issue The firm will send out a User Notice to all sites still running affected software to remind them that the issue resolution is available. For further questions, please call (403 ) 830-8023.
Z-3111-2011 | Aug-31-2011 | 2011 | IntelliVue Clinical Information Portfolio - Philips Medical:865047 IntelliVue Clinical Information Portfolio Release D.03865209 IntelliVue Clinical Information Portfolio Release E.00 and E.01865300 IntelliVue Clinical Information Portfolio NICU LinkIntended for use in data collection; storage; and management with independent bedside devices; and ancillary systems that are connected either directly or through networks. It is indicated for use by healthcare providers whenever there is a need for generation of a patient record and computation of drug dosage. | Philips Healthcare Inc. | 2 | Under certain circumstances; edits to an order can result in unintended scheduled interventions on the IntelliVue Clinical Information Portfolio. A patient may receive interventions or additional medication doses that are not intended. | Philips Healthcare sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated July 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. Customers are asked to follow the guidelines in the "Action to be taken by Customer/User" section of the Urgent Medical Device Correction Notice. Philips has developed a software upgrade to address this problem in the short term. This upgrade will be available immediately. This software upgrade will be performed by Philips service personnel. While customers await the correction of their devices, customers are instructed to follow the instructions for use for the device, which includes the directions to: 1. Always use the medication administration record (MAR) in determining patient care. 2. Always review orders and pending administrations in the administration record after any changes to an order are stored. For the longer term, Philips is developing safeguards intended to minimize the potential for similar errors resulting from order changes and upgrade to be available in September 2011. If you need any further information or support concerning this issue, please contact your local Philips representative.
Z-0896-2011 | Jan-12-2011 | 2011 | Ultrasound System - Acuson Antares Ultrasound SystemDiagnostic Ultrasound system | Siemens Medical Solutions USA; Inc. | 2 | Software safety guard measures on the device malfunctioned to prevent the temperature from going over its limits when the device is used. | Field notification in the form of a Customer Safety Advisory Notice was initiated on 09/30/2009 by the firm, and the notice was mailed to all affected customers Siemens sent a CUSTOMER SAFETY ADVISORY NOTIFICATION letter on September 30, 2009. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to disconnect and reconnect the V5Ms transducer between uses until their local Customer Service Engineer installs the software update to their system. For any questions customers were to contact their local service support person. For questions regarding this recall call 650-694-5398.
Z-3035-2011 | Aug-18-2011 | 2011 | Horizon Medical - Horizon Medical Imaging versions 11.5; 11.5.1; 11.5.2; 11.6; 11.7 and 11.8.Horizon Medical Imaging is a medical image and information management application that is intended to receive; transmit; store; retrieve; display; print and process digital medical images; digital medical video; and associated medical information from various medical imaging systems. | Mckesson Information Solutions LLC | 2 | Images for one patient may appear in the study of another patient. | Mckesson Medical Imaging Group sent a URGENT FIELD SAFETY NOTICE letter dated June 13, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. On February 2, 2011, the McKesson Medical Imaging company initiated a process for applying an image display software update to the HMI systems at the hospital site that initiated the complaint. They were advised to contact McKesson Customer Support at 1-800-663-2533, to indicate whether they would like to receive the software update. They were advised to distribute the notification to all those within their organization who need to be aware and to alert other parties affiliated with their organization that are affected by this advisory notice.
Z-1011-2011 | Feb-02-2011 | 2011 | ABL90 FLEX Blood Analyzer - ABL90 FLEX Blood Analyzer Part Number: 393-090. Intended for use by trained technologists; nurses; physicians and therapists and for use in a laboratory environment; near patient or point-of-care setting. Also; Intended for in vitro testing of samples of heparinized whole blood. | Radiometer America Inc | 2 | Due to a software error; a patient mix up can occur on the ABL90. If a sample is pre-registered and the sample "fails" the next result will inherit the patient data from the previous result. | On 11/16/10 the firm sent Medical Device Field Correction letters to their customers. The letter states that all systems will be upgraded and until each system is upgraded, they instructed their customer to "always query for patient ID prior to introducing the sample." They also instruct each customer to complete the "Fax Return Form" acknowledging receipt of the letter. If customers have questions regarding the letter or the upgrade, they should contact Technical Support at 1-800-736-0600 Opt 4.
Z-2114-2011 | Apr-29-2011 | 2011 | UniCel DxH 800 Coulter Cellular Analysis System - UniCel DxH 800 Coulter Cellular Analysis System Part Number: 629029The UniCel DxH 800 Analyzer is a quantitative; automated hematology analyzer for in vitro diagnostic use in screening patient populations found in clinical laboratories. The UniCel DxH 800 Analyzer provides: - Complete Blood Count (CBC); Leukocyte 5 Part Differential (Diff); Reticulocyte (Retic) and Nucleated Red Blood Cell (NRBC) on whole blood - Total Nucleated Count (TNC) and Red Cell Count (RBC) on Body Fluids (cerebrospinal; serous and synovial) (BF) | Beckman Coulter Inc. | 2 | For Panels containing CBC: The RBC and HCT are improperly corrected and not flagged when there is a WBC Pump Failure and the WBC ~ 140x103/uL. For Body Fluid Panels:The RBC is improperly corrected and not flagged when there is a WBC Pump Failure. IMPACT:Erroneous RBC and/or HCT could affect the interpretation and management of patients at risk with various types of anemias. | The recall communication was initiated on 10/20/2010, with Beckman Coulter forwarding a Product Corrective Action (PCA) letter with PCA Response Form (via US Postal Service for US customers) to all customers who purchased the UniCel DxH 800 Coulter Cellular Analysis System. This letter informs the customers about the problems identified and provides an action/resolution to this issue. When a WBC Pump Failure occurs, the WBC or TNC results will be flagged with 'R' and the WBC Pump Failure System message. The RBC and/or HCT results associated with these affected analyses also require review. This will be corrected in a future version of software. Beckman Coulter asked consignees to complete and return the enclosed response form within 10 days and share this information with laboratory staff and retain this notification as part of laboratory Quality System documentation. Questions were directed to Beckman Coulter Customer Service at 800-526-7694 in the United States or Canada, or consignees local Beckman Coulter representative.
Z-2121-2011 | Apr-29-2011 | 2011 | UniCel DxH 800 Coulter Cellular Analysis System - UniCel DxH 800 Coulter Cellular Analysis System Part Number: 629029The UniCel DxH 800 Analyzer is a quantitative; automated hematology analyzer for in vitro diagnostic use in screening patient populations found in clinical laboratories. The UniCel DxH 800 Analyzer provides: - Complete Blood Count (CBC); Leukocyte 5 Part Differential (Diff); Reticulocyte (Retic) and Nucleated Red Blood Cell (NRBC) on whole blood - Total Nucleated Count (TNC) and Red Cell Count (RBC) on Body Fluids (cerebrospinal; serous and synovial) (BF) | Beckman Coulter Inc. | 2 | RETIC results may be invalid when the instrument user interface allows reflex processing of a pre-diluted specimen in the RETIC test panel. Beckman Coulter has validated pre-diluted specimen processing in CBC panels. RETIC test panel using pre-diluted samples has not been validated.IMPACT:The reticulocyte value is an indicator for assessing red cell production. Erroneous results could impact c | The firm, Beckman Coulter, issued an "URGENT: PRODUCT CORRECTIVE ACTION (PCA)" letter dated October 20, 2010, with PCA Response Form to all customers who purchased the UniCel DxH 800 Coulter Cellular Analysis System. This letter described the product, problem and actions to be taken. The customers were instructed to do the following: 1) Remove all cassettes from the mix station prior to initiating the Diluent Dispense function. Do not initiate Diluent Dispense while specimens are being processed in cassette presentation. 2)Do not use pre-diluted samples for RETIC counts. The instrument is validated for pre-diluted CBC's only, as indicated in the Instructions for Use (IFU). 3)Verify that their information Technology personnel are aware that the Body Fluids data string is only transmitted to the LIS using the US numeric separators and that the Body Fluid results are properly validated at the LIS. 4) Use the lip on the reagent drawer to move the drawer as shown. Avoid touching the drawer slides while replacing containers. Clean all spills promptly to reduce the risk of exposure to blood borne pathogens. 5)Complete and return the enclosed PRODUCT CORRECTIVE RESPONSE FORM within 10 days, share this information with laboratory staff and retain this notification as part of laboratory Quality System documentation. Beckman Coulter will make corrections in future version software. Questions were directed to Beckman Coulter Customer Service at 800-526-7694 in the United States or Canada, or consignees/customers local Beckman Coulter representative.
Z-2128-2011 | May-02-2011 | 2011 | UniCel DxH 800 Coulter Cellular Analysis System - UniCel DxH 800 Coulter Cellular Analysis System Part Number: 629029The UniCel DxH 800 Analyzer is a quantitative; automated hematology analyzer for in vitro diagnostic use in screening patient populations found in clinical laboratories. The UniCel DxH 800 Analyzer provides: - Complete Blood Count (CBC); Leukocyte 5 Part Differential (Diff); Reticulocyte (Retic) and Nucleated Red Blood Cell (NRBC) on whole blood - Total Nucleated Count (TNC) and Red Cell Count (RBC) on Body Fluids (cerebrospinal; serous and synovial) (BF) | Beckman Coulter Inc. | 2 | Body Fluids results may be incorrectly interpreted by the Laboratory Information System (LIS). Body Fluid results ~ 1;000 celis/uL are reported using the US numeric separators to demark thousandths places (comma; ";") and partial units (decimal; "."). An LIS that uses non-US formats could interpret the comma as a partial units separator; which could lead to reporting erroneously low body fluid r | The firm, Beckman Coulter, issued an "URGENT: PRODUCT CORRECTIVE ACTION (PCA)" letter dated October 20, 2010, with PCA Response Form to all customers who purchased the UniCel DxH 800 Coulter Cellular Analysis System. This letter described the product, problem and actions to be taken. The customers were instructed to do the following: 1) Remove all cassettes from the mix station prior to initiating the Diluent Dispense function. Do not initiate Diluent Dispense while specimens are being processed in cassette presentation. 2)Do not use pre-diluted samples for RETIC counts. The instrument is validated for pre-diluted CBC's only, as indicated in the Instructions for Use (IFU). 3)Verify that their information Technology personnel are aware that the Body Fluids data string is only transmitted to the LIS using the US numeric separators and that the Body Fluid results are properly validated at the LIS. 4) Use the lip on the reagent drawer to move the drawer as shown. Avoid touching the drawer slides while replacing containers. Clean all spills promptly to reduce the risk of exposure to blood borne pathogens. 5)Complete and return the enclosed PRODUCT CORRECTIVE RESPONSE FORM within 10 days, share this information with laboratory staff and retain this notification as part of laboratory Quality System documentation. Beckman Coulter will make corrections in future version software. Questions were directed to Beckman Coulter Customer Service at 800-526-7694 in the United States or Canada, or consignees/customers local Beckman Coulter representative.
Z-3077-2011 | Aug-29-2011 | 2011 | MagNA Pure LightCycler - MagNA Pure LC 1.0 (software version 3.0.11).Roche Diagnostics Operations; Inc.An automated instrument used for isolation of nucleic acids (DNA; total RNA; total viral nucleic acids) from different kinds of biological research sample material (whole blood; serum; blood cells; culture cells; tissue; bacteria; fungi) using the specially designed MagNA Pure LC kits; for the purpose of life science research only and automated filling of different kinds of PCR reaction vessels (LightCycler Capillaries; 96-well PCR plates; PCR tube strips) with PCR reaction mixes and template nucleic acid; or other reaction tubes for the making of dilution series; reaction mixes. | Roche Diagnostics Operations; Inc. | 2 | Roche has detected a problem with the MagNA Pure LC 1.0 (software version 3.0.11) and the MagNA Pure LC 2.0 (software versions 1.1.23 and 1.1.24) instruments. Under certain circumstances; only a fraction of each sample is processed which may lead to false negative results; and an impact to patient health if used in a homebrew diagnostic application. | Roche Diagnostics Operations, Inc. sent a letter dated July 21, 2011 entitled "URGENT MEDICAL DEVICE CORRECTION" to all affected customers. The letter included the name product, software versions, and reason for recall and informed them the problem is under investigation. In the meantime, customers are asked to wait 30 seconds after the finalization of the decontamination cycle before initiating the next operation, in order to avoid this issue. For questions on this recall please call Roche at (800) 428-5074.
Z-2890-2011 | Jul-28-2011 | 2011 | BD BACTEC(tm) MGIT (tm) 960 System - BD BACTEC(tm) MGI (tm) System Software; Catalog 445922; Lot V5.02A.Product Usage: The BACTEC(tm) MGIT (tm) System Software is used in the MGIT 960 and MGIT 320 instruments for the rapid detection of mycobacteria from clinical specimens. In addition; an Antimicrobial/Drug Susceptibility Testing (AST or DST) system is used as a rapid qualitative procedure for susceptibility testing of Mycobacterium tuberculosis; from culture; to streptomycin; isoniazid; rifampin; ethambutol; and pyrazinamide. The routine AST set consists of a Growth Control tube and a tube for each drug; as well as a barcoded tube carrier that holds the set. The carrier barcode provides set size and identification. A tube rack capable of holding up to 55 AST tubes in various set combinations is available for tube setup and transport. Large or small AST carrier sets can be placed in the rack. The MGIT AST system is based on growth of the Mycobacterium tuberculosis strain under test as measured by the instrument. AST sets are entered into the BACTEC MGIT instrument and are tested automatically. Analysis of fluorescence in a drug containing tube is compared to the fluorescence of the Growth Control tube to determine susceptibility results. | Becton Dickinson & Co. | 2 | Software contained three additional set configurations which caused the system to reinitialized the AST set carrier defaults. | BD Diagnostic Systems verbally notified all affect customers via the telephone on March 28, 2011. The notification identified the product, the problem, and the action to be taken by the customer. An "URGENT PRODUCT RECALL" follow-up letter was mailed to all customers on April 8, 2011 via UPS. Customers were instructed to complete and fax the attached reply form to acknowledge their receipt of the notification to Regulatory Compliance, 410-316-4256. The letter instructs customers to refer to the modified release notes and IMPORTANT NOTICES outlined in the attached notification letter. If further assistance is needed regarding replacements and all other inquires, please contact BD Technical Services and Support at 1-800-638-8663.
Z-1957-2011 | Apr-12-2011 | 2011 | Varis Vision - RT Chart - Varian Medical Systems; Varis Vision - RT Chart version 7.3.10. For radiation therapy. | Varian Medical Systems; Inc. Oncology Systems | 2 | An issue with the Varis Vision software where treatment can occur without the dynamic Multileaf Collimator component and may deliver a higher than expected dose of radiation. | Varian Medical Systems Urgent Medical Device Correction letters were sent to all affected customers on February 7, 2011.
Z-2407-2011 | Jun-02-2011 | 2011 | XiO Radiation Treatment Planning System - XiO Radiation Treatment Planning System; XiO Release 4.50.00 and aboveUsed to create treatment plans for any cancer patient for who external beam radiation therapy or brachytherapy has been prescribed. | Computerized Medical Systems Inc | 2 | XiO Software: For a specific clinical setup; the MLC leaf positions on the exported DRR images might not match the MLC leaf positions displayed in XiO. | The firm, Elekta, issued an "important Safety Notice" undated, to its customers. The notice indicates the problem and the clinical impact. There is no reasonable workaround for this issue. The notice states a future XiO Release will resolve the issue and the customers will be notified when this solution is available. A return postcard was included with the mailing to verify consignee received the notice. For sites not returning the postcard within one month, the recalling firm will contact them via phone or email to confirm receipt of the notice. The customers were instructed to distribute this notice to any and all users of the CMS software at their organization who are potentially affected by this issue. If you have any questions, please call 314-993-0003 or 800-878-4267.
Z-0037-2012 | Oct-12-2011 | 2011 | LATITUDE¿ Paceart Integration Software - LATITUDE¿ Paceart Integration Software; version 1.00 Model 6472.LATITUDE Paceart Integration is software and therefore does not have any serial numbers. However; the software version affected is version 1.00 and is distributed on USB pen drives. These pen drives are lot number controlled and denoted as Model 6493.When used in conjunction with the LATITUDE EMR Integration software transfers data files from the Boston Scientific LATITUDE Patient Management system to a clinic's Paceart system. | Boston Scientific CRM Corp | 2 | Boston Scientific has determined that under certain conditions; LATITUDE Paceart Integration (LPI) softwaremay not properly transfer data resulting in the incorrect summary of programmed parameter values in the ICD Summary tab display and Current Programming Tachy summary section of the Paceart report. | Consignees were sent on 8/12/11 a Boston Scientific letter dated August 2011. The letter was addressed to "Dear Doctor". The letter described the product and problem. It also gave Background Information, Description of Potential Data Translation Error, and Additional Information and Action information. A letter, including an HRS form, was also sent to all US LATITUDE Clinics that have active LPI software to notify them of the need to install the corrected LPI Software.
Z-2175-2011 | May-10-2011 | 2011 | Philips IntelliVue Clinical Information Portfolio - Philips IntelliVue Clinical Information Portfolio software as follows:865047 IntelliVue Clinical Information Portfolio Release D.03; 865209 IntelliVue Clinical Information Portfolio Release E.0 and E.01; 865300 IntelliVue Clinical Information Portfolio NICU Link | Philips Healthcare Inc. | 2 | Edits to an order can result in unintended scheduled interventions and a patient may receive additional interventions or medication doses that are not intended. | Philips Healthcare notified accounts on 3/16/11 via the Urgent Medical Device Correction Notification. Customer notifications in the United States will be delivered via UPS. Customer notifications outside the US will be managed by Philips representatives in each affected geography. Customers will be asked to follow the guidelines in the Action to be taken by Customer/User section of the Urgent Medical Device Correction Notice until they receive their software correction. The correction will be in the form of a software upgrade and will be performed by Philips service personnel. All customers potentially affected by this issue will receive the software correction as part of the field action.
Z-1907-2011 | Apr-08-2011 | 2011 | 8V3c transducer - 8V3c transducer used with the ACUSON SC2000 ultrasound system. 8V3c transducer model number 10433823; ACUSON SC2000 model number 10433816.Siemens Medical Solutions USA; Inc.Business Unit Ultrasound1230 Shorebird Way;Mountain View; CA 94043High-density array transducer for pediatric and fetal echocardiography applications. | Siemens Medical Solutions USA; Inc. | 2 | When using the 8V3c transducer in color mode; the derated Ispta 3 (Spatial Peak-Temporal Average Intensity) may exceed the limit stated in the instructions for use. | A Customer Safety Advisory Notification was initiated on 2/4/2011 and sent via certified mail. The letter identified the affected product and described when the problem occurs along with the associated potential risk. The firm recommends that users not use the 8V3c until the software issue is fixed. The firm is working to release a software update and customers' local Customer Service Engineer will install new software on their system when it is available. Customers should direct their questions to their local service support person. The information in the letter should be shared with all personnel within the organization who need to be aware.
Z-1010-2011 | Feb-01-2011 | 2011 | Fraxel re:store¿ Dual Laser System - Fraxel re:store¿ Dual Laser System----a laser surgical instrument for use in general and plastic surgery and in dermatology. Model Numbers: MC-SYS-SR1500-D-US; MC-SYS-SR1500-D-US-LOANER; MC-SYS-SR1500-D-UPG-US; MC-SYS-SR 1500-0-1 NTL; MC-SYS-SR1500-D-I-LOANER; MC-SYS-SR1500-D-UPG-INTL; Product is manufactured and distributed by Solta Medical; Inc.; Hayward; CA The Fraxel re:store¿ Dual Laser System is indicated for use in dermatological procedures requiring the coagulation of soft tissue. | Solta Medical; Inc. | 2 | The product has the potential to cause inadvertent laser firing which may lead to patient/operator injury if the problem recurs. | The firm, Solta Medical, issued three "DEVICE CORRECTION" letters, one dated December 15, 2010, the second one date December 17, 2010 and the third one dated December 20, 2010, to all affected customers. The letters described the product, problem and action to be taken by the customers. The customers were instructed to follow the instructions carefully, sign and date the bottom of the instruction form upon completion of the update and promptly return the signed form and USB stick in the enclosed, return-addressed envelope (Product Support, Solta Medical, 25881 Industrial Blvd., Hayward, CA 94545) Note: Solta Medical sales personnel will assist with the software update process so, if your local representative has already updated your system, you do not need to do anything further. If you have any questions, please call Solta Product Support at (510) 259-7291.
Z-3108-2011 | Aug-31-2011 | 2011 | TomoTherapy Hi-Art System - Hi¿Art versions 4.0.x; and HD versions 1.0.x; H-0000-0003.The TomoTherapy HI-ART System is intended to be used as an integrated system for the planning and precise delivery of radiation therapy; stereotactic radiotherapy; or stereotactic radiosurgery to tumors or other targeted tissues while minimizing the delivery of radiation to vital healthy tissue. The megavoltage x-ray radiation is delivered in a rotational; non-rotational; modulated (IMRT); or non modulated (non-IMRT/three dimensional conformal) format in accordance with the physician approved plan. | Accuray Incorporated | 2 | As a result of an internal review; TomoTherapy has identified an issue with Hi??Art versions 4.0.x; and HD versions 1.0.x that we would like to bring to your attention.During DICOM export of plan level images with a non square exported Field of View (FOV); an anomaly in the process of squaring the plan level image may cause the image to shift with respect to ROIs and dose. When the anomaly occu | TomoTherapy sent a "Urgent Medical Device Correction" letter dated July 13, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. This issue will be remedied with HiÂ®Â®Art software version 4.1 and HD software version 1.1 and later. If you have any questions, please contact the TomoTherapy Technical Solutions Center by e-mail "support@tomotherapy.com" . Please refer to the letter for what phone number you should call depending on where you are.
Z-0272-2012 | Nov-23-2011 | 2011 | AdminTool2 - IMPAX CardioVascular (CV) Admin ToolThe IMPAX Cardiovascular suite is a cardiovascular information system providing image archiving and image display and modality study-specific structured reporting. | AGFA Corp. | 2 | Potential for misattribution of data and for a report to become unavailable when the user requests for it to be moved to another study that already has a report associated with it. | An 'URGENT FIELD SAFETY NOTICE" was sent via FED-EX by June 6, 2011 to all the sites using the affected product. The letter described the issue and mitigation. The letter recommends that users not move reports to studies where a report already exists. If a report is being moved because the existing report is irrelevant to the study, then the existing report should be deleted prior to moving the new report to that study. If the report is being moved because the user wants to combine the content of the reports, then the "merge" function should be utilized. The letter also states that a software enhancement is currently available. Acknowledgment via FAX-Back was requested from the sites. Questions regarding the issue should be directed to customers' local Agfa HealthCare Service at 877-777-2432.
Z-2592-2011 | Jun-20-2011 | 2011 | Focal Radiation Treatment Planning System - Focal Radiation Treatment Planning System; Focal Release 4.50.00 and abovefor use in radiation treatment planning using generally accepted contouring methods. The Focal software allows users to access XiO and other RTP systems from a remote PC-based system and provides the user with the ability to edit and modify patient data from a Windows-based PC and transfer the information back to the RTP system. | Computerized Medical Systems Inc | 2 | Focal: for XiO plans calculated in relative dose mode; the Focal Plan Review application will display the DVH and maximum dose correctly. However; the isodose display is not scaled correctly. If the user evaluated the plan based only on the isodose display; they might misinterpret the actual dose delivered to the patient. | Elekta / Computerized Medical Systems Inc sent an undated Safety Notice to consignees beginning on May 21, 2010. The notice identified the product, the problem, and a workaround.
Z-2572-2011 | Jun-15-2011 | 2011 | Focal Radiation Treatment Planning System; - Focal Radiation Treatment Planning System; Focal Release 4.3.1 through 4.34.02.For use in radiation treatment planning using generally accepted contouring methods. The Focal software allows users to access XiO and other RTP systems from a remote PC-based system and provides the user with the ability to edit and modify patient data from a Windows-based PC and transfer the information back to the RTP system. | Computerized Medical Systems Inc | 2 | Focal: In absolute dose mode a beam with a port may be calculated with open normalization; with the result that the displayed dose reflects open normalization while the monitor units and QA plan reflect blocked normalization. This is particularly important when these beams also include compensators. The problem is detectable by looking at the delivered dose to the weight point in the patient as | Computerized Medical Systems Inc sent a "CUSTOMER ADVISORY" letter dated March 2008 to all affected customers. The letter describes the product, the problem and the actions to be taken. The firm provided a workaround for customers that states "Alternatively the plan can be imported directly into XIO". A copy of the Advisory was shipped with software within the affected range to those requesting a software upgrade. The issue was resolved in Focal Release 4.40.00 introduced in June 2008. The firm will send out a User Notice to all sites still running affected software to remind them that the issue resolution is available. For further questions, please call (403) 830-8023.
Z-3084-2011 | Aug-29-2011 | 2011 | Maquet HL Console for Heart Lung Machine HL20 - Maquet HL Console for Heart Lung Machine HL20HL 20 4-umps console base: P/N 917741HL 20 5-pumps console base: P/N 917751Intended for use as an extracorporeal circulation device for perfusion lasting not more than six hours. The HL console is equipped with a rack for accommodating the system's electronic modules. | Maquet Inc. | 2 | Maquet HL 20 Twin Pump Modules may unintentionally increase speeds for a 1.5-2 second period of time before returning to its original speed. | Maquet sent a PRODUCT RECALL NOTIFICATION LETTER/return response forms, dated June 17, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use of the affected product until a software update was performed by their trained technicians. Customers were also asked to forward the information contained in the notification letter to all users and to all appropriate Biomedical Staff. Customers were instructed to acknowledge receipt of the notification letter by completing and faxing back the attached response form to 973-807-9210 whether or not they have any of the affected product. For questions customers were instructed to contact their local MAQUET representative. For questions regarding this recall call 973-709-7994.
Z-0212-2012 | Nov-18-2011 | 2011 | IMPAX CV Echo Measurement Import via Optical Character Recognition (OCR) - IMPAX CV Echo Measurement Import via Optical Character Recognition (OCR) The IMPAX Cardiovascular suite is a cardiovascular Information system providing image archiving and image display and modality study-specific structured reporting. | AGFA Corp. | 2 | Incorrect entries in mapping files could lead to inaccurate measurement display in adult and pediatric echo reports when using the IMPAX CV Echo Measurement Import via Optical Character Recognition (OCR). | The firm, AGFA HealthCare" sent an 'URGENT FIELD SAFETY NOTICE" dated May 27, 2011 via FED-EX by June 6, 2011 to all customers using the affected product. The letter describes the product, problem and actions to be taken. The customers were instructed to distribute this information within their facility to all those who need to be aware and to complete and return the attached FEEDBACK FORM via fax to 864-421-1668 or email to debbie.norris@agfa.com. AGFA will provide software corrections to all affected sites and will contact the customers to schedule the analysis and provide any corrections if required. If you have any questions about this matter, please contact your local Agfa HealthCare Service at 877-777-2432.
Z-3184-2011 | Sep-13-2011 | 2011 | Straumann coDiagnostiX - Straumann coDiagnostiX; version 8.0; used to guide a clinician in the proper implant selection; oral placement and the proper selection of an oral drill.Imaging processing System | Straumann Manufacturing Inc. | 2 | Software version 8.0 has a software anomaly; in which the length is not adjusted accordingly when it is changed after the surgical sleeve position planning step has been completed. | Staumann USA notified accounts by an Urgent Field Safety Notice letter dated August 4, 2011 and delivered via Federal Express Delivery Service. The letter identified the affected product and explained the reason for the recall. Customers were asked to review the letter, take appropriate actions, including completing and faxing the enclosed. Customer Confirmation Form. Straumann released a corrected software upgrade (coDiagnostiX 8.0.2.3672), which is available for download when users go online and activate current software. If the user does not have internet access a replacement CD can be ordered from Technical Support at 866 - 531 -7365. Questions should be directed to Technical Support Monday through Friday from 8:00 am to 8:00 pm Eastern Time.
Z-2201-2011 | May-19-2011 | 2011 | EVOLIS Assay Protocol Files ( APF) Software - Software CD is labeled in part:"BLOOD VIRUS DIVISION***BIO-RAD EVOLIS System APF Version 4.4 Software Version 2.00 2009/05***"The EVOLIS Assay Protocol Files (APF) Software contains assay-specific instructions necessary for the EVOLIS Microplate processing system to process enzyme immunoassays on the EVOLIS instrument; perform data reduction and generate reports. | Bio-Rad Laboratories | 2 | The EVOLIS Assay Protocol Files Software is configured in such a way that may result in a miscalculation of quantitative test results and/or an error in reporting positive/negative sample values. | Bio-Rad Laboratories sent an Urgent Product Correction letter dated July 19, 2010, to all customer sites via FedEx. There are two issues related to The EVOLIS Assay Protocol Files (APF) Software that could potentially affect patient results. Issue 1 involves quantitative sample test/calculations/results interpretation for assays with linear regression, point to point or 4PL data reduction. Issue 2 involves sample test results with " Result Report Flags''. A software patch to correct these two issues will be sent to each customer site. Bio-Rad Laboratories will work with each customer to complete documentation confirming that the software patch has been successfully installed on each instrument. If you have any questions or concerns regarding these two issues, please contact Technical Support at 1-800-224-6723, option 2 and then option 3.
Z-0149-2012 | Nov-02-2011 | 2011 | Medtronic CryoCath Arctic Front CryoAblation Catheter - Medtronic Arctic Front Catheter; Sterile EO; Medtronic CryoCath LP; Kirkland; Quebec; H9H 5 H3; Canada; Model 2AF282 and 2AF231 ( not distributed in US)Product Usage: The Medtronic CryoCath System consists of a Cryoablation console; cardiac Cryoablation catheters; connection components and accessories. The Arctic Front Cardiac Cryoablation Catheter is an over the wire Balloon catheter. The inside of the Balloon at the distal end of the Cryoablation Catheter reaches temperatures of -50¿C or colder when refrigerant is injected from the console to the inside of the balloon. Arctic Front Cardiac Cryoablation catheters are introduced into the vasculature by traditional minimally invasive techniques. The CryoConsole houses the electrical and mechanical components as well as proprietary software for controlling and recording a cryotherapy procedure with Arctic Front and/or the Freezor family of catheters. It stores and controls the delivery of the liquid refrigerant through the coaxial umbilical to the catheter; recovers the refrigerant vapor from the catheter under constant vacuum; and disposes of the refrigerant through the hospital scavenging system. Multiple features are built into the catheter and the CryoConsole system to ensure safety. | Medtronic Inc. Cardiac Rhythm Disease Management | 2 | Medtronic CryoCath has identified specific lots of Arctic Front CryoAblation Catheters that were distributedwith incorrect programming that may result in false expired catheter notifications. Impacted lots are 21571;25375; 25376; 25377; and 25381 of the Arctic Front Catheter Model Number 2AF282. Medtronic CryoCathis communicating this information to the FDA.On affected catheters; the incor | Medtronic sent an "Important Device Recall" letter dated August 2011, hand delivered by Medtronic Field Representatives on August 29, 2011 to all affected customers. The letter described the product affected, the problem, and the actions needed to be taken. The customers were instructed to return all unused catheters from affected lots for credit and replacement. The letter states that Medtronic AF Solution representative will assist customers with the return of affected product. For further questions contact your local Medtronic AF Solutions representative.
Z-1495-2011 | Mar-02-2011 | 2011 | PhD System EIA/IFAversion 2.1A software - PhD System EIA/IFAversion 2.1A software; model number 426-0246; a component of the Bio-Rad PhD System. Manufactured by Bio-Rad Laboratories; France; Distributed by Bio-Rad Laboratories; Hercules; CA.EIA and IFA processing system that consists of a PhD fluid handling station (s) linked to a network computer via ethernet hub. The computer provides worklist generation; data management; data reduction and microplate reader control functions. | Bio-Rad Laboratories Inc | 2 | Incorrect patient results are being generated from out-of-range samples tested withc ompetitive enzyme immunoassays on the PhD system with software version 2.1A. | BIO-RAD sent an URGENT: MEDICAL DEVICE CORRECTION letter dated September 29, 2009, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately examine their customer records to identify all PhD System customers that are using version 2.1A of the EIA/IFA Software and Methods CD-ROM. Customers who are not running competitive immunoassays on the PhD System may continue using the 2.1A software. Customers who run competitive immunoassays must revert to software version 2.0B. Customers were instructed to translate if necessary and add local contact information to the customer letter and distribute to all PhD System customers that are using 2.1A of the EIA/IFA Software and Methods CD-ROM. A Customer Medical Device Correction Response Form must be collected from each customer to ensure that they have received this important communication. Once forms have been collected from all applicable customers, they should be faxed to Bio-Rad CSD Regulatory Affairs Department along with the completed Subsidiary Medical Device Correction Response Form. For questions regarding this recall call 510-741-3954.
Z-2596-2011 | Jun-20-2011 | 2011 | Focal Radiation Treatment Planning System - Focal Radiation Treatment Planning System; Focal Release 4.3.0 through 4.34.01For use in radiation treatment planning using generally accepted contouring methods. The Focal software allows users to access XiO and other RTP systems from a remote PC-based system and provides the user with the ability to edit and modify patient data from a Windows-based PC and transfer the information back to the RTP system. | Computerized Medical Systems Inc | 2 | Focal: when a plan is sent to Focal from XiO and the structure properties are edited in Focal CT Sim activity; these changes result in changes to the MLC or field size for the plan when they should not. XiO will recalculate dose when modified plan is returned to XiO however the user may not notice that it has changed form the plant they had approved earlier. Editing the structure properties (col | Computerized Medical Systems Inc sent a Customer Advisory, dated 11/30/07, was sent to consignees December 2007 to notify them of the issue and how to avoid the workaround. The advisory listed the problem, clinical impact, and a workaround. The firm will be sending each site a reminder that the issues is resolved.
Z-2190-2011 | May-12-2011 | 2011 | ARTISTE / RT Therapist v4.1 system - ARTISTE / RT Therapist v4.1 system.The ONCOR and PRIMUS system may be updated to utilize the RT Therapist v4.1 software.Siemens Healthcare; Concord; CA.The intended use is to deliver X-ray photon and electron radiation for the therapeutic treatment of cancer. | Siemens Medical Solutions USA; Inc | 2 | When using the ARTISTE Linac and RT Therapist v4.1 along with an imported cropped reference image; the image maybe processed incorrectly; thus resulting in a mistreatment. | Distribution of the Urgent: Medical Device Correction Customer Advisory Notice began on March 24, 2011 via certified mail or by the Siemens Customer Service Group. The letter identified the issue and when it occurs. Users were asked to verify the isocenter position of the reference image in the Portal Review application before using a reference image imported from the treatment planning system. Siemens is currently working on a solution and will notify customers as soon as the update is available. The letter should be included with the Owner Manual of the system in the chapter for "Safety Advisory Letters" until an update is provided.
Z-3003-2011 | Aug-11-2011 | 2011 | Linear Accelerator - ARTISTE Linear Accelerator with RT Therapist v4.1 system.Siemens Medical Solutions; Kemnath; Germany.Distributed by Siemens Medical Solutions; Concord; CA 94520.Product Usage: Linear Accelerators used to deliver X-ray photon and electron radiation for therapeutic treatment of cancer. | Siemens Medical Solutions USA; Inc | 2 | A potential safety issue was discovered when a pause for an imaging segment is inserted (where an imaging segment does not exist) with the ARTISTE and RT Therapist 4.1 system will continue to the next image. | Siemens sent an "Urgent: Medical Device Correction/ Customer Safety Advisory Notice" letter to all affected customers starting June 3, 2011. The letter describes the product, problem, and instructions to prevent mistreatment. The letter states that Siemens is working on an updated version of RTT 4.1 that will fix the issues. Users will be notified as soon as the update for the system is available. The customers were instructed to include Customer Safety Advisory notice in their syngo RT Therapist 4.1 System Owner Manual chapter "Safety Advisory Letters" where it should remain.
Z-1871-2011 | Apr-15-2011 | 2011 | Toshiba Infinix-i - Toshiba Infinix-i; INFX-8000CThe expected usage of these devices is for fluoroscopy studies for various anatomical regions of the human body. | Toshiba American Medical Systems Inc | 2 | Under some conditions during high doses procedures the cumulative air kerma display may reach its maximum capability and fail to display the accrued value. | An Important: Electronic Product Radiation Warning notification was sent to all consignees. The letters described the reason for the recall and identified the affected product. The system software will be modified to allow the shift of the decimal to one position to the right allowing for a display of 99999.9 mGy. This will allow for the retention of the resolution until such time as the system exceeds the 9999.99 mGy range. A Toshiba Service person will contact customers when the software is available. Installed systems with the current version of software ( version 4.23) will begin to receive the corrected software April 2011. Previous versions of the software will begin to receive corrected software June 2011. The software will be provided free of charge. Questions related to the letter should be directed to the Regulatory Affairs Department at 800-421-1968.
Z-2964-2011 | Aug-09-2011 | 2011 | Lenstar LS 900 EyeSuite Biometry version i2.000; - Lenstar LS 900 EyeSuite Biometry version i2.000; Four (4) Protective Cardboard Containers Per Each System.Haag-Streit USA; Inc.A non-invasive; non-contact OLCR (Optical Low Coherence Reflectometer) Biometer used for obtaining ocular measurements and performing calculations to assist in determination of the appropriate power and type of IOL (Intraocular Lens) for implantation after removal of the natural crystalline lens following cataract removal. | Haag-Streit USA Inc | 2 | A patient record may be linked to an incorrect/inconsistent set of calculation results under specific circumstances when using EyeSuite Software version i2.000. | Haag-Streit USA, Inc. sent an "Urgent Field Safety Notices" dated June 20, 2011 to all affected customers. The letter included a description of the affected product and problem. The firm instructed their customers to make sure all new and current templates are saved and upgrade to EyeSuite version i4.000. Customers are asked to complete and return the Software Update Verification Form. For additional information contact the Haag-Streit at (513) 398-3937.
Z-2947-2011 | Aug-04-2011 | 2011 | BrightView X Gamma Camera; 3/8". - BrightView X Gamma Camera; 3/8".Philips Medical Systems; Inc.Intended to produce images depicting the anatomical distributions of single photon emitting radioisotopes within the human body for interpretation by medical personnel. | Philips Medical Systems | 2 | Non-conformances with BrightView X/XCT systems running version 2.0.1 including data transfer; XCT acquisition; Pixel Mask for image reconstruction and table moving incorrectly. | Philips Medical Systems, Inc. set a "Customer Information" letter dated July 8, 2011 to all affected customers. The letter informed them of an upcoming software updated (version 2.5). For questions please contact 800-722-9377 (option 5).
Z-3205-2011 | Sep-14-2011 | 2011 | Philips Healthcare PCR Eleva Radiological Image Processing System - PCR Eleva Radiological Image Processing System with software R.1.2 Philips Healthcare; Inc.A computed radiography image processing system. | Philips Healthcare Inc. | 2 | The mirror icon on the bottom of each image could be interpreted as a RIGHT marker for side indication if the image is mirrored again within a PACS system. This could potentially lead to a misdiagnosis | Philips Healthcare Inc. sent out "Urgent --Medical Device Correction" notices dated June 10, 2011 to all affected customers. The letter described the product and problem. Customers were informed the "burned-in Mirror-Icon" may be misinterpreted as a Right marker. Customers are advised to preferably use lead markers during each X-ray acquisition or alternatively electronic annotations on each image for side indication. The presentation of these markers within each image will clearly show the difference between the intended Left/Right indication and the little mirror ICON. Philips will install improved software that prevents the misinterpretation of the symbol and will schedule the installation of the updated software. For questions on this recall call Phillips Healthcare Inc. at 978-687-1501.
Z-2059-2011 | Apr-26-2011 | 2011 | PHILIPS MODEL MX 16-slice Computed Tomography Scanners - MX 16-slice Computed Tomography Scanners; equipped with Software Version 1.1.1. Model #E989605657371.The MX 16-slice CT Scanner System can be used as a whole body computed tomography X-ray system featuring a continuously rotating X-ray tube and detector array with multi-slice capability up to 16 slices simultaneously. The acquired X-RAY transmission data is reconstructed by computer into cross-sectional images of the body from the same axial plane taken at different angles. The system is suitable for all patients. | Philips Medical Systems (Cleveland) Inc | 2 | The Philips MX 16-slice CT scanners are subject to recall/field correction due to a problem encountered with the units when using the 'Combine Viewing' Option. The firm determined that inaccurate dimension measurements display when using the Combine Viewing Option. Moreover; if the dimension measurements from merged images created using this option are used; there is a potential for applying inac | Philips Healthcare issued an 'URGENT- Field Safety Notice MX 16-slice' (dated 3/10/2011) concerning a software issue to customers. The notification informed the firm's users/customers of the potential for inaccurate dimension measurements produced by the MX 16-slice CT scanners when using the Combine Function Viewing Option. Moreover, the letter notified customers of the potential for patient misdiagnosis in the event that inaccurate dimension measurements are used by a physician for reference. The Field Safety Notice described the following: 1) How to identify the affected product; 2) What actions should to be taken by the customers with regard to avoiding the issue -discontinue use of the Combine Viewing Option and instead use the MPR Batch for diagnostic reviews until a planned software upgrade is installed by the recalling firm; and 3) The firm's plans to carry out the installation of a field change order kits to correct the issue. The planned software upgrade is expected to be completed on or before 6 months after the new software is released. If you need any further information or support concerning this issue, please contact your local Philips representative or local Philips Healthcare office. For North America and Canada contact the Customer Care Solutions Center (1-800-722-9377, option 5: Enter Site ID or follow the prompts).
Z-0083-2012 | Oct-21-2011 | 2011 | Instrument Manager software - Instrument Manager -Software Versions 8.05; 8.06; 8.07; 8.08; and 8.09.It is a stand-alone software package that provides connectivity between laboratory instruments and Laboratory Information Systems; aiding the flow of laboratory information; including test orders and results; specimen information; and patient information. Instrument Manager also provides data management functions that aid in the interpretation and management of laboratory results. The Instrument Manager software is put on a CD which is distributed in a box. User documentation consists of user manuals; PDF copies of which are on the CD Intended to be used by trained laboratory personnel as an aid in laboratory instrument interfacing and laboratory data management. | Data Innovations; Inc. | 2 | The "previous result" retrieved from the Specimen Management Database may not be the most recent "previous result". | Data Innovations sent an "IMPORTANT PRODUCT NOTICE" dated May 11, 2009 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter offers temporary workarounds for customers to use until corrected versions of the software is complete. Versions of the corrections will be provided on a staggered schedule. Customers will be notified by e-mail. Questions regarding this notice are directed to Support at 802-658-1955.
Z-0213-2012 | Nov-18-2011 | 2011 | DICOM Option for the Norland Illuminatus Software Revs - DICOM Option for the Norland Illuminatus Software Revs 4.2.0 though 4.3.1. Models Norland XR-600; XR-800; XR-46; EXCELL; XR-36 and Eclipse DXA Bone Densitometers.The XR 46 performs DXA scans of the AP Spine; Hip; Forearm; Lateral Spine; and Whole Body; as well as other user selectable sites. It provides BMD (g/cm2); Area (cm2); and BMC (g) values. It compares these values to gender and ethnic matched reference populations and provides T-Score and % young Reference; Z- Score and % age matched; and long term and short term change values. This includes sBMD (mg/cm2). The XR 46 performs soft tissue assessment and provides lean mass; fat mass; percent fat; and total soft tissue values for all scan sites; including Whole Body. The XR 46 includes a Report writer that allows the operator to customize the appearance of their report stand to automate the tedious part of the report. It also allows the user to enter ranges of patient values and the specific statements they want to be printed on their report for each of these ranges. The bone density measurements from the XR 46 can be used as an aid to physicians in determining fracture risk. | Cooper Surgical; Inc. | 2 | The DICOM Software; Norland Illuminatus DXA Software with the DICOM Option (Part Number 434A133) version 4.2.0 through 4.3.1; which is installed on the XR-600; XR-800; XR-46 and EXCELL DXA Bone Densitometers; could lead to improper patient information getting into a previous patient's record. | The firm, CooperSurgical, sent a "Notice of RECALL-URGENT" letter dated September 23, 1011 to its customers. The letter described the product, "Problem/Situation", and "When can this problem occur?". CooperSurgical will send a new revision of Illuminatus software to the customers once available. If you have any concerns about your existing database of patients, please contact our Norland Technical Support team at (203) 601-5200 Extension 9888 or 9889.
Z-2167-2011 | May-10-2011 | 2011 | Philips - Philips NM3 Monitors; Model 7900PHILIPS NM3 Monitor; English - P/N 1051674PHILIPS NM3 Monitor; English Refurbished - P/N U1051674PHILIPS NM3 Monitor; English; Demo Unit - P/N DU1051674PHILIPS NM3 Monitor; Spanish - P/N 1060462NM3 is a multi-parameter patient monitor that utilizes the FloTrak Elite module to perform some of the calculations. | Respironics Novametrix; LLC. | 2 | Software: two displayed respiratory parameters are incorrect due to errors in the system software: MValv (alveolar minute ventilation) and Vt/kg (tidal volume/patient weight). In both instances; the displayed parameters are higher than actual. | Respironics issued Field Action Notification Letter via Fed'x on 1/4/11, notifying customers of the issue and to instruct them on conditions for continued use and return of the device. The device will be updated with revised software to correct the calculation errors. Any questions or need additional information about this recall action, Contact Respironics US Customer Service at 1-800-345-6443 for customers residing in the USA or +011 (203)697-6488 for International customers.
Z-1021-2011 | Feb-02-2011 | 2011 | Vital Connect - Vital Connect 4.0 & 4.1; radiological image processing softwareA Medical diagnostic software system intended to process; analyze; review; and distribute multi-dimensional digital images acquired from a variety of imaging devices including: CT MR; CR/DR/DX; SC; US; NM; PET; XA and RF; etc. Vital connect has the following intended use: Vessel Probe is intended for viewing the anatomy and pathology of a patient's coronary arteries. Clinicians select any artery to view the following anatomical references: the highlighted vessel in 3D; two rotate-able curved MPR vessel views displayed at angles orthogonal to each other; and cross sections of the vessel. | Vital Images; Inc. | 2 | There is a potential alignment error of fused PET/CT images in the Vitrea Core component of Vitrea Enterprise Suite 1.2 and 1.3. The error occurs during panning of the fused PET/CT images. Because the PET data has a larger voxel size than the CT data; it is possible to pan the CT image in more frequent; smaller steps than the PET image; so that the fused PET image does not move until the CT data | The firm, Vital Images, Inc,., sent a "Urgent Software Alert" letter dated August 23, 2010 to all customers. The letter described the product, problem and actions to be taken by the customers. The customers were instructed to not use the PET/CT fusion functionality until the software updates are provided and to re-examine any past results. If you have any questions or comments, contact Customer Support at support@vitalimages.com or at 1-800-208-3005.
Z-3122-2011 | Sep-01-2011 | 2011 | SIGNA 3.0T MR750 SYSTEM - GE Healthcare; 3.0T Signa MR750 (Marketing Name: GE Discovery MR750)Product Usage: The GE Sigma MR750 System is a whole body magnetic resonance scanner designed to support high resolution and high signal-to-noise ratio images in short exam times. It is indicated for use as a diagnostic imaging device to produce axial; sagittal; coronal; and oblique anatomical images; spectroscopic images; parametric maps; and/or dynamic images of the structures and organs of the entire body; including; but not limited to head; neck; TMJ; spine; breast; heart; abdomen; pelvis; joints; prostate; blood vessels; and musculoskeletal regions of the body. The images produced by the GE Sigma MR750 reflect the spatial distribution or molecular environment of nuclei exhibiting magnetic resonance. These images and spectra; when interpreted by a trained physician yield information that may assist in diagnosis. | GE Healthcare; LLC | 2 | GE Healthcare has become aware of two potential safety issues associated with MR products (Optima MR450w; Discovery MR450; Discovery MR750; Brivo MR355; Optima MR360) that may impact patient safety.1) Inline post-processed (i.e. automated post-processing) image locations are not correctly annotated under some conditions. 2) Reference images displayed in Functool FiberTrak do not match with | GE Healthcare sent an "Urgent Medical Device Correction" letter dated July 27, 2011 to customers with affected systems. The letter described the recalled product, safety issues and recommended actions to be taken. The letter provided immediate mitigation (Safety Instuctions) that address each issue. The letter states that the recalling firm will contact customers to arrange to provide a software upgrade to all affected systems at no cost. For any questions concerning this notification contact your local service representative.
Z-2243-2011 | May-23-2011 | 2011 | ARTISTE ; ONCOR and PRIMUS digital accelerators with syngo RT Therapist v4.1 and 4.2 system; - ARTISTE ; ONCOR and PRIMUS digital accelerators with syngo RT Therapist v4.1 and 4.2 system;Therapist Express Basic; syngo RT Therapist and syngo RT Therapist Connect/MOSAIQ OIS.The ONCOR and PRIMUS may be updated to utilize the RTT v4.x software.Distributed by Siemens; Healthcare; Concord; CA 94520Radiation therapy | Siemens Medical Solutions USA; Inc | 2 | Siemens found a potential safety issue related to patient treatment with table offsets on syngo RT Therapist (RTT) 4.1 and 4.2 which could result in dose being delivered to wrong location. | Siemens initiated an Urgent Medical Device Correction letter on March 31, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to: use only relataive table setup with all planned table values '0' if applicable. In cases where the treatment plan contains at least one table parameter other than '0' to consider the following: Use only one PVG plus resulting offset per fraction. Once an offset has been applied do not use 'apply offset' at RTT for any further offset. Select the PVG and treatment beams in the same auto-sequence. In case this is not possible, the table position for each treatment beam, in the fraction, needs to be verified separately (especially for treatments with beams having different table posiitons. If an in-session resumption or individual download of beams has to be performed, make sure that the treatment table is in the correct position. In order to prevent any unintended table movement select manual table movement for all directions in the syngo RT Therapist service UI. Scenario 2 - Manual change of table posiiton followed by a PVG calculated offset. Always deliver the PVG's at the planned table position. Avoid manual change of table posiiton prior to applicaiton of PVG. Select the PVG and treatment beams in the same auto-sequence. In case this is not possible, the table position for each treatment beam, in the fraction, need to be verified separately (especially for treatments with beams having different table positions). If an in-session resumption or individual download of beams has to be performed, make sure that the treatment table is in the correct position. In order to prevent any unintended table movement select manual table movement for all directions in the syngo RT Therapist service UI. Customers were informed that a software update would be provided as soon as possible. Customers were asked to include the Customer Sa.
Z-3195-2011 | Sep-13-2011 | 2011 | TSX-301B: AQUILION Premium; System; X-Ray; Tomography; Computed - TSX-301B: AQUILION Premium; System; X-Ray; Tomography; ComputedX-Ray Tomography | Toshiba American Medical Systems Inc | 2 | Toshiba America Medical Systems is initiating a field correction on their TSX-301B: AQUILION Premium; System; X-Ray; Tomography; Computed;TSX-301A: AQUILION One; System; X-Ray; Tomography; Computed;TSX-20IA: AQUlLlON LB; System; X-Ray; Tomography; Computed; and TSX-1O1A: AQUILION 16; System; X-Ray; Tomography; Computed due to several software issues that cause the scanners to function incorr | Toshiba America Medical Systems (TAMS) sent an URGENT MEDICAL DEVICE CORRECTION letter dated June 22, 2011, to all affected customers. Customers were instructed that a revised software to prevent occurrence of the problem would be installed on their system. Their Toshiba Service Representative would contact them for an appointment to install the software to their system when it became available. Customers were requested to take temporary measures until corrective measures were taken. 1) In dual scanoscopy, do not press the short key between the completion of the first scanoscopy and the start of the second scanoscopy. 2) In dual scanoscopy, do not click the (Scano Skip) button immediately before the completion of the first scanoscopy. 3) When the SplitCase function is set to ON, do not select two or more orders simultaneously from the appointment list. 4) When an image comment is to be preset by making a selection from the preset memo, do not select any image comment with more than 10 characters. Customers were asked to share the information with all users and reviewing radiologists as well as clinical engineering or Biomedical group at the facility. Customers were asked to complete and return the attached form and fax it to 877-349-3054. For any questions call (800) 421-1968 or contact your local Toshiba Representative at (800) 521-1968.
Z-2562-2011 | Jun-14-2011 | 2011 | Focal Radiation Treatment Planning Workstation - Focal Radiation Treatment Planning Workstation. Focal Release 4.51.00 and above.Allows users to access XiO and other RTP systems from a remote PC-based system and provides the user with the ability to edit and modify patient data form a Windows-based PC and transfer the information back to the RTP system. It is intended for use in radiation treatment planning using generally accepted contouring methods. | Computerized Medical Systems Inc | 2 | Focal: If a plan involving Dynamic MLC Delivery is sent from XiO to Focal and called up there in Plan Review; the angle of the gantry is set to zero for all beams. If this plan is then sent from Focal to the R&V System; the gantry angles all remain at zero which could lead to patient mistreatment. | Elekta CMS Software sent a USER NOTICE dated November 5, 2010, to all affected customers. The notice identified the product, the problem, and the action to be taken by the customer. A return postcard was included for the customer to confirm receipt of the User Notice. The firm intends to correct the software so it does not modify anything about the plan sent from XiO. Customers were instructed to contact their local customer support office with any questions.
Z-2147-2011 | May-06-2011 | 2011 | ACUSON X300 - Siemens ACUSON X300 Systems with software versions 3.0.00 and 3.0.02; model numbers 10348531; 10348532 and 10348533; manufactured by Siemens Medical Solutions; Mountain View; CA.Intended for General Radiology; Fetal; Abdominal; Intraoperative; Pediatric; Small Parts; Neonatal/Adult Cephalic; Cardiac; Transesophageal; Pelvic; Transcranial; OB/GYN; Urology; Vascular; Musculoskeletal; Superficial Musculoskeletal; and Peripheral Vascular applications. | Siemens Medical Solutions USA; Inc. | 2 | Mistreatment risk-- when a user defined formula is used to calculate a value in a vascular exam; data is switched laterally when it is internally transferred and displayed on the summary report within the system. | A field correction was initiated and a Customer Safety Advisory Notification was sent out on 7/29/09. The letter identified the affected product and stated that there was a potential for user defined calculations to be incorrect within the vascular package when using the listed software versions of the product. The letter also discussed the steps that customers could take to avoid the potential risk of the issue. A mandatory update software was developed and will be installed free of charge. Customers should have been contacted to schedule a time for the installation. Questions were to be directed to the local service support person.
Z-1064-2011 | Feb-08-2011 | 2011 | Linear Medical Accelerator ARTISTE MV system - Linear Medical Accelerator ARTISTE MV systems; when equipped with a syngo RT Therapist; Therapist Express Basic; syngo RT Therapist Assist or Therapist Express Assist workspace with software version NET VA10A for RTT 4.1.122 or NET V10B for RTT 4.1.125; manufactured by Siemens Medical Solutions; Concord; CA.Intended use: medical charged-particle radiation therapy system | Siemens Medical Solutions USA; Inc | 2 | Potential for Mistreatment-- when treatment plans intended for non-ARTISTE machinese are loaded; if the plans contain unspecified beam parameters; attributes or unknown block codes; the therapy system may remove those parameters. | The firm,Siemens, sent two Customer Advisory letters (TH004/09/S and TH008/09/S) contained in the Update Instructions to all consignees/customers with active devices. The letters describe the product, problem and action to be taken by the customers. The customers were instructed to follow the updated instructions. Siemens service representative will sign a document noted as CORRECTED upon completion of the software installation. If you have any questions, call 1-925-602-8157.
Z-0209-2012 | Nov-17-2011 | 2011 | SoftPath GUI - SoftPath GUI versions: 3.17.5 to 3.17.9; 4.2.1 to 4.2.3; and 4.3.0 to 4.3.0.4. SoftPath is a laboratory information system to be used in a clinical laboratory by knowledgeable; trained; and experienced personnel. SoftPath enables the user to document steps in the receiving and processing of surgical; gynecological; medical cytological; and autopsy case specimens within the Pathology labs; for the purpose of generating Pathology reports for requesting physicians and label creation for identification of specimen/case material. | SCC Soft Computer | 2 | Soft Computer; Clearwater; FL; initiated a correction on the following SoftPath GUI versions: 3.17.5 to 3.17.9; 4.2.1 to 4.2.3; and 4.3.0 to 4.3.0.4. SoftPath architecture does not correctly handle AP macros and data fields inside report sections that are stored in the database. Improper information can be displayed in all places where the diagnosis from the database is read and displayed. | The firm, SCC Soft Computer, sent a "CORRECTION COMMUNICATION" notice on July 13, 2011 to its customers. The notice describes the product, problem and actions to be taken. The customers were instructed to restrict security access to the template set-up files and to acknowledge receipt of the notice; grant SCC personnel permission to check their report templates and correct any that are affected; indicate if they wish to schedule the utility to be run and if they would like to have the SoftWeb parameter PathPDFLinkOnly set to "Y". SCC will review the templates on the customers systems; discuss any irregularities requiring change and move or remove AP macros and data elements as needed. If you have any questions, contact the Director of Quality Management at (727)789-0100.
Z-1999-2011 | Apr-15-2011 | 2011 | GemStar SP Infusion Suite Database; Version 1.0 - GemStar SP Infusion Suite Database; Version 1.0; an Rx PC software application; Hospira; Inc.; Lake Forest; IL 60045; list 13092-04-01A stand-alone PC application for administration of medications specific to the GemStar Infuser | Hospira Inc. | 2 | The software enables programming of the GemStar device for a Bolus Lower Lockout Limit outside of the hard limits entered into the GemStar SP Infusion Suite Database PC application. | The firm, Hospira, sent "URGENT DEVICE FIELD CORRECTION" letters dated March 23, 2011 to their customers on the same date. The letter described the product, problem and preventative action to be taken by the customers until the software can be upgraded. The customers were requested to contact Hospira Global Product Safety and Complaints at 1-800-441-4100 to report adverse events, and Hospira Advanced Knowledge Center at 1-800-241-4002 for technical assistance. The accounts were also requested to complete the attached Reply Form and return it to Hospira via fax at 1-262-577-6921. For additional information and technical assistance, contact Hospira Advanced knowledge Center at 1-800-241-4002, option 4 (available 24 hours a day/7 days per week).
Z-2598-2011 | Jun-20-2011 | 2011 | Monaco Radiation Treatment Planning Workstation - Monaco Radiation Treatment Planning Workstation; Monaco Release 2.03.00 and aboveTo create treatment plans for any cancer patient for whom external beam intensity modulated radiation therapy (IMRT) has been prescribed. | Computerized Medical Systems Inc | 2 | Monaco: When the isocenter location is edited in the IMRT Activity screen; the shift coordinates are not updated in the IMRT Plan report. | Computerized Medical Systems Inc drafted a User Notice which explains the reason for correction, the clinical impact, and the workaround method. The User Notice was sent January 13, 2011. A return postcard is included for the customer to confirm receipt of the User Notice. The firm intends for this issue to be corrected in Monaco Release 3.00 in March 2011.
Z-0311-2012 | Nov-30-2011 | 2011 | Roche Cobas c111 Analyzer - Roche Cobas c 111 AnalyzerThe Roche Cobas c 111 analyzer is an in-vitro diagnostic analyzer capable of performing clinical chemistry; specific protein and electrolyte tests. Analytes are measured photometrically or turbidimetrically; the analyzer also has an optional ISE module for measuring sodium; potassium and chloride. | Roche Diagnostics Operations; Inc. | 2 | Internal investigations have revealed a potential printer malfunction on the Cobas c 111 analyzer; where one or more numbers or characters can be missing on the hard copy result printout; or the printer may skip leaving blanks which may result in an incorrect print layout. The missing number of characters can be in any position; therefore it could be in the middle of a result; (e.g. 48.53 mmol in | Roche Diagnostics Operations, Inc sent a Urgent Medical Device Correction dated October 18, 2011, via UPS Ground (receipt required) to all affected customers and one distribution center, with instructions to disable the Auto Print Result feature until new software is available. All printed results will need to be confirmed against the results indicated on the display screen. Questions are to be directed to Roche Diagnostics Technical Support at 1-800 428-2336,.
Z-2113-2011 | Apr-29-2011 | 2011 | DOCTORS REVIEW SYSTEM - Doctors Review System; OB-Windows/OB-View software release numbers 4.1 through 4.1.3; Digisonics; Inc.; 3701 Kirby Drive; Houston; TX 77098OB-Windows provides professional reporting and analysis of fetal growth throughout the gestational period; which can help evaluate clinical situations. OB-View provides image management for the user to capture; view; read; and edit images. | Digisonics; Inc | 2 | Study would on occasion have fetal heart rate and/or cervical length on the report that was not entered nor did it belong to that patient. Software displayed incorrect data for fetal heart rate and cervical length in OB ultrasound reports. | The firm, Digisonics Inc, notified the affected customers of software release via e-mail on January 24, 2007 and follow-up e-mail on February 7, 2007. The e-mails described the product, problem and the actions taken. The firm was able to trap the error and after this report the developer was able to determine the cause and a new executable was sent and installed at the customers site. If you have any question, please call 800-940-3240 or e-mail: service@digison.net.
Z-0325-2012 | Dec-01-2011 | 2011 | Brachytherapy System; Remote Afterloader - VariSource iX; Software Version 1.1; Part Number AL20020000 VariSource iX Afterloader.The VariSource iX Series High Dose Rate Afterloader system is a computer controlled remote electra/mechanical system used for medical purposes; for placing a NiTiNol wire incorporating an irradiated iridium seed internally or close by; a malignant tumor or tumor bed in a practice known as brachytherapy. The device has up to 20 channels for treatment. | Varian Medical Systems; Inc. | 2 | VariSource iX console software may cause incorrect dwell times to be delivered to patient and may result in a mistreatment. | Varian Medical Systems sent a "MEDICAL DEVICE RECALL" letter dated April 1, 2010 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. Customers were asked to conduct a simple test to establish if their machine is affected by the problem. Device users were requested to complete an attached receipt verification form and return it to Varian BrachyTherapy via fax or e-mail. Call the Contact Service Dispatch number at 1-800-864-1672 for questions regarding this detail.
Z-2977-2011 | Aug-10-2011 | 2011 | ARTISTE/RT Therapist v4.2 and/or Oncologist v4.2 systems - ARTISTE/RT Therapist v4.2 and/or Oncologist v4.2 systems.Manufactured by Siemens AG; Medical Solutions; Kemnath; Germany.Delivery of X-ray photon and electron radiation for the therapeutic treatment of cancer. | Siemens Medical Solutions USA; Inc | 2 | A potential safety rick related to transferring reference images for offset calculation between a syngo RT Therapist v 4.2 or a syngo RT Oncologist 4.2 and other versions of the COHERENCE/syn go RT Therapist or Oncologist or vice versa. An incorrect offset may be applied for patient positioning and result in dose being delivered to the wrong location. | Siemens sent an "URGENT: MEDICAL DEVICE CORRECTION/CUSTOMER SAFETY ADVISORY NOTICE" letter dated June 3, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. The letter provides preventive measure instructions for customers to use until an update for the system is available.
Z-1434-2011 | Feb-24-2011 | 2011 | CRS - Cardiovascular Review StationThe IMPAX Cardiovascular suite is a cardiovascular information system; providing image archiving; image display and modality/study specific structured reporting. | AGFA Corp. | 2 | Echocardiography volume measurements performed within CRS were not reflected correctly in associated Results Management (RM) reports. | The firm, AGFA Healthcare, sent a "Safety Notice" dated January 27, 2011 to all customers. The notice described the product, problem and action to be taken by firm and customer. AGFA informed the customers that a software upgrade was made between December 2008 and February 2010 and instructed the customers to distribute this information to all individuals within their facility and to complete and return the attached URGENT SAFETY NOTICE FEEDBACK FORM via fax to 864-421-1664 or email to debbie.norris@agfa.com. Note: Service Bulletin DD+DIS100.09E, was issued March 2009 and an upgrade from Software Versions 2.12.05 through 2.14.03.SU1 to Software Version CRS 2.14.03.SU2 was required for the impacted sites. Most identified consignees/customers required the software upgrade, but some consignees/customers had already been upgraded prior to the event and did not require additional upgrades. If you have any questions about this matter, please contact the Post Market/QARA North America via email at debbie.norris@agfa.com or call at 864-421-1754.
Z-1960-2011 | Apr-12-2011 | 2011 | SYSTEM 1E liquid chemical sterilant processing system - STERIS SYSTEM 1E Liquid Chemical Sterilant Processing SystemThe SYSTEM 1 E Liquid Chemical Sterilant Processing System is intended for the liquid chemical sterilization of manually cleaned impressible; reusable; critical and semi-critical heat sensitive medical devices; including endoscopes and their accessories. Devices processed in a SYSTEM IE have been chemically sterilized using S40 Sterilant Concentrate; a peracetic acid liquid chemical sterilant and rinsed with extensively treated; potable water. | Steris Corporation Hopkins Facility | 2 | The printout; display and Operator Manual for the STERIS System 1E Liquid Chemical Sterilant Processing System are subject to recall/correction by the firm because in the event of an incomplete or aborted cycle; the language of the warning messages used in the printout; display; and Operator Manual omits the qualifying term: *liquid chemically* (not **sterilized); when alerting the user to inco | The firm, STERIS, Issued an 'Urgent Field Correction Notice' dated March 14, 2011 addressed to their Hospital Administrator, or Manager, and Risk Manager customers on March 18, 2011. The notification letter describes the product subject to field correction and the problem that exists with the language omission present in the printout display and Operator Manual messages which warn of an incomplete or aborted processing cycle. The notification goes on to state that there is no health hazard involved in this error, as the users are still made aware that the processing cycle has not been successfully completed. The notification goes on to remind the user/customers that in case of any incomplete, cancelled or aborted cycle, the users are instructed to process devices in the cycle following the directions contained in the Operator Manual. Lastly, the notification letter informs the user/customers that the company will be installing new software on the recalled systems in order to revise the language on the cycle printout and the display screen warning of a disrupted cycle. It also states that a STERIS Service Representative will be contacting each facility in order to schedule installation of the new software. The notification also provides the customers with a 1-800 telephone number for STERIS Field Service Dispatch, to be used if the customers have questions or want to request additional information. It also provides a local telephone no. 1-440-392-7455, for obtaining additional information from the firm's Director, Div. of Low Temperature Sterilization at: 1-440-392-7455.
Z-1023-2011 | Feb-02-2011 | 2011 | CyberKnife Robotic Radiosurgery System Device - Radiation therapy device. Medical charged particle radiation therapy system with software version. Model Numbers: Multi Plan Treatment Planning System; Software version 3.0; 3.1; 3.5.0; and 3.5.1.For the treatment planning and image guided radiosurgery and precision radiotherapy of lesions; tumors and conditions anywhere in the body when radiation treatment is indicated. | Accuray Inc | 2 | When a user modified the align center of an existing plan during the treatment planning process the effective depths of each beam in the plan were not updated by the system. These depths were incorrect and led to an incorrect dose calculation. Software patch /update to address this issue. Implementation at affected sites will initiate; once the patch is available. | The firm, Accuray, sent an "URGENT DEVICE CORRECTION" letter dated January 18, 2010, to all customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed to load the plan and recalculate the dose to reflect the current state of beam, compare the reloaded plan and the recalculated plan to verify maximum dose and prescription values, and re-prescribe and save to overwrite to existing plan if necessary. The firm is developing a patch/software update to address the issue and a Accuray Customer Support person will contact the customer regarding further resolutions. If you have any questions or concerns regarding this issue, contact Accuray Customer Support at +1-877-668-8667(USA) or +1-408-716-4700 (non USA) or email: customersupport@accuray.com.
Z-0363-2012 | Dec-02-2011 | 2011 | SoftPath GUI - SoftPath GUI versions: 4.2.3.3-4.2.3.6 and 4.3.0.5-4.3.0.12.SoftPath is a laboratory information system to be used in a clinical laboratory by knowledgeable; trained; and experienced personnel. SoftPath enables the user to document steps in the receiving and processing of surgical; gynecological; medical cytological; and autopsy case specimens within the Pathology labs; for the purpose of generating Pathology reports for requesting physicians and label creation for identification of specimen/case material. | SCC Soft Computer | 2 | On 07/20/2011 SCC Soft Computer; Clearwater; FL; initiated a correction on the following SoftPath GUI versions: 4.2.3.3-4.2.3.6 and 4.3.0.5-4.3.0.12. A client reported that the HPV results on a patient report were incorrect. The HPV results were from the previous order on the patient and not the current order. All affected clients were notified about the issue on 07/20/2011. | SCC Soft Computer sent a "CORRECTION COMMUNICATION" dated July 20, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. SCC personnel will review the customized Lab Results AP macro and/or the Micro Results AP macro on customer systems and will either discuss with them the irregularities requiring change or reconfigure macros as necessary.
Z-1392-2011 | Feb-22-2011 | 2011 | V5M TEE Transducer - V5M TEE Transducer; when used with an ACUSON S2000 ultrasound system that has cardiac options installed and is at software version VA10; VA10A; VA10B; VA10C; VA15; VA15A; VA15B; VA16 or VA16A; manufactured by Siemens Medical Solutions; Mountain View; CA.Multipurpose mobile software controlled diagnostic ultrasound system and on-screen display for thermal and mechanical indices related to bio-effect mechanisms. Its function is to acquire primary or secondary ultrasound echo data. The V5Ms is a multi frequency; multiplane; phased sector array transducer. It can be introduced into the esophagus to obtain images of the heart structure. | Siemens Medical Solutions USA; Inc. | 2 | If the transducer is not disconnected from the system after an automatic shutdown due to thermal limits being exceeded; then the next time it is used; the thermal safety guards do not take effect. Normally; the thermal limit before automatic image freeze (and therefore power shut off) occurs at 43 degrees Centigrade. | Siemens sent a Customer Safety Advisory Notification letter on August 28, 2009, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed that their local Customer Service Engineer would install a software update as soon as it becomes available. Customers were instructed to pass the Customer Safety Advisory Notification letter to all who need to be aware of these issues until the corrective action is completed. For questions regarding this recall call 650-969-9112 or 800-422-8766.
Z-1928-2011 | Apr-08-2011 | 2011 | Hamilton-G5 Ventilator - Hamilton-G5 Ventilator with software version V2.00x and Neonatal Option or V2.01x and Neonatal software. Hamilton-G5: catalog number 159001;Neonatal Software catalog number: 159700; Neonatal Option catalog number: 159187. Hamilton MedicalIntended for use in the hospital and institutional environment where healthcare professionals provide patient care. | Hamilton Medical Inc | 2 | The adaptive volume/pressure regulation can be automatically and permanently reduced to PEEP + 3 mbar when a closed suctioning maneuver is performed on the patient and may be undetected by the operator. | A Medical device Safety Alert and Corrective Action along with a cover letter, was sent to all locations on February 16, 2011. The letter identified the affected product along with the reason for recall. The letter states that operators deactivate the "high tidal volume" alarm as described in appendix K.4 of the operator's manual until the ventilator is upgraded with a corrected software. Distributors will upgrade the affected devices with the corrected software as soon as possible. Questions can be directed to 800-426-6331, ext 215.
Z-2246-2011 | May-24-2011 | 2011 | IMPAX 6.5 - IMPAX 6.5 Client Software.IMPAX 6.5 combines the traditional activities of RIS informatics management with PACS image management to provide a powerful platform for imaging-based planning; interpretation; and results distribution. The IMPAX 6.5 Client focuses on the integration of PACS; RIS and Speech applications into a single delivery of information. A single IMPAX 6.5 Client application can be used by a range of users on any appropriate; networked workstation they have access to. | AGFA Corp. | 2 | Software issue that occurs after dictation of a study. | Agfa Corp. sent an "URGENT SAFETY NOTICE" customer letter via Agfa email to customers on 4/14/2011, along with a "URGENT SAFETY NOTICE FEEDBACK FORM," The letter identified the product , the problem, and the action to be taken by the customer. It was discovered a software issue in the IMPAX 6.5. The upgrade has already been released to address the software problem. Customers were instructed to fax back the feedback form as soon as possible to (864) 421-1664. If you have any questions, please contact your local Agfa Healthcare organization.
Z-3276-2011 | Sep-23-2011 | 2011 | Logix Order Entry (OE) Software - LOGIX Order Entry (OE) Software; all versions; catalog 2M8410; Baxter Healthcare; Deerfield; IL 60015; The software is included in a number of Baxter's pharmacy compounders. The software assists the pharmacist in management of compounding; creating and authorizing nutritional solutions intended for IV administration.Product Usage: The LOGIX OE application is an order entry system intended for use with Baxters LOGIX CM software. The software program aids the pharmacist in creating and authorizing nutritional solution orders. | Baxter Healthcare Corp. | 2 | There is the potential for an order entry error during use of the LOGIX Order Entry (OE) software. When adding ingredients to a TPN order; the user can accept the pre-populated "per units" of "Bag" or select other "per units". A dosing error may occur if the user does not select the prescribed "per units" for the added ingredient. | Baxter Healthcare Corp. sent an "Urgent Device Correction" letters dated August 15, 2011 which included LOGIX OE screen illustrations of a representative neonate and adult patient, via first class mail to all affected customers. The letter identified the product, problem and actions to be taken. The customers were instructed to follow safe compounding practice, which includes verifying that the ingredient per units entered match the TPN prescription prior to preparation. Baxter requests that customers communicate this issue to all personnel responsible for LOGIX order entry and order authorization at their facility. Customers are instructed to retain a copy of the letter at each computer on which LOGIX OE software is installed. The customers were also requested to complete the attached customer reply form and fax it to Baxter at the number provided on the form. Any questions regarding the communication were directed to Baxter's Medical Information Services at 1-800-422-2751.
Z-2561-2011 | Jun-14-2011 | 2011 | Monaco Radiation Treatment Planning Workstation. Monaco Release 2.03.00 and above. - Monaco Radiation Treatment Planning Workstation. Monaco Release 2.03.00 and above. The Monaco RTP system is used to create treatment plans for any cancer patient for whom external beam intensity modulated radiation therapy (IMRT) has been prescribed. | Computerized Medical Systems Inc | 2 | Monaco: For VMAT plans; the measured dose of radiation delivered is lower than the software calculated dose by about 5% in Monaco Release 2.03.00 and 2.04.00. | Computerized Medical Systems, Inc. sent a "USER NOTICE" on January 13, 2011, to all affected users. This notice identified the product, the problem, and the action needed to be taken by the customer. A return postcard is included for the customer to confirm receipt of the user notification. The firm intends for correction in the dose calulation engine to appear in the Monaco Release 3.00 in March 2011. Any further questions please call (403) 830-8023.
Z-0005-2012 | Oct-04-2011 | 2011 | Cyberonics - VNS Therapy Demipulse Model 103 Generator Cyberonics; Inc. Houston; Texas; Model 103 Generators are provided in a sterile single use packages as part of the VNS Therapy System. Product labeling may be found at http://us.cyberonics.com/en/vns-therapy-for-epilepsy/healthcare-professionals/vns-therapy/manuals-page.Product Usage: The VNS Therapy System is indicated for use as an adjunctive therapy in reducing the frequency of seizures in adults and adolescents over 12 years of age with partial onset seizures; which are refractory to antiepileptic medications and for the adjunctive long-term treatment of chronic or recurrent depression for patients 18 years of age or older who are experiencing a major depressive episode and have not had an adequate response to four or more adequate antidepressant treatments. | Cyberonics; Inc | 2 | An investigation was initiated based on a report from the field in which an Intensive Follow-up Indicator (IFI) message was unexpectedly received by a medical professional when using Model 250 version 8.0 software to interrogate a patient's Model 103 Generator. | The firm, Cyberonics Inc., initiated containment actions on March 17, 2011 to prevent additional affected devices from being shipped. The firm initiated contact with implanting facilities in the United States on May 6, 2011 and international facilities and distributors on May 10, 2011 to remove product from the field that has not yet been implanted. Additionally, a "Field Safety Alert" letter was sent to all affected medical professionals. The letter identifies the affected product, the issue, and the action(s) to be taken by the physicians. Consignees are instructed to complete and return an attached Effectiveness Check Form. If assistance is needed physicians may contact Clinical Technical Support at 866-882-8804 or via email at clinicaltechnicalservices@cyberonics.com.
Z-1474-2011 | Mar-02-2011 | 2011 | MicroSpeed Uni Motor System - Aesculap MicroSpeed Uni Control Box; Part #GD670Intended for high speed cutting; sawing; drilling and manipulation of soft tissue and bone in the fields of Spine; ENT; Neuro and Maxillo facial Surgery. | Aesculap Inc | 2 | Potential safety issue in that the control box can misinterpret the information sent from the low speed motor if the control magnet inside the motor changes its posthion. | AESCULAP sent an IMPORTANT CORRECTION NOTICE letter dated November 22, 2010, to all affected consignees. The letter identified the product, the problem, and the action the consignee should take. Consignees were instructed to not use the motors for thread cutting and removal of screws. If any unusual reactions of any cutting tools were observed consignees should check to ensure that the motor is not running in the opposite direction. If any error message is displayed, consigness were instructed to not use the motor. Consignees were also instructed to check both directions of running before use and do not use a motor which produces an error message in one direction. Aesculap would be providing all customers with a software update to the MicroSpeed Uni control box either by the unit being returned to Aesculap for update or by firm representative doing the upgrade on site at their facility. For questions regarding this recall call 800-258-1946 x5254 or 800-258-1946 x5267.
Z-2781-2011 | Jul-11-2011 | 2011 | DICOM Query/Retrieve; version 1.1.4 - DICOM Query/Retrieve; version 1.1.4. Individual part # 9733808; or as part of a system order; DICOM Q/R Software Kit; part # 9733933. DICOM Query/Retrieve software is installed on Medtronic StealthStation systems.Product Usage: DICOM Query/Retrieve software enables the user to query a hospital Picture Archiving and Communication System (PACS) for patient DICOM scans. The user can then transfer the identified scan to a Medtronic StealthStation system. | Medtronic Navigation; Inc. | 2 | Software allows incorrect patient or patient history to be displayed. | Medtronic Navigation sent a "FIELD SAFETY NOTICE" letter dated June 2, 2011 to all affected customers. The letter identified the product, problem and description of the corrective action. The letter states that Medtronic field representatives will visit each consumer/user and upgrade the version of DICOM Q/R to 1.1.5. Further information is available by contacting Medtronic Navigation Technical Services at 1-800-595-9709.
Z-3214-2011 | Sep-15-2011 | 2011 | 8K Cardioplegia Monitor - Sarns 8000 Cardioplegia Monitor; Catalog Number 16414; also known as PROG COMP BD AMPRO 420 CPG 8K; Catalog Number 816373The Sarns 8000 Cardioplegia Monitor is indicated to measure extracorporeal line pressure and signal when the pressure exceeds the operator-set limit. | Terumo Cardiovascular Systems Corporation | 2 | The Cardioplegia Monitor for the Sarns Modular Perfusion System 8000 may fail to track cardioplegia delivery volume. | Terumo Cardiovascular Systems sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated August 29, 2011 to all affected customers. The letter describes the product, problem, and action to be taken. The letter advises customers to make sure all users are aware of the notice. Customers are instructed to complete and return an attached Customer Response Form. Terumo will replace the software on all monitors of the affected product. Customers with questions or concerns should call Terumo CVS Customer Service 1-800-521-2818.
Z-2291-2011 | May-25-2011 | 2011 | Logix Order Entry (OE) Software - LOGIX Order Entry (OE) Software; all versions; catalog 2M8410; Baxter Healthcare; Deerfield; IL 60015; The software is included in a number of Baxter's pharmacy compounders. The software assists the pharmacist in management of compounding; creating and authorizing nutritional solutions intended for IV administration. | Baxter Healthcare Corp. | 2 | The LOGIX OE automated backup does not retain the three most recent backup files and instead will only retain the last three backups of the previous year. The software stores backup files in numeric order rather than chronological date order. | Baxter Healthcare Corp. sent an URGENT DEVICE CORRECTION letter dated March 29, 2011, via first class mail to all affected customers. The letter identified the product, the problem and the action to be taken by the customer. It was discovered there is an issue with the LOGIX OE automated backup not retaining the three most recent backup files and instead will only retain the last three backups of the previous year. The accounts were requested to turn off the automated backup function on the LOGIX OE Software, and perform manual backups following the instructions included in the Urgent Device Correction letter, and retain a copy of the letter at each computer on which LOGIX OE Server is installed. The customer were also requested to complete the attached customer reply form and fax it to Baxter at the number provided on the form. Any questions regarding the communication were directed to Baxter's Medical Information Services at 1-800-422-2751.
Z-2733-2011 | Jul-05-2011 | 2011 | SJM Confirm; Implantable Cardiac Monitor (ICM) - SJM Confirm; Implantable Cardiac Monitor (ICM) Models: DM2100; DM2102 (IDE).This product is a lead less non-therapeutic implantable cardiac monitor and is implanted to aid clinicians in diagnosing intermittent symptoms which may be due to arrhythmia | St Jude Medical CRMD | 2 | The recall was initiated because St. Jude Medical made new software available for the SJM Confirm DM2100 implantable cardiac monitor (ICM) that incorporates enhancements to the devices sensing algorithm. These enhancements are intended to provide improved sensitivity and specificity for diagnosing arrhythmias. The user is presented with the option to upgrade new software into the SJM Confirm ICM | St. Jude Cardiac Rhythm Management Division sent an "Important Product Correction" letter on March 25, 2011 to all affected customers. The letter explains the reason for the recall and provided recommendations including a software upgrade by SJM Field Clinical Engineer or Technical Services. Customers were recommended to fill out and return a "Physician Acknowledgement Form". For additional information please contact your local St.Jude Medical Sales Representative, Field Clinical Engineer or the Technical Services Department at 800-722-3774.
Z-2789-2011 | Jul-11-2011 | 2011 | ONCOR/PRIMUS Linear Accelerator systems with RTT v2.x AND an ARTISTE Linear Accelerator/RT Therapist - ONCOR/PRIMUS Linear Accelerator systems with RTT v2.x AND an ARTISTE Linear Accelerator/RT Therapist v4.1 system with the LANTIS OIS.The Linear Accelerators are used to deliver X-ray photon and electron radiation for the therapeutic treatment of cancer. | Siemens Medical Solutions USA; Inc | 2 | Potential safety issue of possible data loss for the transfer of the same patients treatment data between the LANTIS OIS and the RTT v2.x software and the RT Therapist version 4.1 system | Siemens Customers Service group began providing Update Instructions and installing the software update on affected machines beginning May 25, 2011. A customer information letter was also provided.
Z-1849-2011 | Mar-30-2011 | 2011 | HEMOCUE Hb 201 DM ANALYZER - HEMOCUE AB Hb 201 DM ANALYZER; in vitro diagnostic device for quantitative hemoglobin determination. | Hemo Cue; Inc. | 2 | Incorrect configurations settings in software. As a result; certain error codes do not trigger appropriately. These error codes are: E01; E02 and E05 which are all related to the optronic performance of the analyzer. Possible incorrect results. | Hemocue AB sent Customer Notification "Notice of Medical Device Correction on HemoCue Glucose 201 DM and HemoCue Hb 201 DM Analyzers" dated October 7, 2010, to customers notifying them of incorrect settings on error codes. The corrective action consists of an update of the configuration settings on the the affected HemoCue Glucose 201 DM and Hb 201 DM analyzers that are in the field. The affected units in the field will be corrected in accordance with a rework procedure that has been established by HemoCue AB.
Z-0320-2012 | Dec-01-2011 | 2011 | XVI R4.5 X-ray Volume Imaging System - Elekta Synergy XVI R.4.5 and R4.6ProducT Usage: To be used as part of radiation therapy treatment process. | Elekta; Inc. | 2 | The software did not operate as expected. | Elekta sent an IMPORTANT NOTICE letter dated October 14, 2011 to all affected consignees. The letter identified the affected product, the problem, and the actions to be taken. This notice contained precautionary measures that are required for the operation of the equipment. The notice advise all customers to follow any instructions or recommendations covered in the Notice. The notice instructs Users to file this document in the Important Notice section of the appropriate User Manual. For questions contact your local Elekta representative.
Z-2577-2011 | Jun-16-2011 | 2011 | XiO Radiation Treatment Planning System - XiO Radiation Treatment Planning System; XiO Release 4.0.0 through 4.40.00Used to create treatment plans for any cancer patient for who external beam radiation therapy or brachytherapy has been prescribed. | Computerized Medical Systems Inc | 2 | XiO: When the patient is positioned with their feet toward the CT gantry; the coordinates exported to the LAP IsoMark laser positioning system are not correct. | The firm, CMS, Inc., sent a "XiO Customer Advisory" letter dated April 8, 2008 to all affected customers. The letter describes the product, the problem, and the actions to be taken ( workaround). The letter also states the problem will be resolved in a future release. No release number or date was given. If you have any questions, please call 408-380-8023.
Z-2193-2011 | May-13-2011 | 2011 | Digisonics; Digiview and ERS software releases 3.6.44; 3.6.4.6; and 3.6.5 - Digisonics; Digiview and ERS software releases 3.6.44; 3.6.4.6; and 3.6.5; Digisonics; Inc.; 3701 Kirby Drive; Houston; TX 77098; Tel: (800) 940-3240; Fax: (713) 529-7999Software used to make on-screen (2-D; M-Mode and Doppler) measurements of left ventricle; right ventricle; etc. | Digisonics; Inc | 2 | Left ventricle volume was inaccurate when the volumes were measured from images having different scales. | The firm, Digisonics, notified consignees/customers via e-mail entitled "Distribution of Notification of Biplane Error" dated February 6, 2009. The email described the product, problem and actions to be taken. The customers were instructed to immediately contact Digisonics if their site calculates biplane volumes-a utility will be provided which will run on their database and identify all studies in which biplane volumes are calculated. The customers were also instructed to share the following important information with all users of DigiView at their site. The email stated -until the software patch has been installed, biplane volumes should be calculated only from pairs of images having the same depth scale. Replacement DigiView/ERS software that prevents further occurrence of this problem will be provided if it appears their site could be impacted. The Distribution of New Software Release was sent on March 6, 2009. If you have any questions, please call Digisonics Customer Support at 800-344-4266.
Z-1008-2011 | Feb-01-2011 | 2011 | CyberKnife Treatment Planning System - CyberKnife Treatment Planning System; a subsystem of the CyberKnife Robotic Radiosurgery System; with MultiPlan Treatment Planning System Software version 3.5Medical charged-particle radiation therapy system; intended for treatment planning and image-guided stereotactic radiosurgery and precision radiotherapy of lesions; tumors and conditions anywhere in the body when radiation treatment is indicated. | Accuray Inc | 2 | If electron density values are left empty; the calculation of radiation dose in a patient will be modeled as air-like density material rather than correct density. A plan may be created and saved; thus creating the risk of mistreatment. | The firm, Accuray, sent an "URGENT DEVICE CORRECTION" letter dated December 29, 2009, to all customers. The letter described the product, problem, and action to be taken by the customers. The customers were instructed to verify correctness of the CT number to relative electron density and Mass Density Calibration curves. Note: A patch is being developed to address this issue. Accuray Customer Support will contact the customers regarding further resolution. If you have any questions or concerns regarding this issue, please contact Accuray Customer Support at +1-877-668-8667 (USA) or +1-408-716-4700 (non USA) or by email: customersupport@accuray.com.
Z-2061-2011 | Apr-25-2011 | 2011 | Siemens ACUSON X300 Premium Edition Systems with software versions 5.0.00 and 5.0.01 - Siemens ACUSON X300 Premium Edition Systems with software versions 5.0.00 and 5.0.01 (model numbers 10348531; 10348532 and 10348533); when using the V5Ms transducer (model number 08263703) and the Plus option (10427795)Product is an ultrasound imaging system intended for General Radiology;Fetal; Abdominal; Intaoperatve; Pediatric; Small parts; Neunatal/Adult Cephalic; Cardiac; Transesophageal; Pelvic; Transcranial; OB/GYN; Urology; Vascular; Musculoskeletal; Superficial Musculoskeletal; and peripheral Vascular Applications. The system also provides for the measurement of anatomical structures and for analysis packages that provide information that is used for clinical diagnosis purposes. The Arterial Health Package software provides the physician with the capability to measure Intima Media Thickness and the option to reference normative tables that have been validated and published with peer-reviewed studies. The information is intended to provide the physician with an easily understood tool for communicating with patients regarding the state of their cardiovascular system. The V5Ms is a multifrequency;multi-plane; phased sector array transducer; which can be introduced into the esophagus to obtain heart structure images. Transesophageal echocardiography is used to obtain unobstructed views of the heart surrounding tissue. | Siemens Medical Solutions USA; Inc. | 2 | Product Failure-- When the 64-Channel System V5M is used in the 128-Channel mode; an over current fault occurs which shuts the system down. | Siemens sent a Customer Safety Advisory Notification letter on July 24, 2009. The letter identified the product, the problem, and the action to be taken by the customer. A software update will be conducted in the field. Until that update occurs, customers were instructed to image without the CW doppler to avoid a system shut down. Customers were instructed to contact their service support person with any questions. For questions regarding this recall call 650-694-5993.
Z-2166-2011 | May-10-2011 | 2011 | DigiView software - DigiView software release 3.4.4.4 through 3.5.5.2. ; 2008 Digisonics; Inc; 800-940-3240.Digiview is a digital image management and reporting system with an associated clinical database. It provides capture; review and storage of images together with report creation and data basing. | Digisonics; Inc | 2 | Failure of software to update some measurements when tracings or points are amended after measurement calculations have been performed. | Digisonics, Inc., notified consignees by telephone and e-mail on 10/17/2008, informing them that a problem had been found in the Save Linked Images functionality of DigiView Release 3.6 (and all releases which start with 3.6 such as 3.6.1.). The problem is the failure of the software to update some measurements when tracings or points are amended after measurement calculations have been performed. A utility program, called Linked Image Measurement Verify was being supplied to all sites as a patch to the DigiView program. Release 3.6.6 was to be provided to all customers having previous 3.6 version. Problem was fixed in version 3.6.6 and 3.7.1. If you have any questions, please contact Digisonics Customer Support at (800) 940-3240.
Z-2180-2011 | May-10-2011 | 2011 | DigiView software - Digisonics; DigiView and ERS software release 3.6.4.4 through 3.6.5.2+; 2008 Digisonics; Inc.; 3701 Kirby Drive; Houston TX 77098; 800-940-3240.DigiView is a digital image management and reporting system with an associated clinical database. It provides capture; review; and storage of images together with report creation and data basing. | Digisonics; Inc | 2 | Data from one patient's study had the potential for being put into another patient's report. | Digisonics notified end users via e-mail January 20, 2009 and January 21, 2009, that Data from one patient's study was being put into another patient's report. The problem occurred when a System bug was found in DigiView Release 3.6.4.4, and higher versions in which under certain circumstances (specific workflow) a failure to switch to a patient's images occurs. Two software patch versions will be made available to sites on or after February 5, 2009, to correct the reported problem. If you have any further questions, please contact Digisonics Customer Support at (800 ) 940-3240.
Z-0146-2012 | Nov-01-2011 | 2011 | RIO Robot Arm Interactive Orthopedic System - Labeling for RIO Robot Unit is comprised of three main components: 20399 RIO SURGICAL ARM; 201251 RIO GUIDANCE MODULE; 200294 RIO CAMERA STAND ASSEMBLY***PN 203999 SN ROB 125 2010-12 V 100/120/230 A 9.6/8.0/4.2 Hz 50/60 Class I Equipment. Conforms to IEC 60601-1/A2: 1995; EN 60601-1/A2: 1995 UL 60601-1: 2003; CAN/CSA-C22.2 No. 601.1-M90.***Manufactured in USA***MAKO SURGICAL CORP. 2555 DAVIE ROAD; FT. LAUDERDALE; FL 33317.RIO System - The Tactile Guidance System v2.0 is intended to assist the surgeon in providing software defined spatial boundaries for orientation and reference information to anatomical structures during orthopedic procedures. The Tactile Guidance System v2.0 is indicated for use in surgical knee procedures; in which the use of stereotactic surgery may be approriate; and where reference to rigid anatomical bony structures can be identified relative to a CT base model of the anatomy. These procedures include unicondylar knee replacement and/or patellofemoral knee replacement. | Mako Surgical Corporation | 2 | MAKO Surgical Corp. is recalling their RIO Robotic Arm Interactive Orthopedic System (RIO) due to software issue that exist that could potentially result in a bone resection. No adverse events reported. | The firm, MAKO Surgical Corp., sent an "RECALL NOTICE 2.4. SITES" letter dated February 1, 2011 to its customers. The letter describes the product, problem, actions to be taken which includes a workaround and an updated user guide identifying the appropriate workflow. The customers were instructed to follow the user guide and workflow and to complete and return the enclosed confirmation memo to their MAKOplasty Specialist or return it in the pre-addressed, pre-paid envelope provided. MAKOplasty included with the letter new steps in the surgical workflow in the use of the RIO system to ensure error value accuracy; worked on a software update that will not allow the user to proceed with error values that do not support the required accuracy-this software was to be supplied to customers in February 2011, and trained MAKOplasty Specialist will make available written updates to the RIO use instructions as well. If you have any questions or concerns, please do not hesitate to speak with your MAKOplasty Specialist or contact the Manager of Customer Relations at 954.927.4022 x447.
Z-0177-2012 | Nov-09-2011 | 2011 | Capintec CRC 25 PET - Capintec CRC 25 PET; item number 5130-3217 with Software Revisions 2.01 and 2.02PET dose calibrator | Capintec Inc | 2 | Software error in Rev. 2.01 and Rev 2.02 for Capintec CRC 25PET | Capintec, Inc. sent a Software Error Alert letter dated July 8, 2011, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to contac their Capintec Customer Support Representative to receive a software upgrade. For questions call Capintec Customer Support at 1-800-631-3826.
Z-0277-2012 | Nov-23-2011 | 2011 | SoftLab GUI versions 4.0.2.8-4.0.2.10; 4.0.3.12-4.0.3.13; and 4.0.4.1-4.0.4.4. - SoftLab GUI versions 4.0.2.8-4.0.2.10; 4.0.3.12-4.0.3.13; and 4.0.4.1-4.0.4.4. SoftLab is a laboratory information system to be used in a medical research or clinical laboratory by knowledgeable; trained; and experienced personnel. SoftLab enables the user to handle data-processing and laboratory activities; for the purpose of storing; managing; querying; and report laboratory data. | SCC Soft Computer | 2 | On 05/03/2011 SCC Soft Computer; Clearwater; FL initiated a correction on the following SoftLab GUI versions 4.0.2.8-4.0.2.10; 4.0.3.12-4.0.3.13; and 4.0.4.1-4.0.4.4. A client reported the PT and INR results did not get flagged. The PROT should have been flagged as HIGH and the INR should have been flagged as PANIC. Neither test was flagged. All affected clients were notified of the issue on 05/03 | SCC Soft Computer initiated a Field correction on May 3, 2011. All customers with affected software have been notified of this issue. Clients were instructed to have a "valcheck" utility run to determine if there are any tests in their system affected by the issue. They were supposed to indicate whether they wanted to schedule to receive the patch or hot fix. If you have any further questions please call ( 727 ) 789-0100.
Z-1831-2011 | Mar-28-2011 | 2011 | enGen TM Laboratory Automation System - The enGenTM Laboratory Automation System consists of 3 main components including Instrument Management Subsystem (Instrument Manager); Automated Track Subsystem (TCAutomation) and Configuration File developed by Ortho-Clinical Diagnostics (OCD). The Instrument Manager is a Windows-based software product that operates on a Pentium-class (or higher) PC processor. The instrument management subsystem is capable of connectivity with clinical analyzers; a Laboratory Information Systems (LIS); barcode label printers and other third party devices. The TCAutomation is a pre-analytical and post-analytical sample management system. The configuration file enables operational support of the site specific layout and instrument selection. The VITROS 5; 1 FS Chemistry System with enGenTM Laboratory Automation System is intended for use in the in vitro quantitative measurement of a variety of analytes of clinical interest; using both VITROS Chemistry Products Slides (colorimetric endpoint; rate; ion-selective electrode; and immunorate methods) and VITROS Chemistry Products MicroTip liquid reagents (spectrophotometric and spectrophotometric Immunoassay methods.) | Ortho-Clinical Diagnostics | 2 | A software anomaly that can occur with analyzers connected to the Bypass modules listed below on an enGen" Laboratory Automation System using TCAutomation" Software Version 3.1.1 and below. When this anomaly occurs; it is possible for a sample other than the intended sample to be aspirated; potentially leading to test results being assigned to the wrong patient without error notification. | The firm, Ortho Clinical Diagnostics (OCD), sent an "URGENT PRODUCT CORRECTION NOTIFICATION" letter dated January 27, 2011, by Federal Express overnight mail to all customers. Foreign affiliate consignees/customers were notified by email informing them of the issue on the same day. The letter described the problem, product and actions to be taken. The firm also included temporary procedures to help prevent the occurrence of the software anomaly. The customers were instructed to follow the enclosed Procedure to STOP the Automation System and/or Bypass Module; complete and return the attached Confirmation of Receipt - Important Response Required form by February 4, 2011, via fax to: Ortho Clinical Diagnostics, 1-888-557-3759 or 1-585-453-4110, and display this communication near the enGen Laboratory Automation System to ensure compliance by all operators until updated TCAutomation Laboratory Automation System software is installed. A Technical Bulletin will be sent to customers in the near future to be stored with user documentation. Ortho will implement a TCAutomation Laboratory Automation System software change in the near future to resolve this anomaly. If you have any questions regarding this notification, please call Customer Technical Services at 1-800-421-3311 (Options 1, 2).
Z-3160-2011 | Sep-07-2011 | 2011 | Picis ED PulseCheck-Software - Picis ED PulseCheck-Software; transmission & storage; patient dataService Pack v5.0 L and MThis product is a non-alarming software patient information system that compiles an electronic medical record utilizing commonly available hardware. A patient record is populated with information from various sources; such as health care professionals; medical devices connected to the system; and data that arrives via hospital and laboratory information systems. The application stores this information in a database and displays it to healthcare providers in different formats for informational purposes. The product is intended for use by healthcare professionals. | Picis Inc. | 2 | Software does not display appropriate allergy interaction warning | Picis issued notification on 08/01/11 to affected clients, explaining the potential for the error, with instructions to monitor documented patient allergies closely when prescribing medication orders in ED PulseCheck, using Muitum drug database content. Picis will track the progress on acknowledgement of notification letter receipt through the implementation of software upgrade, for all potentially affected Customers. For questions regarding this recall call 781-557-3000.
Z-2564-2011 | Jun-14-2011 | 2011 | Monaco Radiation Treatment Planning System - Monaco Radiation Treatment Planning System; Monaco Release 2.0.1 and aboveUsed to create treatment plans for any cancer patient for who external beam intensity modulated radiation therapy (IMRT) has been prescribed. | Computerized Medical Systems Inc | 2 | Monaco: due to a software defect; the Radiological Depth (cm) and SSD to DRP (cm) parameters are reported incorrectly on the Beam Summary Report. This occurs for all Monaco IMRT plans. | Computerized Medical System sent a "USER NOTICE" letter dated June 24, 2010, to all affected users. This notice identified the product, the problem, and the actions to be taken. There is no current workaround. All customers will be mailed a postcard to be returned as proof of receipt and understanding of the notification. Consignees will be contact via telephone or e-mail after June 30, 2010 if the postcard is not returned This notice is to be distributed to any and all users of CMS software at your organization who are potentially affected. The notice states the issue will be resolved in Monaco Release 2.04, but no date for this release was given. For further questions, please call (403) 830-8023.
Z-1834-2011 | Mar-28-2011 | 2011 | Coherence Dosimetrist; v 2.2 - Coherence Dosimetrist; version 2.2.Intended use: Medical charged-particle radiation therapy system | Siemens Medical Solutions USA; Inc | 2 | A problem in the software has the potential to be associated with the dose; when a template is created the user has the possibility to predefine organs. For the Target the user defines the max. dose and the min dose; for the OAR (Organs at Risk) the user only defines the max. dose. The firms engineering team determined that the software defect was repeatable. | Affected consignees will be notified by the Customer Service Engineer to schedule a software update beginning November 2009.
Z-2611-2011 | Jun-22-2011 | 2011 | Pathfinder Holter Analyzer - Pathfinder Holter Analyzer is a Programmable Diagnostic Computer. The brochure is labeled in part: "***Pathfinder Holter Analyzer ***Spacelabs Healthcare Company Headquarters***5150 220th Ave. SE; Issaquah; WA 98029***PO Box 7018; Issaquah; WA 98027-7018***Telephone: 425-657-7200***In North America: 800-522-7025***www.spacelabshealthcare.com***".The Reynolds Pathfinder 700 Holter Analyzer is intended to be used to analyze magnetic tapes of ambulatory electrocardiograms made on compatible Holter Recorders. The system will detect various arrhythmias and will measure ST elevation or depression. The system acts in both a automatic mode (tape analyzed without operator intervention) and in an interactive mode; (operator can intervene and affect analysis). | Spacelabs Healthcare; Llc | 2 | A potential hazard associated with the use of Spacelabs Healthcare Pathfinder Holter Analyzer software. The second 24 hours of a 48 hours EVO recording will not be analyzed. | Spacelabs Healthcare sent a "URGENT-MEDICAL DEVICE CORRECTION" letter dated June 3, 2011 to all customers. The letter describes the product, problem, and the actions to be taken by the customers. Spacelabs Healthcare recommended that until software upgrades can be done, the customers should weigh the benefits versus the risks when deciding whether or not to continue using the product. Spacelabs Healthcare also recommended that customers advise their staff that the second 24 hours of a 48 hour EVO recording would not be analyzed. Field service engineers will be contacting customers to set up a time and date to install the revised version of the Pathfinder software. Spacelabs will upgrade Pathfinder software regardless of whether the consignees are using EVO recorders to ensure this hazard does not happen in the future. Consignees can call Spacelabs Healthcare at 1-800-522-7025 and select 2 for Technical Support, or outside the US, 44 1992 507740 for any questions about this recall.
Z-2604-2011 | Jun-20-2011 | 2011 | XiO Radiation Treatment Planning System - XiO Radiation Treatment Planning System; XiO Release 4.50.00 and above | Computerized Medical Systems Inc | 2 | XiO Software: when a treatment machine is defined in Source File Maintenance (SFM); the MLC Model is set to Elekta; and beams using this treatment machine are added to a new plan in Teletherapy; XiO automatically adds an MLC to the beam. This is because the Elekta linac will always have an MLC in place when the plan is actually delivered. | A Safety Notice, dated 5/20/10, was sent to Elekta / Computerized Medical Systems Inc clients in May 2010. The notice indicates the problem, the clinical impact. The notice states a future XiO Release will resolve the issue. A return postcard is included with the mailing to verify consignee received the notice. For sites not returning the postcard by 5/28/10, the recalling firm will contact them via phone or email to confirm receipt of the notice. The notice also informed the consignees they would be contacted when a solution was available.
Z-2292-2011 | May-25-2011 | 2011 | WaveLight FS 200 Femtosecond Laser System - WaveLight FS200 Femtosecond Laser System Alcon Product Code 8065990714 WaveLight Model 1025 Software Version: Version Blue; Service Pack 2.The WaveLight FS 200 Laser System is an ophthalmic surgical laser indicated for use in refractive surgery for LASIK flap cutting; lamellar and perforating keratoplasty; and intrastromal surgery for implants. | Wavelight AG | 2 | The software calculating the patterns were found to has a rounding error for certain parameter ranges. These rounding errors could result in an incomplete cut in the donor or patient tissue. | WaveLight AG sent an URGENT MEDICAL DEVICE CORRECTION letter April 18, 2011 to the affected customers. The letter identified the product, problem and the action to be taken by the customer. The letter advised the customers to discontinue keratoplasty cuts with the WaveLight FS200 until the new version can be installed. Alcon will contact you to schedule a visit of a service engineer at your site, to disable the keratoplasty function in your device, until a new software version can be installed to fix this problem. If you have experienced a product problem or adverse event related to this medical device correction, please contact Alcon Medical Safety at 1-800-561-6466, option 3. Customers were requested to complete the attached reply form to acknowledge receipt of this notification and fax it to ( 817) 916-9087.
Z-2714-2011 | Jun-30-2011 | 2011 | SoftReports versions 1.1.6.x and 1.1.7.x - SoftReports versions 1.1.6.x and 1.1.7.x SoftReports version 1.1.6 released 2/2008; Version 1.1.7 released 6/2008.SCC Soft Computer 5400 Tech Data Drive Clearwater; FL 33760.SoftReports is a report designer and web-based report launching tool to be used by knowledgeable; trained; and experienced personnel. It provides the ability to create ad hoc; user defined queries; and reports using data from other SCC products. SoftReports enables users to customize standard reports generated by SCC systems to meet their specific needs. This system performs the following functions: 1. The ability to customize report layouts using beginner; intermediate; or advanced options. 2. The ability to import and export reports. 3. The ability to provide custom desktops to launch SQL's within other SCC applications. 4. The ability to preview and modify layouts specific to users' needs. 5. The ability to modify elements of reports using the comprehensive control formatting toolbar. 6. The ability to incorporate charts and graphics into reports. 7. The ability to modify existing query templates. 8. The ability to create independent and embedded queries. 9. The ability to pars SQL statements and notify users of errors within the syntax. 10. The ability to easily update criteria parameters to indicate input needed for executing queries. 11. The ability to view and print any report that is created using the Designer. 12. The ability to schedule; execute; and track ad-hoc reports. This system does not perform the following functions: 1. SoftReports is not configured to write back to the database. 2. SoftReports does not support user-created Ad Hoc reports being used for diagnostic purposes. 3. SoftReports is not intended to alter the meaning of patient result reports (test reports). | SCC Soft Computer | 2 | SCC Soft Computer is issuing a correction for SoftReports Versions 1.1.6.x released 02/2008 and 1.1.7.x released 06/2008. Potential hazard - when using SoftReports Layout Designer; results displayed in the report may not match the patient information in the report header. | SCC Soft Computers sent a "RISK TO HEALTH NOTIFICATION TASK" letter dated June 9, 2010 to affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers were offered an upgrade or hot fix. The letter requested acknowledgement of receipt of the task notification and permission to install the hot fix.
Z-2560-2011 | Jun-14-2011 | 2011 | Monaco RTP System - Monaco Radiation Treatment Planning System; Monaco Release 1.0.0 through 1.0.2. Used to create treatment plans for any cancer patient for who external beam intensity modulated radiation therapy (IMRT) has been prescribed. | Computerized Medical Systems Inc | 2 | Monaco: If the user selects Monte Carlo Photon as the Secondary Algorithm (to be used for final dose calculation through segments); and selects SKIP Forward button at any time during stage 2 optimization; the system may only use the pencil beam algorithm for final computation of dose even though Monte Carlo was selected. Other than specific output to the console window; there will be no other in | Computerized Medical Systems, Inc. sent a "MONACO CUSTOMER ADVISORY" letter dated June 2008 to all affected customers. The letter describes the product, problem, and the action to be taken. The customer was instructed to use caution when using the Skip Forward button in stage 2 according to written instructions provided in the letter. This defect exists in Monaco Release 1.0.0 through 1.0.2 and will be resolved as soon as possible. Note: The issue was resolved in Monaco Release 2.0.1 and above. Questions should be directed to the manufacturer's representative telephone number at 408-830-8023.
Z-2897-2011 | Jul-29-2011 | 2011 | Aplio Artida Ultrasound Diagnostic System SSH-880CV - Aplio Artida Ultrasound Diagnostic System SSH-880CV.Toshiba American Medical Systems; Inc.A diagnostic ultrasound system. | Toshiba American Medical Systems Inc | 2 | Toshiba America Medical Systems; Inc. is initiating a field correction on the Aplio Artida Ultrasound Diagnostic System because the panel operation for the Aplio Artida may become disabled (i.e.; system may hang up) when processing to display patient information (heart rate; time; VCP counter; ect.). | Toshiba America Medical Systems, Inc. sent out an "Urgent Medical Device Correction" dated December 21, 2010 to all affected customers. The letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were notified that Toshiba will issue a field modification to correct the software, and if the problem occurs, they were instructed to shutdown the system holding down the power supply button for 5 seconds for 5 seconds and then restart the system. A form to send back to Toshiba is included. For questions on this recall call (800) 421-1968 or contact your local Toshiba Representative at (800) 521-1968.
Z-2645-2011 | Jun-23-2011 | 2011 | ERGO++ - ERGO++ Stereotactic Radiation Treatment Planning System; ERGO++ Release 1.6.3 and 1.6.3.1Product Usage: Used to create treatment plans for any cancer patient for who external beam radiation therapy has been prescribed. It is an accessory to linear accelerators used for radiation therapy. It is indicated for use in the planning of 3-dimensional radiation therapy. | Computerized Medical Systems Inc | 2 | ERGO Release 1.6.3 is overestimating the MU values. | Computerized Medical Systems, Inc. sent an "USER NOTIFICATION" dated December 11, 2009, to all affected customers. The notice identified the product, problem and suggested a workaround to minimize the risk. The notice stated that this workaround does NOT substitute the need for regular Quality Assurance. Customers are instructed to review page 39 of user manual for further explanation of warnings. Customers were asked to complete the Notice for Customer section and return to 3D Line Medical System. This issue has been resolved in Release 1.7.0 and above. No follow-up notice will be distributed to customers. For further information please contact: CMS Software (314) 812-4460 (St Louis Office).
Z-2216-2011 | May-19-2011 | 2011 | ARTISTE ; ONCOR and PRIMUS accelerators - ARTISTE ; ONCOR and PRIMUS accelerators that utilize the RT Therapist v4.2 system in combination with syngo RT Oncologist 4.2.The ONCOR and PRIMUS may be updated to utilize the RT Therapist v4.2 software. The RT Oncologist v4.2 is currently only on Non-US systems. Distributed by Siemens Healthcare; Concord; CA 94520Intended use: Radiation therapy | Siemens Medical Solutions USA; Inc | 2 | Siemens has become aware of a potential safety issue that may result in mistreatment when customers are configured with the ARTISTE Linac RT Therapist v4.2 in combination with RT Oncologist v.4.2 and use the remote "Online 2D Image Review. | The firm, Siemens, sent an "Urgent: Medical Device Correction Customer Safety Advisory Notice" on March 31, 2011 via Siemens Customer Service Group or certified mail to its customers. The notice describes the product, problem and actions to be taken. The customers were instructed to not use the "Online 2D Image Review" on the Syngo RT Oncologist 4.2; to always calculate the offset directly on the RTT 4.2 before it is applied for the treatment with the patient on the treatment table; include this notice in their syngo RT Therapist 4.2 System Owner Manual chapter "Safety Advisory Letter" where it should remain until the measures defined in the notice are performed, and in the interest of safety, perform the preventative measures noted in this notice and inform all affected personnel immediately. Siemens is currently working on a solution for the incremental loading workflow in syngo RT Therapist 4.2 and will provide a software update as soon as possible. Siemens will notify the customers as soon as an update for their system will be available. UPDATE: The release of Safety Update UI TH010/11/S to distribute the software update for RTTv4.2.94O04 began on April 15, 2011. If you have any questions, contact Regulatory Affairs at 925-602-8157.
Z-1257-2011 | Feb-15-2011 | 2011 | Toshiba AQUILION ONE - Toshiba TSX-301A AQUILION ONE (MPDCT0297EA) - Whole Body X-Ray CT Scanner | Toshiba American Medical Systems Inc | 2 | Incorrect Dose Calculation. The software that performs the pre-scan dose calculation uses an incorrect parameter for an exam Plan using a tube voltage of 80kV or 100kV. This results in a pre-scan expected dose that is approximately 10% lower than what the displayed value should be when the scanner executes the exam Plan. | Toshiba America Medical Systems (TAMS) issued notification to consignees 9/7/2010, via USPS return receipt mail. The letter contains a return reply form that is to be faxed to TAMS as proof of notification.
Z-0917-2011 | Jan-18-2011 | 2011 | Philips MultiDiagnost Eleva with Digital Image (DI) Rel.2 system - Philips MultiDiagnost Eleva with Digital Image (DI) Rel.2 system.The Philips MultiDiagnost Eleva is intended use as a multi-functional universal system; general RF; Fluoroscopy; Radiography and Angiography can be performed along with more specialized interventional applications. | Philips Medical Systems North America Co. Phillips | 2 | Images from two different patients may be combined in one patient folder in the Digital Image (DI) or may be combined as one (printed) image. It is possible that an image of a previous patient registered in the DI will be overwritten with an image of a later patient. | On March 6, 2009, Philips sent an "URGENT-Device Correction Easy Diagnost Eleva, MultiDiagnost Eleva, and OmniDiagnost Eleva with DI Rel. 2 systems notification" to their consignees/customers. The consignees/customers are advised that before diagnosing images, always check the stored (Digital Image, Viewform or PACS) or printed images to ensure that these represent the data and images captured from the examination of the respective patient. Do not use images if they do not represent the patient that was examined. Consignees will be contacted by Philips to schedule a software upgrade of their units. The consignees can contact Philips Healthcare Call Center at 800-722-9377, #5, #2 and reference "FCO 70800070" with any questions about this recall.
Z-2622-2011 | Jun-23-2011 | 2011 | Interplant - Interplant Radiation Treatment Planning System; Interplant Release 3.4.0Used when planning and performing prostate brachytherapy implant procedure. It provides ultrasound probe alignment and radioactive seed implantation. | Computerized Medical Systems Inc | 2 | Interplant Software: If the trajectory needle with a non-zero retraction is adjusted in any of the planning views (Transverse; 3D) and the program is closed and reopened; the Needle Definitions page in the Dose Plan Summary (Plan View) displays the needle with a 0.00 retraction even when the needle is displayed correctly in the planning view. If the Eliminate Leading Spacers function is applied t | Computerized Medical Systems, Inc. sent a "CUSTOMER ADVISORY" letter dated December 1, 2008, to all affected customers. The letter identified the product, the problem, and the actions to be taken. Customers were instructed to discontinue use of the product. There is no workaround at this time. This issue was resolved in Interplant Release 3.4.1 software. All possibly affected users have been shipped the Interplant Release 3.4.1 software in May 2010. If you have any questions, please call 314-993-0003 or toll free 800-878-4267.
Z-1029-2011 | Feb-04-2011 | 2011 | SEQUENCER for Mosaiq - SEQUENCER for Mosaiq interfaced to Electa SL75-5 machines.Sequencer Verify and Record System is to assist in the process of patient treatment on a radiotherapy treatment machine with software control. | Impac Medical Systems Inc | 2 | Computer interface is defected. Computer based verification system for recording patient treatment on a radiotherapy treatment machine fails to record treated field. | A User Notice was sent by the firm to its consignees. It identified the affected product and the reason for the recall. It also discussed the clinical impact and interim mitigations. The firm states in the letter that a fix for the problem and been developed and is available. Customers are to contact IMPAC Support for the patch to the SL75-5 interface. If customers have questions, concerns, or requests, they can email support@impac.com or one of the numbers listed in the letter for the applicable country. The letter should be distributed to all users affected by the issue at the customer's organization.
Z-1498-2011 | Mar-03-2011 | 2011 | SSH-880CV Aplio Artida; System; Diagnostic Ultrasound - SSH-880CV Aplio Artida; System; Diagnostic Ultrasound | Toshiba American Medical Systems Inc | 2 | Toshiba America Medical Systems is initiating a field correction on the SSH-880CV Aplio Artida Diagnostic Ultrasound System due to software issues. | Toshiba America Medical Systems (TAMS) sent an URGENT MEDICAL DEVICE CORRECTION letter to consignees via USPS return receipt mail. The letter contained a return reply form that is to be faxed to TAMS at 877-349-3054 as proof of notification. Customers were instructed to share the information with all users and reviewing radiologists as well as clinical engineering or biomedical group at their facility. For questions regarding this recall call 800-421-1968 or 800-521-1968.
Z-2762-2011 | Jul-08-2011 | 2011 | syngo Dynamics 9.0 - syngo Dynamics Picture Archiving Communication SystemThe software is a picture archiving and communication system intended for acceptance; transfer; display; storage; archive; and manipulation of digital medical images including quantification and report generation. | Siemens Medical Solutions; USA; Inc | 2 | For systems upgraded to syngo Dynamics version 9.0; if a legacy report generated by version 3.x or earlier is opened in the version 9.0 portal; the legacy report may be deleted and replaced with a new report from the version 9.0 reporting system or a blank report. | Siemens Medical Solutions, USA, Inc sent a letter entitled "Customer Safety Advisory Notice" to all direct accounts on May 10, 2011. The letter stated the name and version of the software, problem noted, a work-around to be implemented until a service patch is created to fix the problem, and that the letter should be placed in the software's operator's manual to be seen by all who utilize the software.
Z-2204-2011 | May-19-2011 | 2011 | CHVH-001A/1B Variable Helical Pitch (vHP) Software for TSX-101A: Aquilion 64 - CHVH-001A/1B Variable Helical Pitch (vHP) Software for TSX-101A: Aquilion 64; System; X-Ray; Tomography Computed | Toshiba American Medical Systems Inc | 2 | The recall was initiated by Toshiba America Medical Systems (TAMS) because the software associated with TSX-101A: Aquilion 64 can prevent scanning from being performed at the specified helical pitch (couch movement speed). | Toshiba America Medical Systems (TAMS) forwarded an Urgent: Medical Device Correction letter, dated 4/6/10, with attached Customer Reply Form, via US Postal Service to all customers who purchased the CHVH-001A/1B variable helical pitch (vHP) software for TSX-I0IA: Aquilion 64. The letter provides the customers with an explanation of the problem identified and actions to be taken. Customers were informed that a Toshiba service representative would contact them to schedule an appointment in order to install a corrective measure that will prevent the occurrence of the problem. Toshiba would like customers to set a scan range different from that for the previous scan for the second and subsequent vHP scans, if multiple vHP scans with the same scan conditions are performed in Auto Mode. Customers were instructed to complete and return the attached form and fax it to the toll free number at the top of the form. The form can also be sent via e-mail to raffairs@tams.com. Customers with any questions can contact the Director of Regulatory Affairs at (800) 421-1968 or contact their local Representative at (800) 521-1968.
Z-2594-2011 | Jun-20-2011 | 2011 | FOCAL Workstation - Focal Radiation Treatment Planning System; Focal Release 4.40.00 and above.Product Usage: for use in radiation treatment planning using generally accepted contouring methods. The Focal software allows users to access XiO and other RTP systems from a remote PC-based system and provides the user with the ability to edit and modify patient data from a Windows-based PC and transfer the information back to the RTP system. | Computerized Medical Systems Inc | 2 | Focal: the couch position can be moved between scans of the same CT study. When the images are imported into Focal; the software is not checking the DICOM Image Position (Patient) tag and therefore is not aligning the images sets if there are different couch positions indicated. A shift between image sets will exist of the same amount that the CT couch was moved. | Elekta CMS Software sent an "ADVISORY NOTICE" dated January 19, 2009 to its affected customers. The notice identified the product, the problem, and the acton to be taken. Customers were informed that the workaround method available is to (use the image fusion feature already in Focal to relign the scans). The firm states that this issue is resolved in Focal version 4.62.00, which was released in March 2011. If you have any questions, please call 314-993-0003 or toll free 800-878-4267.
Z-2293-2011 | May-25-2011 | 2011 | syngo Dynamics version 9.0 - syngo Dynamics version 9.0; Siemens Medical Solutions; USA; Inc.; Ann Arbor; MI.The syngo Dynamics is a picture archiving and communications system intended for acceptance; transfer; display; storage; archiving and manipulation of digital medical images; including quantification and report generation. It is not intended to be used for the reading of mammography images. | Siemens Medical Solutions; USA; Inc | 2 | If one or more graphs are in a table on the report; and if any of those graphs are deleted; it may lead to reports failing to save and/or marking the study as read might leave the study in an unread state or might mark the study as read but leave the report with a "Preliminary" watermark displayed. | The firm, SIEMENS, sent a "Customer Safety Advisory Notice" letter dated March 24, 2011 to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to place graphs in sections and not tables in the report; avoid deleting graphs from the reports in run-time and include this safety advisory notice in their operator's manual where it should remain until the service patch is applied. SIEMENS will be issuing a service patch to fix the defect. The letter states the patch will be available in the next few weeks. If you have any questions, call +1 (734) 205-2400.
Z-2203-2011 | May-19-2011 | 2011 | SSH-880CV Aplio Artida - Software for SSH-880CV Aplio Artida; System Diagnostic Ultrasound | Toshiba American Medical Systems Inc | 2 | Toshiba America Medical Systems (TAMS) has become aware that the FLEX M software has reversed left and right annotation and post FLEX M measurement error; which may result in incorrect values. | The recall communication was initiated on 1/6/10 with Toshiba forwarding an Urgent: Medical Device Correction letter with attached Customer Notification Reply Form (via US Postal Service for US Customers) to all customers who purchased the FLEX-M mode on Aplio Artida Diagnostic Ultrasound System. The letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to avoid performing M image measurement using images saved in FLEX-M mode until the corrective action is completed. If customers need to perform measurement for images saved in FLEX-M mode, they were instructed to use manual calibration. Customers were instructed to note that even though the corrective action is taken, the existing saved images (saved images before the corrective action is taken) are not modified. Therefore, when referring to the measured values on the images which are saved in FLEX-M mode before the corrective action is taken, customers were instructed to carefully handle the measurements results because they may be incorrect. Customers were instructed to complete the attached form and fax it to the toll free number at the top of the form. The form can also be sent via e-mail to raffairs@tams.com. If customers have any questions regarding the medical device correction letter, they can contact the Director of Regulatory Affairs at (800) 421-1968 or their local Toshiba representative.
Z-0293-2012 | Nov-28-2011 | 2011 | Ingenia 1.5T and 3.0T MRI Systems - Ingenia 1.5T and 3.0T MRI Systems; Model Numbers: 781396 and 781377; 510(k) #K110151. System units equipped with software version: 4.1.1 are subject to recall/field correction. This MRI system is indicated for use as a diagnostic device. It can produce cross-sectional images; spectroscopic images and/or spectra in any orientation of the internal structure of the head; body; or extremities. Image appearance is determined by the spatial distribution and flow as well as by many different NMR related properties of the tissue and anatomy studied and the MRI scan technique applied. These images and/or spectra; when interpreted by a trained physician; provide information that may assist the diagnosis. | Philips Medical Systems (Cleveland) Inc | 2 | The firm determined that during scanning procedures; for some larger patients; some reddening of the skin of can occur in certain instances due to RF heating when the torso of the patient is positioned within the body coil. | Philips Medical Systems ( Cleveland) Inc. sent a Urgent Field Safety Notice dated August 3, 2011, to all affected customers. Philips Field Service Engineers were dispatched to the affected customer sites to deliver the recall notification and to perform the necessary software corrections. The letter lists the affected MRI systems; describes the problem; the actions to be taken by the customer/user; it also informs the customer of the upcoming software correction which will be applied to each unit by the Philips Field Service Engineers in order to address this issue. If any additional information is required, the customer is instructed to contact their local Philips Field Service Engineer. For further questions please call (440) 483-7000.
Z-1245-2011 | Feb-15-2011 | 2011 | GE Centricity Laboratory - GE Centricity Laboratory software; GE Healthcare IT; Barrington; IL 60010. The intended use: the Centricity Laboratory is used to facilitate the general clinical; anatomic pathology and cytology laboratory workflow such as order entry; results entry; instrument interfacing; results reporting and patient record retrieval. The Centricity Laboratory software is used to facilitate the general clinical; anatomic pathology and cytology laboratory workflow such as order entry; results entry; instrument interfacing; results reporting and patient record retrieval. It integrates with other Hospital Information Systems (HIS) through HL7 interface by order entering and report processing. The centricity is intended to interface with various lab devices through standard HL7 interfaces to record; report and distribute lab results. | Ge Healthcare It | 2 | The GE Centricity Laboratory System Software has safety issues with the "Telephone List" function; the "Anatomic Pathology" application; and the Instrument Interface IO module that may impact patient safety. | GE Healthcare, sent "Urgent Medical Device Correction" letters dated December 12, 2010 to GE Centricity Laboratory software customers on the same date. The letters advised the users of the product, the problem, and the action to be taken by the customer. The letters provided the users with specific safety instructions to follow until the software was updated. The customers were requested to complete the enclosed customer reply form and fax the completed form to 847-620-2780. Any questions should be directed to their GE Service Representative or the GE Help desk at 888-778-3375.
Z-2467-2011 | Jun-08-2011 | 2011 | HIVISION Preirus; Avius ultrasound systems - Hitachi HIVISION Preirus/Avius Ultrasound equipped with V03.02 Software - Doppler Trace Measurement program. 510 (k) Numbers: Preirus - K093466 (6/1/10); Avius - K102901 (10/27/10).The Preirus and Avius are intended for use by trained personnel (doctors; sonographers; etc) for the diagnostic ultrasound evaluation of Abdominal; Cardiac; Intra-operative; Feta; Pediatric; Small Organ; Peripheral vessel; Biopsy; Trans-rectal; Trans-vaginal; Musculoskeletal; Neonatal Cephalic; Adult Cephalic; Endoscopy; Intra-luminal; Genecology; Urology and Laparoscopic clinic applications. The Modes of Operation of the HIVISION Preirus are B mode; M mode PW mode (Pulsed Wave Doppler); CW mode (Continuous Wave Doppler) Color Doppler; Amplitude Doppler (Color Flow Angiography); TDI (Tissue Doppler Imaging); 3D Imaging; 4D Imaging; RealTime Tissue Elastography; and Real Time Virtual Sonography. | Hitachi Medical Systems America Inc | 2 | Software anomaly. The Doppler Trace Measurement feature on certain models of the firm's diagnostic ultrasound systems; may; under certain conditions whereby multiple measurements are done sequentially; record incorrect values. The recorded value is displayed to the user on a report; however depending upon the measurement area; displayed velocity values could be lower than at the true velocity lead | The firm, Hitachi, sent a "DEVICE CORRECTION NOTICE" letter dated April 15, 2011 to its customers. The letter describes the product, problem and actions to be taken. The customers were inform of the nature of the software error in the Doppler Trace Measurement feature; the circumstances under which this error can occur; and the safety issue involved should incorrect values generated lead to a false diagnosis of stenosis in a patient. The letter provides the customer with a workaround solution in order to avoid this issue. The customer were also instructed to call the HMSA Applications Helpline at 877-871-2777 to reprogram their system to temporarily eliminate the noted safted issue and to immediately have all selections unchecked on the Doppler trace measurement result store setting set up screen. The notification letter also states that Hitachi Medical Systems America (HMSA) will be having their service personnel schedule a visit at each customer location and install a software upgrade on each affected system unit in order to directly address, and permanently resolve this issue. If you need any further information regarding this notice, please contact the Manager of Regulatory Affairs at 800-800-3106 x3720 or email to: thistlewaited@hitachimed.com.
Z-1848-2011 | Mar-30-2011 | 2011 | HEMOCUE AB GLUCOSE 201 DM ANALYZER - HEMOCUE AB GLUCOSE 201 DM ANALYZER; in vitro diagnostic device for quantitative determination of blood glucose level. | Hemo Cue; Inc. | 2 | Incorrect configurations settings in software. As a result; certain error codes do not trigger appropriately. These error codes are: E01; E02 and E05 which are all related to the optronic performance of the analyzer. Possible incorrect results. | Hemocue AB sent Customer Notification "Notice of Medical Device Correction on HemoCue Glucose 201 DM and HemoCue Hb 201 DM Analyzers" dated October 7, 2010, to customers notifying them of incorrect settings on error codes. The corrective action consists of an update of the configuration settings on the the affected HemoCue Glucose 201 DM and Hb 201 DM analyzers that are in the field. The affected units in the field will be corrected in accordance with a rework procedure that has been established by HemoCue AB.
Z-2915-2011 | Aug-01-2011 | 2011 | GEMINI TF Big Bore - GEMINI TF Big Bore computer tomograph software version 3.6. Catalog #882476; Gantry S/N 9001-9023.The GEMINI TF Big Bore is a diagnostic imaging systems for fixed or mobile installations that combines Positron Emission Tomography (PET) and X-ray Computed Tomography (CT) systems. The CT subsystem produces cross-sectional images of the body by computer reconstruction of x-ray transmission data. The PET subsystem produces images of the distribution of PET radiopharmaceuticals in the patient body (specific radiopharmaceuticals are used for whole body; brain; heart and other organ imaging). Attenuation correction is accomplished by CTAC. The device also provides for list mode; dynamic; and gated acquisitions. Both subsystems (PET and CT) can also be operated independently as fully functional; diagnostic imaging systems including application of the CT scanner as a radiation therapy simulation scanner. | Philips Medical Systems (Cleveland) Inc | 2 | Two non-safety related defects have been found on the Philips GEMINI TF Big Bore running version 3.6; if these defects were to recur; it could impact the performance or reliability of the system. Philips is issuing a software update to all Philips GEMINI TF Big Bore customers to correct these defects. | Philips Medical Systems (Cleveland) Inc. sent an "GEMINI TF BIG BORE CPE CUSTOMER LETTER" dated October 20, 2011 to all affected customers. The letter describes the product, problem, and updates that are available to the customers. The letter provides additional information about acquisition workflows for SUV Calibrations and Validations. For support concerning this upgrade, contact the Customer Care Solutions Center at 1-800-722-9377.
Z-1592-2011 | Mar-07-2011 | 2011 | Straumann coDiagnostiX - Straumann coDiagnostiX; dental implant planning system. Software Version 7.0 including updates 7.0.2 and 7.0.2.10 | Straumann Manufacturing Inc. | 2 | Software: Surgical Error may occur when the implant length is changed after positioning the surgical sleeve. | Straumann notified accounts of this correction by letter dated September 30, 2010 and delivered via Federal Express Delivery Service. The upgrade is available on-line by logging into the coDiagnostiX system and enabling the automatic upgrade to occur.
Z-0032-2012 | Oct-07-2011 | 2011 | Philips Medical - MDC PACS - release R2.3 SP1 .Phillips HealthcareAndover; MassachusettsA software application that is used for receiving; managing; archiving; distributing and recording medical images onto portable digital media (including but not limited to Compact Disk and DVD). | Philips Healthcare Inc. | 2 | Clinical measurements on true size printouts may be inaccurate | Philips Medical notified all affected customers with a "Field Safety Notice" dated August 2011. The letter provides instructions to the customer for how to avoid the problem. Customers are informed that until the software upgrade is recieved from Philip, do not use the true-size printing feature if the system is configured for more than one DICOM printer. Customers will be instructed that they will receive a software upgrade when it is available. For questions on this recall contact Philips representative at (800) 722-9377.
Z-2183-2011 | May-11-2011 | 2011 | TEMPUS IC PROFESSIONAL - TEMPUS IC PROFESSIONAL Patient Monitor; Part No. 00-1002; Manufactured by: Remote Diagnostics Technologies Ltd.; The Old Coach House; Farleigh Wallop; Hants; UK RG25 2HT.Indicated for 3 Lead ECG monitoring and 12 Lead ECG recording; non-invasive blood pressure (NIBP); respiration; end-tidal C02 (ETCO 2); pulse oximetry (SP02) and tympanic temperature. | Remote Diagnostic Technologies Ltd. | 2 | It was discovered that when a 12 Lead ECG is recorded and monitoring is quickly restarted; the ECG can cease functioning. | Remote Diagnostic Technologies, Ltd notified its consignees of the problem and the work-around via letter with the subject line "ECG Issue - Field Action Required" on 03/25/2011. The letter stated that customers can avoid the issues by viewing the recorded ECG for at least 13 seconds. Two diagrams were attached with the letter depicting the conditions under which the issue can occur. Customers can continue to use the device normally in all other respects. The letter states that the firm is working on a software modification which will permanently fix the issue and prevent it from occurring. Customers are to contact the Regulatory Affairs and Operations Manager if they have any questions.
Z-0359-2012 | Dec-02-2011 | 2011 | RayStation - RayStation Version 2.0.0.15 Consult instruction for use. RaySearch Laboratories AB Sveavagen 25; SE-111 34 Stockholm Sweden. support@raysearchlabs.comRayStation is a software system designed for treatment planning and analysis of radiation therapy. The treatment plans provide treatment unit set-up parameters and estimates of dose distributions expected during the proposed treatment; and may be used to administer treatments after review and approval by the intended user. The system functionality can be configured based on user needs. The intended users of RayStation shall be clinically qualified radiation therapy staff trained in using the system. | RAYSEARCH LABORATORIES AB | 2 | The firm has found an issue with DICOM RT Image representation of DRR images from RayStation 2.0 which needs to be clarified in the user manual. The symptoms involve the Pixel Spacing and Image Position attributes in the RT images being incorrect; thus enabling the user to make measurements that are erroneous and possibly misalign the patient. | The firm, RaySearch Laboratories, sent an email with an "Urgent Field Safety Notice Medical Device Correction" notice dated October 19, 2011 to its customers. The notice describes the product, problem and actions to be taken. The customers were instructed to pass on the information in the field safety notice to everyone involved and create routines to compensate for the incorrect DRR images; if the DICOM DRR images are not used to make any form of absolute measurements, no action needs to be taken. Raysearch laboratories will provide a script to update these DICOM files to have the correct values upon request. For further information please contact RaySearch Laboratories AB Quality and Regulatory Affairs Manager at +46 722 366 110 or david.hedfors@raysearchlabs.com; RaySearch Americas Director of Customer Support at +1 877 778 3489 or freddie.cardell@raysearchlabs.com.
Z-2401-2011 | May-31-2011 | 2011 | Xpert Infinity Software - Xpert Infinity Software; V 4.1a.Cepheid; Sunnyvale; CA 94089system software for GeneXpert Infinity 48 | Cepheid | 2 | A mismatch of Patient ID or Sample ID and the Host Order from the hosted LIS and its requested test results could occur when running in batch mode and cartridges are introduced in rapid succession. | Cepheid sent an Corrective Action Notification letter dated April 13, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Cepheid recommended customers using 4.1a software on the INFINITY 48 System that is interfaced to a LIS with Host Ordering to make these alterations in operationof the system. After scanning the Patient ID (optional). Sample ID and Cartridge barcode, the host test order information should apear highlighted in the Order Test workspace. Once the highlighted host order information appears select SUBMIT. Warning. If the host test order informaiton is not highlighted, cancel the order by selecting CLOSE. Re scan Patient ID (optional). Sample ID and cartridge barcode and verify highlighted host test order before SUBMIT. Contact Cepheid Technical Support at 888-838-3222 in the US or 33 5 63 82 53 19 outside of the US for any questions regarding this Corrective Action Notice.
Z-2468-2011 | Jun-08-2011 | 2011 | Newport e360 Ventilator - Newport e360 Ventilator; Model Number: e360The e360 Ventilator System is intended to provide continuous (endotracheal or tracheostomy [ET] tube) or non-continuous (mask) ventilatory support and monitoring for infant; pediatric; and adult patients requiring tidal volumes equal to or greater than 20 milliliters (mL). The device is for use by prescription only. The intended environments include hospital; hospital-type; and intra-hospital transport environments. Hospital use typically includes general care floors; operating rooms; special procedure areas; emergency rooms; and intensive and critical care areas within the hospital. Hospital-type use includes facilities such as or similar to surgicenters; sub-acute centers; and special nursing facilities outside of the hospital. Intra-hospital transport includes patient transport within the hospital or hospital-type facility. | Newport Medical Instruments Inc | 2 | Cregarding e360 ventilators that were shipped with a different software version from what the customers ordered. | Newport Medical Instruments Inc. initiated a field correction on October 27, 2009, to address complaints received regarding e360 ventilators that were shipped with a different software version from what the customers ordered. Newport Medical Instruments provided the correct software to the Newport sales or service providers who were responsible for the units affected. They installed the correct software into the ventilators at the customer sites. No written Correction Notice was provided since the corrections were done at the customers' request. Should you have any questions or require additional information, please call (714) 427-5811 extension 344.
Z-2069-2011 | May-19-2011 | 2011 | Lockheed Martin Aculight Capella R-1850 Laser - Lockheed Martin Aculight Capella-r 1850 Laser. This device has the Accession number 0612339.The Capella R-1850 is a Class 4 solid state diode laser product designed to emit 1850mn laser radiation at maximum pulse energy of 5 mJ (@1ms). The output port connects to a user-selected 200-600 micron diameter optical fiber. The Capella R-1850 is labeled in part: "***LASER APERTURE***Aculight Corporation 22121 20th ave SE Bothell; WA 98021***Feb 2007***Model R-1850; SN: 007***". | Lockheed-Martin Aculight | 2 | It has been determined that there are certain software driven operating sequences involving off-on cycling of the key-switch that do not result in a reset of the Stim Enable button; and emissions can resume after a standard 5 second countdown without pressing this button a second time. | On 11/22/10, Lockheed Martin Aculight sent a Memo to their Customers and Users informing them of the "Stim Enable" button reset error. The defect involves the function of the "Stim Enable" button on the front panel of the Capella R-1850. In December 2010, the firm sent a second lletter with instructions on how to return the Capella R-1850 Laser back for repair. The firm's plan to correct the defect involved removal of existing software, installation of revised software, and documented verification testing, per written instructions, to confirm resolution of the Stim-Emit button failure to reset. Corrections were to be made at no cost to the customer. A Lockheed Martin Aculight customer service representative will contact you shortly and provide detailed information on how to return the Capella-R 1850 laser product in your possession for testing and repair on this defect. Should you have any questions, please call (425) 482-1100.
Z-2000-2011 | Apr-15-2011 | 2011 | LANTIS 6.1 Oncology Information System with MEDOne Module - Siemens LANTIS 6.1 Oncology Information System with MEDOne Module; Linear Medical Accelerators; Product Numbers: 4503178; 5835991; 5851055; Product is manufactured by IMPAC Medical; Inc.; Sunnyvale; CA.Medical charged-particle radiation therapy system. | Siemens Medical Solutions USA; Inc | 2 | The product has a software problem; it does not calculate the Creatinine Clearance correctly for patients less than two years old. | On 3/6/2009, a Safety Advisory Letter was distributed to all affected users, with a description of the problem and identifying the affected product. Customers were instructed not to use the LANTIS to calculate Creatine Clearance for any patient under two years of age. Customers should direct their questions with regards to the letter to their local Siemens Service organization. A Siemens Representative will update the systems.
Z-2119-2011 | Apr-29-2011 | 2011 | 4D Intergrated Treatment Console - The 4D Integrated Treatment Console; Model Number: H51; Reference: 4DITC versions 8.6 and below; File Reference: CP-2007-0288; Product is manufactured and distributed by Varian Medical Systems Inc.; Palo Alto; CA The 4D Integrated Treatment Console function is designed to assist the operator of radiation therapy device in providing accurate treatment setups for each patient by monitoring setup parameters and preventing the radiation therapy device from commencing irradiation when any parameter is out of conformance with the treatment plan. | Varian Medical Systems; Inc. Oncology Systems | 2 | A software anomaly; in that the Multileaf Collimator (MLC) message displayed on the 4DITC on versions 8.6 and below; does NOT interlock beam delivery. | Varian Medical Systems sent an Urgent Medical Device Correction Letter on February 28, 2011, to all affected customers. Varian has created the procedure described in CTB-GE-579 to inform users on proper interfacing of the demountable BrainLAB m3 on the Varian Clinac. The procedure ( CTB-GE-579) is enclosed and posted to the customer site: http://www.MyVarian.com. This design fix in 4DITC v10.1 or later will be provided to all affected sites. If you require further clarification, please feel free to contact your local Varian Customer Support District or Regional Manager. USA and Canada - 1-888-VARIANS (888-827-4265) Europe - + 41 41 749 8844.
Z-0900-2011 | Jan-13-2011 | 2011 | Horizon Cardiology ECG Management - Horizon Cardiology ECG ManagementThis is a software application designed for use in various hospital departments to import; display; store; analyze; distribute; and process ECG procedures from resting ECG devices. It is also intended to provide analysis or reanalysis or testings ECGs and to provide preliminary data for editing and confirmation by an over-reading physician. Horizon Cardiology ECG Management may provide a serial comparison of ECG Data to facilitate the review of the patient's current ECG with previous ECG's of the same patient. | McKesson Provider Technologies - Medical Imaging Group | 2 | When importing ECG files from Mortara ECG carts using DICOM interface; the Horizon Cardiology ECG Management System does not display the amplitudes of the waveform properly. | McKesson Medical sent an advisory notice letter dated June 16, 2010, to all affected customers. Advisory notice ADV01646 was distributed to 13 affected clinical sites from August 30, 2010 to September 28, 2010. The letters identified the product that is affected, the circumstances under which patient safety issue may occur, the risk, and the immediate actions that need to be taken along with a temporary fix. Customers were instructed to use the ECG cart printout to verify results of procedures acquired by Mortara ECG carts. Users should be aware of the problem when referrring to past ECG procedures and when performing a visual comparison with ECG procedures acquired by Mortara ECG carts. Customers should disable the Glasgow Serial Comparison for ECG procedures. McKesson Support would contact the customer to coordinate the deployment of the software update on all the affected applications and Horizon Cardiology client applications. For any questions regarding this recall, customers should call the Customer Support department at 1-866-777-0202 and speak to their Support Manager.
Z-2692-2011 | Jun-28-2011 | 2011 | EOS System - Software versionsEos software versions: 1.2.4; 1.2.5; 1.2.6;1.2.7;1.2.8; 1.2.9;1.2.10; .2.11;1.2.12; 1.2.13; 1.3.2;1.3.3; 1.3.4; 1.4.0; 1.4.1; 1.4.2.sterEOS software versions: 1.2.3.3788; 1.2.4.4286; 1.3.2.3373; 1.3.3.3671; 1.3.4.3740; 1.3.5.3915; 1.3.6.4272;1.4.1.4813;1.4.2.4850.EOS system is a digital radiography system used in general radiographic examinations. sterEOS workstation is a medical image review station for acceptance; transfer; display; storage; and digital processing of 2D X-ray images. | Eos Imaging | 2 | Print true size; print requests from the EOS acquisition station or the sterEOS image review station may; in certain cases; be affected by a different enlargement factorduring printing. T | EOS imaging sent a "FIELD SAFETY NOTICE" letter dated April 13, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. The letter instructs customers to wait for the printing queue to empty before sending a second "print true queue" request with different print settings. Customers should then click the button shown on the instruction sheet to view the print queue. EOS imaging will provide a software update to correct all affected systems at no cost to the customers. An EOS representative will contact customers to schedule a time to upgrade the software. Questions can be answered by calling 33 1 55 25 63 90.
Z-0345-2012 | Dec-01-2011 | 2011 | Odyssey - OdysseyA radiation treatment planning system | PerMedics; Inc. | 2 | Anomaly 1 - Invalid Export of Beam Table and or Collimator Angle for DICOM RT Ion; DICOM Image and RTP Link export files:During PerMedics' internal code reviews for lEC 61217 compliance verification; an issue was identified by PerMedics engineers in three export files: D1COM RT Ion Plan; D1COM RT Image and RTP Link. In some cases; this issue can result in table and/or collimator angle values t | PerMedics sent an Odyssey TPS URGENT NOTICE OF CORRECTION letter dated November 4, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to: Anomaly 1 User Corrective Action: Install the mandatory new version correcting the software as soon as it is available (see PerMedics Corrective Action below). In the meantime, the following recommendations were provided if they planned to use Odyssey RT Ion Plan, DICOM RT Image, and/or RTP Link export: Workaround for the invalid exported table and/or collimator angle errors: If, prior to the software correction, any of these affected file types were exported from Odyssey, manually verify that the table and collimator angles are correct. Additionally, apply appropriate changes to any angles deemed incorrect that were obtained from Odyssey's DICOM RT Ion Plan, DICOM RT Image, and/or RTP Link files. Only those machines with table and/or collimator angles that increase in the clockwise direction will have the incorrect angles. Anomaly 2 User Corrective Action: Install the mandatory new version correcting the software as soon as it is available (see PerMedics Corrective Action below). In the meantime, the following recommendations are provided if the user - Changes the patients scheduled start date prior to the software correction: Workaround for Incorrect Monitor Unit Values when Changing Treatment Start Date: If, prior to the software correction, the user changes the patient's treatment start date on any fully-computed IMRT plan, IMAT plan or plan containing a "Region DVH" prescription, the user should force a full recommendation oft hose patient plans prior to fixing and exporting the plan. The following procedure can be used to force a full re-computation of the plan: - Open the patient's plan worksheet - Change the algorithm type of the plan - Change the algorithm type back to its original value - Close t.
Z-2563-2011 | Jun-14-2011 | 2011 | GE Centricity Laboratory Instrument Interface Data Innovations Instrument Manager Software - GE Centricity Laboratory Instrument Interface to Data Innovations Instrument Manager software; GE Healthcare IT Solutions; Barrington; IL 60010. The Centricity Laboratory is designed to facilitate the general clinical; anatomic pathology and cytology laboratory workflow such as order entry; results entry; instrument interfacing; results reporting and patient record retrieval. It integrates with other Hospital Information Systems (HIS) through HL7 interface by order entering and report processing. The centricity is intended to interface with various lab devices through standard HL7 interfaces to record; report and distribute lab results. | Ge Healthcare It | 2 | The GE Centricity Laboratory Instrument Interface to Data Innovations is processing a preliminary result instead of the final result from the Process System Manager when there is a test rerun conducted. | GE Healthcare sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated May 3, 2011 to all affected customers. The letter describes the product, problem, and action to be taken by the customer. The letter provided the users with specific safety instructions to follow until the software was updated. The customers were requested to complete the enclosed customer reply form and fax the completed form to 847-620-2780. Any questions should be directed to their GE Service Representative or the GE Help desk at 888-778-3375.
Z-3091-2011 | Aug-30-2011 | 2011 | Rapid Programmer - Software version 3.3 for Rapid Programmer devices; Model 3832.Software version 3.3 for Rapid Programmer devices; a personal digital assistant (PDA) based programmer supporting implantable pulse generators of the Eon; Genesis; and Renew family. | Advanced Neuromodulation Systems | 2 | Product programmed with unapproved software was distributed. | St. Jude Medical Neuromodulation contacted all affected customers by telephone between June 2, 2009 and June 5, 2009. Customers were given information about the product, problem and action to be taken. All affected devices were returned to the firm as instructed.
Z-0355-2012 | Dec-02-2011 | 2011 | SoftPath GUI Versions: 4.3.0.10; 4.3.0.11; and 4.3.0.12. - SoftPath GUI Versions: 4.3.0.10; 4.3.0.11; and 4.3.0.12. SoftPath is a laboratory information system to be used in a clinical laboratory by knowledgeable; trained; and experienced personnel. SoftPath enables the user to document steps in the receiving and processing of surgical; gynecological; medical cytological; and autopsy case specimens within the Pathology labs; for the purpose of generating Pathology reports for requesting physicians and label creation for identification of specimen/case material. | SCC Soft Computer | 2 | On 06/13/2011; SCC Soft Computer initiated a correction on the following SoftPath GUI Versions: 4.3.0.10; 4.3.0.11; and 4.3.0.12. A client reported when they verified the contents of the fax reports; they found one of the specimen diagnoses to be missing. Parts A through D were present as well as Part F; however; Part E diagnosis was missing for one of the patient reports. All affected clients wer | SCC Soft Computer initiated a Field Correction on June 13, 2011. All customers with affected software have been notified of this issue. SCC recommends until the corrected patch is loaded to temporarily use one of the following alternative working solutions: 1. Use MS Word Editor functionality within SoftPath application for creating patient reports. In order to ensure Text Control editor is not inadvertently activated, the templates for Text Control editor will be deleted and reinstalled at the clients request by their SCC representative after receipt of the correction. 2. In Report Template set-up, the bottom margin should be made slightly larger than the size of the footer plus its distance from the bottom page border. For further explanation, please consult your SCC representative for further details, or call (727) 789-0100.
Z-1723-2011 | Mar-22-2011 | 2011 | Elekta Synergy XVI R4.5 - Elekta Synergy XVI R4.5Radiation therapy treatment of malignant neoplastic diseases; as determined by a licensed medical practitioner. | Elekta; Inc. | 2 | If another patient is selected on the imaging system (XVI) release 4.5) during the transmission of end of treatment data from Desktop Pro R7.01 to R&V system; the XVI information is prioritized and the end of treatment data is never received by the R&V system. | Important Notice A341 titled "Mandatory Upgrade Desktop Pro R7.01 Service Pack and Cancelled Notices" dated January 26, 2011, is in distribution to all affected customers. Field Change Order 20000501011, "Service Pack 2 upgrade to Desktop Pro R7.01 Service Pack 1", dated October 25, 2010 will be utilized by Elekta Service personnel to upgrade the Desktop Pro to include Service Pack 2. The Important Notice letter instructed customers to file the letter in the Important Notice section of the appropriate User Manual. Before doing so, users are to ensure that Desktop Pro R7.01 SP2 is installed on their treatment control system. If it is not installed, customers are to contact their local Elekta representative. Once the installation of Service Pack 2 has taken place, the notices listed in Table 1 of the letter will no longer be applicable and should be removed from the User Manual. Customers should direct their questions to their local Elekta representative.
Z-0278-2012 | Nov-23-2011 | 2011 | Softlab - SCC Soft Computer; Softlab Laboratory information System with DORA versions: 2.0.0.4 with Hot Fix 1.2808; 1.2809; 1.2956; or 1.3476; 2.0.0.5; 2.0.0.6; 2.0.0.7 with Hot Fix 1.6595; and 2.1.0.3. | SCC Soft Computer | 2 | Potential for patient data to be sent to the wrong patient record. A software change affected how the Softlab handles recovery from certain types of table access conflicts. When such conflicts occur; this function has the potential to cause faulty writing to the database. | Clients were notified through SCC Soft Computer's proprietary Task Management System On 03/22/201. All clients have received a hot fix that corrected the issue or made the hosparam change that prevents the issue.
Z-2569-2011 | Jun-15-2011 | 2011 | Monaco Radiation Treatment Planning Workstation - Monaco Radiation Treatment Planning Workstation. Monaco Release 2.04.00 and above.Used to create treatment plans for any cancer patient for who external beam intensity modulated radiation therapy (IMRT) has been prescribed. | Computerized Medical Systems Inc | 2 | Monaco: In XiO IMRT (Dynamic MLC Delivery) software; the radiation beam angle is set to zero. If different beam angles are selected via Monaco software; the beam angles will remain at zero. | Computerized Medical Systems, Inc. sent a User Notice to the affected consignees on October 27, 2010, explaining the reason for correction, the clinical impact, and the workaround method. (VERIFY USE NOTICE DATE) A return postcard was included for the customer to confirm receipt of the User Notice. The firm intends for this issue to be corrected in Monaco Release 3.00 in March 2011.
Z-0344-2012 | Dec-01-2011 | 2011 | GEMINI TF 64 Diagnostic X-ray/PET Imaging System - Philips GEMINI TF 64 Diagnostic CT X-ray/PET Imaging System; Catalog/Model #882471; 510(k) #K052640. System units equipped with software version: 3.5.2 are subject to recall/field correction. The GEMINI TF I6; and TF 64 are diagnostic imaging systems for fixed or mobile installations that combine Positron Emission Tomography (PET) and X-ray Computed Tomography (CT) systems. The CT subsystem produces cross-sectional images of the body by computer reconstruction of x-ray transmission data. The PET subsystem produces images of the distribution of PET radiopharmaceuticals in the patient body (specific radiopharmaceuticals are used for whole body; brain; heart and other organ imaging). Attenuation correction is accomplished by CT AC. The device also provides for list mode; dynamic; and gated acquisitions. Both subsystems (PET and CT) can also be operated independently as fully functional; diagnostic imaging systems including application of the CT scanner as a radiation therapy simulation scanner. | Philips Medical Systems (Cleveland) Inc | 2 | In response to complaints received by Philips Medical concerning certain identified software errors reported by their customers; the firm is conducting a recall/software correction to several of their GEMINI PET/CT X-ray system units in the field. The firm believes that if these software defects were to recur; it could impact the performance of reliability of the these PET/CT system units. | Philips Healthcare sent a "CPE CUSTOMER LETTER" dated October 28, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. Additionally, device operation information is provided to customers. A Philips Field Service Engineer will be visiting each customer site to install the necessary software upgrades. Contact the Customer Care Solutions Center at 1-800-722-9377 for questions concerning this notice.
Z-0139-2012 | Oct-31-2011 | 2011 | Picis Anesthesia Manager - Picis Anesthesia Manager; PACU Manager;Critical Care Manager-8.2 Service Pack 4 or higher (Software; transmission & storage; patient data)This product is a non-alarming software patient information system that compiles an electronic medical record utilizing commonly available hardware. A patient record is populated with information from various sources; such as health care professionals; medical devices connected to the system; and data that arrives via hospital and laboratory information systems. The application stores this information in a database and displays it to healthcare providers in different formats for informational purposes. The product is intended for use by healthcare professionals. | Picis Inc. | 2 | Periop/Critical Care Manager EHR Software Application displayed one patient's demographic and health information in place of another patient that was in current view. | Picis Inc. sent a "Safety Notification letter" dated August 12, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The customers were instructed until this service pack is applied, we recommend: * For patients who have been transferred from other facilities within the same VISN, you should discontinue the use of these Picis applications. Manually check the patient's name with that displayed in the Picis patient blue banner to ensure you catch all of these cases. * For all other patients, you may safely continue the use of these Picis applications. For further questions please call (781) 557-3000.
Z-2206-2011 | May-19-2011 | 2011 | Brilliance 64 CT X-Ray System - Brilliance 64; Computed Tomography X-Ray System; Model Numbers: 728231 and 728326; including the Ingenuity CT. The recalled units are equipped with software versions: 3.5.21028 and 3.5.25028. Intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. | Philips Medical Systems (Cleveland) Inc | 2 | A program software error in the algorithm produces image artifacts that could possibly lead to a misdiagnosis of patients who have small metal implants. | A 'Field Safety Notice - Urgent - Medical Device Correction' notification was sent to the firm's customers via Certified mail on 3/8/2011. The notification letter explains the nature of the Metal Artifact Reduction/Metal Artifact Correction software error; the Brilliance CT X-Ray systems subject to correction and how to identify them; and the circumstances under which this error can occur. It also asks that the user stop running the MAR application on any of the affected Brilliance systems identified with the recalled serial numbers and software versions alluded to in the medical device correction notification. Lastly, the notification states that Philips Medical will provide a solution that will fix the issue and the notification lists contact names; telephone numbers; and E-mail addresses which the customer can contact for further information concerning this device issue, or customer support assistance.
Z-3199-2011 | Sep-13-2011 | 2011 | GEMINI TF16 Diagnostic Imaging System - Philips GEMINI TF16 Diagnostic Imaging Systems; Model Number: 882470; 510(k) #K052640; Serial Numbers: 7501; 7503; 7507; 7508; 7011; 7047; and 7174. The recalled GEMINI TF16 system units equipped with software versions: 3.5.1.1 and 3.5.2.1 are subject to recall/software correction. The GEMINI TFI6; TF64 and TF Big Bore (BB) are all diagnostic imaging systems for fixed or mobile installations that combine Positron Emission Tomography (PET) and X-ray Computed Tomography (CT) systems. The CT subsystem produces cross-sectional images of the body by computer reconstruction of x-ray transmission data. The PET subsystem produces images of the distribution of PET radiopharmaceuticals in the patient body (specific radiopharmaceuticals are used for whole body; brain; heart and other organ imaging). Attenuation correction is accomplished by CT AC. The device also provides for list mode; dynamic; and gated acquisitions. Both subsystems (PET and CT) can also be operated independently as fully functional; diagnostic imaging systems including application of the CT scanner as a radiation therapy simulation scanner. | Philips Medical Systems (Cleveland) Inc | 2 | The firm is conducting a recall/software correction to several of their GEMINI PET/CT X-ray system units in the field. The firm believes that if these software defects were to recur; it could impact the performance of reliability of the these PET/CT system units. | The firm , Philips Healthcare, issued a 'Customer Information Letter' dated August 9, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were not required to perform any action. A Philips' Field Service Engineer will visit each customer site and install the necessary software upgrade and Philips will track the progress of the software corrections as they take place. For customers in North America, if you need any further information or support concerning this update, please contact our Customer Care Solutions Center at 1-800-722-9377. Selection option 5 for "All Imaging Systems". Enter your site ID number. If you do not know your site ID, please hold and then select option 5 for "Nuclear Medicine" and finally select option 2 for "PET" support. In all other countries, the local Philips Healthcare office should be contacted at 440-483-3000.
Z-1435-2011 | Feb-25-2011 | 2011 | RM; RMAT - IMPAX CV Results Manager/Results Manager Administration ToolThe IMPAX Cardiovascular suite is a cardiovascular information system providing image archiving and image display and modality study-specific structured reporting. The suite is comprised of modular software that runs on industry standard personal computers and servers; utilizing Windows 2000/2003/XP operating systems. Image data storage/display utilizes DICO< (Digital Imaging and Communications in Medicine) data primarily from independent medical devices. | AGFA Corp. | 2 | User could modify a measurement's units and that change would not be reflected in echocardiography formulas used within the Adult echocardiography RM. | The firm, Agfa Healthcare, sent individual letters dating from April 15, 2009 to August 26, 2009, to the customers. The letters described the product, problem and actions to be taken. Agfa provided the affected customer with a complete list of their impacted patients and the measurements. Agfa also scheduled unit corrections with their customers and received acknowledgement of the letters at the time of scheduling the unit corrections. Service Bulletin DIS026.09E was issued January 2009 to inform users that Agfa will no longer allow customers to request unit changes for measurements that have formulas. Note: Technical Support ran diagnostic scripts for each customer and documented in GSC cases for each consignee/customer. If you have any questions, contact Agfa Healthcare at (949) 830-4959.
Z-3208-2011 | Sep-14-2011 | 2011 | Brilliance iCT CT X-Ray System - Brilliance iCT Computed Tomography X-Ray System; 510 (k) #K060937; Model #728306; Serial Numbers: 100019; 100023; 100087; 100103; and 100170. The recalled system units are equipped with operating software version; 3.2.0.19011.The Brilliance ICT and Brilliance ICT SP are Computed Tomo grapy X-ray Systems intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. These devices may include signal analysis and display equipment; patient; and equipment supports; components and accessories. | Philips Medical Systems (Cleveland) Inc | 2 | Software errors. Philips Medical's Field Service Engineers are installing software upgrades to certain of their Brilliance iCT and Brilliance ICT SP CT X-ray system units in order to correct non-conformances evidenced in these systems due to software errors. | Philips Medical ( Cleveland) Inc. indicated in the Correction and Removal Report dated June 21, 2011, that a Philips Field Service Engineer will be contacting each site in order to schedule a time to implement the Field Change Order. After which time, the Field Service Engineer will visit each customer site to install the software upgrade and the firm will be tracking the progress of this corrective action. For further questions please call (440) 483-7000.
Z-0360-2012 | Dec-02-2011 | 2011 | Brilliance CT X-Ray System - Brilliance 64 Computed Tomography System; Model #728231; 510(k) #K 033326; Serial Numbers: 95504; 91003; 300030; 95130; and 95157.The Brilliance 64 is a Computed Tomography X-Ray System intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at difference angles and planes. This device may include signal analysis and display equipment; patient; and equipment supports; components and accessories. | Philips Medical Systems (Cleveland) Inc | 2 | Software errors. A number of non-conformances were logged during in-house testing and external validation testing of Brilliance 64 and Ingenuity CT software version 4.0 Beta 2. The non-conformances are due to software errors identified in these device systems. | Philips Healthcare released a 'Field Change Order' (FCO #72800542) on September 15, 2011. Philips sent an URGENT - Medical Device Correction letter dated October 10, 2011, to all affected customers. The letter listed the affected device systems; described the nature of the software problem; and any possible hazard inherent to the situation. The letter was mailed by Philips to their customers on October 27, 2011. The letter provided information telling the customer how to identify the affected products and what action to take. In this instance, the customer was asked to review the Customer Release Notes which were provided to each customer location once the software update was installed. The letter notified the customer that a Philips representative would be installing a software upgrade in order to address the referenced software issues free-of-charge, through a Field Change Order. The letter closes by providing information which all customers in Canada and North America can use to contact the firm through a toll-free telephone number 1-800-722-9377, option 5; enter side ID or follow the prompts. For questions regarding this recall call 440-483-4918.
Z-0362-2012 | Dec-02-2011 | 2011 | SoftReports - SoftReports version 1.1.8.2.4 with Hot Fixes 1.7199.7 or 1.7200.7. SoftReports is a web-based report launching tool to be used in a hospital or laboratory environment by knowledgeable; trained; and experienced personnel. SoftReports enables the user to schedule and generate reports for the purpose of viewing and printing data stored in SCC product databases. | SCC Soft Computer | 2 | On 10/25/2011 SCC Soft Computer; Clearwater; FL initiated a correction on the SoftReports version 1.1.8.2.4 with Hot Fixes 1.7199.7 or 1.7200.7. When SoftReports Designer opened a layout; the system did not load the layout with the previously saved value in the NewPage property. This caused multiple patient records to be printed on the same page. All affected clients were notified on the issue on | SCC Soft Computer sent a "Correction Notification" letter dated October 25, 2011 to all affected customers. The letter described the product, problem, and actions to be taken by the customers. The hot fixes that were previously loaded which introduced this issue were corrected and delivered to all affected customers as a mandatory correction in hot fixes 1.7199.8 and 1.7200.8 and higher.
Z-2771-2011 | Jul-11-2011 | 2011 | UniCel DxC 880i; 860i;680i and 660i Synchron Access Clinical Systems - UniCel DxC 880i; 860i;680i and 660i Synchron Access Clinical SystemsThe Unicel DxC Integrated Workstation takes samples loaded from a single point of entry through a UCTA (UniCel Closed Tube Aliquotter). The UCTA functions as the sample processing manager by aliquotting and routing samples to the DxC and the Dxl analyzers according to the programming requirements | Beckman Coulter Inc. | 2 | The recall was initiated after Beckman Coulter confirmed reports of inaccurate Access Intrinsic Factor Ab test results displayed on; or printed from; the UniCel Dxl portion of an integrated workstation listed above. The problem is limited to numerical Intrinsic Factor Ab results produced on integrated systems running UniCel Dxl system software version 4.0; 4.1; or 4.2. The DxC portion of your inte | Beckman Coulter, Inc. sent out an "Urgent: Product Corrective Action" letter dated May 19, 2009 to all affected customers. The letter explained the reason for recall and provided recommendations including: -System software currently under development will correct this problem, and will be made available soon. Until then, please refer to UniCel DxC reports, user interface information, and LIS transmissions, which contain accurate Access Intrinsic Factor Ab test results. -Refer to the UniCel DxC system documentation for information on displaying and printing results from the DxC portion of an integrated workstation, or from your laboratory information system (LIS) -Past results for these assays only need to be reviewed if they were reported from the Dxl portion of an integrated workstation -Complete and return the enclosed response form For additional information please contact the Customer Technical Support at 1-800-854-3633, or your local Beckman Coulter representative.
Z-1820-2011 | Mar-24-2011 | 2011 | Brilliance 64 Computed Tomography X-Ray System Model #728231; - Brilliance 64 S/N 9089; 9551; 9548.The Brilliance 64 is a Computed Tomography X-Ray System intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. | Philips Medical Systems (Cleveland) Inc | 2 | During the version 2.6 software upgrade; the system was not recalibrated with the infant phantom for the infant protocol. | Philips sent an URGENT - Medical Device Correction letter dated January 31, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers may consider reviewing any infant images that may have been affected during the Field Test to confirm reading accuracy. For further information or support concerning this issue, customers could contact their local Philips representative or local Philips Healthcare office for North America and Canada contact the Customer Care Solutions Center (1-800-722-9377, option 5: Enter Site ID or follow the prompts).
Z-0335-2012 | Dec-01-2011 | 2011 | Brilliance CT X-Ray System - Brilliance 64 Computed Tomography X-Ray System; 510(k) #K033326; Model Number: 728231; Serial Numbers: 9745; 90169; and 95414. The recalled device systems are equipped with Version 3.5 Beta 1 software. | Philips Medical Systems (Cleveland) Inc | 2 | Software errors. A number of non-conformances were found for the Brilliance Big Bore and Brilliance 64 Computed Tomography X-ray Systems equipped with software version; 3.5 Beta1. The non-conformances stem from software errors identified in the operating software. | Philips Healthcare released a 'Field Change Order' (FCO #72800531) on September 12, 2011. Philips sent an - URGENT Medical Device Correction letter dated October 17, 2011, to all affected customers. The letter listed the affected device systems; described the nature of the software problem; and any possible hazard inherent to this situation. The letter informed customers on how to identify the affected products and what action to take. In this instance, the customer was asked to review the Customer Release Notes which were provided to each customer location once the software update is installed. The letter notified the customer that a Philips representative would be installing a software upgrade in order to address the referenced software issues free-of-charge, through a Field Change Order. Customers were instructed to contact their Philips representative or local Philips Healthcare Office for further information or support for this issue. For North America and Canada customers were instructed to contact the Customer Care Solutions Center (1-800-722-9377, option 5: Enter Site ID or follow the prompts).
Z-0284-2012 | Nov-23-2011 | 2011 | Symbia T Series SPECT-CT System - Siemens Medical Solutions USA; Inc.; Symbia T Series SPECT-CT (Single-Photon Emission Computed Tomography and Computed Tomography) System. | Siemens Medical Solutions USA; Inc. | 2 | Errors could potentially impact image interpretation. The acquisition software is exhibiting a defect in the energy window position when an isotope is peaked. There is a potential patient safety issue when running syngo MI Applications version VA60A with Foresight detectors firmware version v1.5.9.12 on the Symbia systems. | Siemens Medical Solutions USA, Inc. sent an "Urgent Field Correction Recall" letter dated October 14, 2011 to all affected customers. The letters informed customers of the acquisition software defect in the energy window position when an isotope is peaked, which could potentially impact image interpretation, and that the AEB firmware version on their imaging system has been changed to a previously released version until acquisition software service pack is available. For questions or assistance they were instructed to contact Siemens Medical Solutions USA, Inc. at (800) 888-7436.
Z-3183-2011 | Sep-09-2011 | 2011 | LATITUDE¿ Patient Management System Communicator - LATITUDE¿ Patient Management System Communicator; Model 6482.The LATITUDE Patient Management system is intended for use to remotely communicate with a compatible Boston Scientific pulse generator and transfer data to a central database. | Boston Scientific CRM Corp | 2 | Boston Scientific has identified two occurrences where the LATITUDE system did not process a data payload; which is the information uploaded to the LATITUDE system from a remote interrogation of the implanted device. These occurrences resulted in non-displayed payloads because the data is not posted to the website. If the payload is not displayed; the pertinent data will not be available on the | Boston Scientific notified all affected customers by telephone on May 25, 2011. The customers were informed that there was an interrogation for one of their patients that was not able to be displayed on the website. Boston Scientific detected that a payload failure had occurred, but could not conclude what information was contained within the payload. Each patient has had previous and subsequent payloads that have processed and displayed on the LATITUDE website.
Z-0307-2012 | Nov-30-2011 | 2011 | Elekta Leksell Gamma - Elekta Leksell Gamma PlanProduct Usage: Gamma Plan is a computer based system designed for Gamma Knife treatment planning. | Elekta; Inc. | 2 | There is a risk that the user may interpret the instructions in the manual too literally; and adjust the TMR 10 dose description to obtain an ETCPD that matches previous clinical prescriptions; which in turn could lead to a prescription dose that deviates a few percent from the intended dose. | Elekta sent an IMPORTANT NOTICE letter dated June 23, 2011 to affected customers. The notice identified the product, problem, and corrective actions to be taken by the user. The letter advise all customers to follow any instructions or recommendations covered in the Notice. The letter states that Elekta is working on a new software version and customers will receive the future version when released. For questions contact your local Elekta representative.
Z-2064-2011 | May-16-2011 | 2011 | Defibtech Lifeline AED and ReviveR AED - Defibtech Lifeline AED and ReviveR AED; Model: DDU-100 Series AEDsSoftware version 2.004 or earlierThe product under correction is the DDU-100 AED. The AED is indicated for use on victims of sudden cardiac arrest (SCA) when the patient is unconscious; unresponsive and not breathing. The indications for use also state that for patients under 8 years old; child/ infant electrode pads are to be used and that therapy should not be delayed to determine the patients exact age or weight. The AED must also be used by or on the order of a physician. These AEDs are sold under the brand names Lifeline AED and ReviveR AED. | Defibtech; LLC | 1 | Software with version 2.004 software or earlier may cause an affected AED to cancel shock during the charging process | The firm, Defibtech, sent an "URGENT MEDICAL DEVICE SAFETY INFORMATION AND CORRECTION" letter dated March 11, 2011 to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to do the following: 1) Remove the pads from the storage pouch on the back of their DDU-100 series AED and inspect the 9-digit serial number label. 2)Carefully note the 9-digit serial number on the back of the AED, then visit the Defibtech website at www.defibtech.com/fa11 and enter their serial number to determine if their AED is affected; email their serial number and contact information to fa@defibtech.com or call Defibtech directly at 1-877-453-4507 or 1-203-453-4501 and Defibtech will inform them if their unit is affected. 3) If their AED is affected, confirm receipt of this notification with Defibtech as soon as possible by website and/or confirmation card. Complete and return the confirmation form via fax to 1-203-738-1072. Note: "it is very important for you to contact us either through the Defibtech website or return this card to confirm that you have received this notice. This information will expedite the process of upgrading your AEDs." If your DDU-100 series AED serial number DOES NOT fall within the listed range, your AED is not affected by this recall. A letter was also sent to the distributors and direct accounts that have received the affected product by first class mail with a prepaid, return written acknowledgement card. The field action consists of a software change to the AED. All affected AEDs operating with software version 2.004 or earlier are to remain in the field and in service throughout the field action process. Upon receipt of contact information from affected customers, Defibtech will send an upgrade kit to customers consisting of a software data card, instructions on how to perform the upgrade, and a pre-paid upgrade confirmation postcard that must be returned to Defibtech upon completion.
Z-0051-2012 | Oct-18-2011 | 2011 | Outlook ES Safety Infusion System - Outlook ES pump is intended for use with B. Braun Medical Inc. Horizon Pump IV Sets. Used to regulate the flow of primary and secondary fluids when positive pressure is required. The infusion system is capable of delivering fluid from a negative head height (when the IV fluid container is lower than the pump); and provides clinically accepted volumetric accuracy for all standard IV fluids; including blood; lipid; and Total Parenteral Nutrition (TPN). Model number: 621-300ES; 621-400ES; 621-400ESR.To deliver a broad range of drugs with use with the B. Braun Medical Inc. Horizon Pump IV Sets. | B Braun Medical; Inc | 1 | B Braun Outlook ES Safety Infusion System "Run" Light Emitting Diodes (LEDs) pumps stop infusing and a backup alarm sounds; but the "Run" LEDs advance as if the pumps were infusing. | B. Braun Medical, Inc. sent an "URGENT: FIELD CORRECTION" letter dated September 7, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. A firm representative will contact all affected customers in order to schedule the field correction. The correction consists of revising the Door and Main Processor software, inspecting the main board, and replacing the board, if warranted. A Field Correction Information Form was included in the letter for customers to complete and return. Contact Technical Product Support at 1-800-627-7867 for questions regarding this notice.
Z-0837-2011 | Jan-10-2011 | 2011 | B. Braun Outlook - Outlook 400ES Safety Infusion System; Model Number 621-400ES. For use with B. Braun Medical Inc. Horizon Pump IV Sets | B Braun Medical; Inc | 1 | Possible pump stop with audible alarm; but no visual error indication that the pump is not delivering the medication. Under certain conditions in a network environment; Management Processor may become non-responsive; ceasing normal operation. | B. Braun Medical Inc. notified consignee by letter on August 24, 2010.